## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 282
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 13.7087 | 0.32712 | 13.0544 | 14.3629 | 12.7273 | 14.6901 | 2.386% | 4.772% | 7.159% | 13.0233 | 14.3941
PhotonAssay
Au, ppm | 14.0508 | 0.25668 | 13.5374 | 14.5641 | 13.2807 | 14.8208 | 1.827% | 3.654% | 5.48% | 13.3482 | 14.7533
4-Acid Digestion
Ag, ppm | 0.61106 | 0.021483 | 0.56809 | 0.65403 | 0.54661 | 0.67551 | 3.516% | 7.031% | 10.55% | 0.58051 | 0.64161
Al, wt.% | 2.9473 | 0.11414 | 2.719 | 3.1756 | 2.6049 | 3.2897 | 3.873% | 7.745% | 11.62% | 2.8 | 3.0947
As, ppm | 2327.7175 | 102.5313 | 2122.655 | 2532.7801 | 2020.1237 | 2635.3114 | 4.405% | 8.81% | 13.21% | 2211.3317 | 2444.1034
Ba, ppm | 99.7694 | 2.9737 | 93.822 | 105.7169 | 90.8482 | 108.6907 | 2.981% | 5.961% | 8.942% | 94.781 | 104.7579
Be, ppm | 0.92801 | 0.056307 | 0.81539 | 1.0406 | 0.75909 | 1.0969 | 6.068% | 12.14% | 18.2% | 0.88161 | 0.97441
Bi, ppm | 0.93214 | 0.049625 | 0.83289 | 1.0314 | 0.78326 | 1.081 | 5.324% | 10.65% | 15.97% | 0.88553 | 0.97874
Ca, wt.% | 17.4672 | 0.90998 | 15.6472 | 19.2871 | 14.7372 | 20.1971 | 5.21% | 10.42% | 15.63% | 16.5938 | 18.3405
Cd, ppm | 1.0798 | 0.051382 | 0.97707 | 1.1826 | 0.92569 | 1.234 | 4.758% | 9.517% | 14.27% | 1.0258 | 1.1338
Ce, ppm | 36.8927 | 1.8874 | 33.1179 | 40.6676 | 31.2305 | 42.555 | 5.116% | 10.23% | 15.35% | 35.0481 | 38.7374
Co, ppm | 3.7994 | 0.24767 | 3.3041 | 4.2948 | 3.0564 | 4.5425 | 6.519% | 13.04% | 19.56% | 3.6095 | 3.9894
Cr, ppm | 32.5159 | 6.3052 | 19.9054 | 45.1263 | 13.6002 | 51.4315 | 19.39% | 38.78% | 58.17% | 30.8901 | 34.1417
Cs, ppm | 8.2399 | 0.38673 | 7.4664 | 9.0134 | 7.0797 | 9.4001 | 4.693% | 9.387% | 14.08% | 7.8279 | 8.6519
Cu, ppm | 19.1511 | 1.3813 | 16.3885 | 21.9137 | 15.0072 | 23.2951 | 7.213% | 14.43% | 21.64% | 18.1936 | 20.1087
Dy, ppm | 2.6079 | 0.12097 | 2.3659 | 2.8498 | 2.2449 | 2.9708 | 4.639% | 9.277% | 13.92% | 2.4775 | 2.7383
Er, ppm | 1.461 | 0.058203 | 1.3445 | 1.5774 | 1.2863 | 1.6356 | 3.984% | 7.968% | 11.95% | 1.3879 | 1.534
Eu, ppm | 0.65426 | 0.032191 | 0.58988 | 0.71864 | 0.55769 | 0.75083 | 4.92% | 9.84% | 14.76% | 0.62155 | 0.68697
Fe, wt.% | 1.0641 | 0.039954 | 0.98424 | 1.1441 | 0.94428 | 1.184 | 3.755% | 7.509% | 11.26% | 1.0109 | 1.1174
Ga, ppm | 7.1461 | 0.45916 | 6.2278 | 8.0645 | 5.7686 | 8.5236 | 6.425% | 12.85% | 19.28% | 6.7888 | 7.5034
Gd, ppm | 3.0464 | 0.19697 | 2.6524 | 3.4403 | 2.4555 | 3.6373 | 6.466% | 12.93% | 19.4% | 2.8941 | 3.1987
Hf, ppm | 1.6265 | 0.14075 | 1.345 | 1.908 | 1.2042 | 2.0487 | 8.654% | 17.31% | 25.96% | 1.5452 | 1.7078
Ho, ppm | 0.50778 | 0.030613 | 0.44655 | 0.569 | 0.41594 | 0.59962 | 6.029% | 12.06% | 18.09% | 0.48239 | 0.53317
In, ppm | 0.059486 | 0.0039973 | 0.051491 | 0.06748 | 0.047494 | 0.071477 | 6.72% | 13.44% | 20.16% | 0.056511 | 0.06246
K, wt.% | 1.8512 | 0.093467 | 1.6643 | 2.0382 | 1.5708 | 2.1316 | 5.049% | 10.1% | 15.15% | 1.7587 | 1.9438
La, ppm | 20.5254 | 1.4698 | 17.5858 | 23.465 | 16.116 | 24.9348 | 7.161% | 14.32% | 21.48% | 19.4992 | 21.5517
Li, ppm | 21.5676 | 1.1628 | 19.242 | 23.8933 | 18.0792 | 25.0561 | 5.392% | 10.78% | 16.17% | 20.4892 | 22.646
Lu, ppm | 0.18917 | 0.016089 | 0.15699 | 0.22135 | 0.1409 | 0.23743 | 8.505% | 17.01% | 25.52% | 0.17971 | 0.19862
Mg, wt.% | 3.3938 | 0.16418 | 3.0655 | 3.7222 | 2.9013 | 3.8864 | 4.838% | 9.675% | 14.51% | 3.2241 | 3.5635
Mn, wt.% | 0.025112 | 0.00085573 | 0.0234 | 0.026823 | 0.022544 | 0.027679 | 3.408% | 6.815% | 10.22% | 0.023856 | 0.026367
Mo, ppm | 13.5416 | 0.50909 | 12.5234 | 14.5598 | 12.0143 | 15.0689 | 3.759% | 7.519% | 11.28% | 12.8645 | 14.2187
Na, wt.% | 0.065404 | 0.0066848 | 0.052035 | 0.078774 | 0.04535 | 0.085459 | 10.22% | 20.44% | 30.66% | 0.062134 | 0.068674
Nb, ppm | 5.7269 | 0.30726 | 5.1124 | 6.3414 | 4.8051 | 6.6487 | 5.365% | 10.73% | 16.1% | 5.4406 | 6.0132
Nd, ppm | 17.4517 | 0.73088 | 15.9899 | 18.9134 | 15.259 | 19.6443 | 4.188% | 8.376% | 12.56% | 16.5791 | 18.3243
Ni, ppm | 21.8135 | 1.4372 | 18.9391 | 24.6879 | 17.502 | 26.125 | 6.588% | 13.18% | 19.77% | 20.7228 | 22.9042
P, wt.% | 0.031805 | 0.0019798 | 0.027846 | 0.035765 | 0.025866 | 0.037745 | 6.225% | 12.45% | 18.67% | 0.030215 | 0.033396
Pb, ppm | 18.7597 | 1.136 | 16.4876 | 21.0318 | 15.3516 | 22.1679 | 6.056% | 12.11% | 18.17% | 17.8217 | 19.6977
Pr, ppm | 4.6041 | 0.1744 | 4.2553 | 4.9529 | 4.0809 | 5.1273 | 3.788% | 7.576% | 11.36% | 4.3739 | 4.8343
Rb, ppm | 126.2468 | 4.8183 | 116.6102 | 135.8833 | 111.7919 | 140.7016 | 3.817% | 7.633% | 11.45% | 119.9344 | 132.5591
Re, ppm | 0.022087 | 0.002436 | 0.017215 | 0.026959 | 0.014779 | 0.029395 | 11.03% | 22.06% | 33.09% | 0.020982 | 0.023191
S, wt.% | 0.71732 | 0.049082 | 0.61916 | 0.81548 | 0.57007 | 0.86457 | 6.842% | 13.68% | 20.53% | 0.68145 | 0.75319
Sb, ppm | 15.2068 | 0.83196 | 13.5428 | 16.8707 | 12.7109 | 17.7027 | 5.471% | 10.94% | 16.41% | 14.4464 | 15.9671
Sc, ppm | 3.9361 | 0.24641 | 3.4433 | 4.4289 | 3.1969 | 4.6753 | 6.26% | 12.52% | 18.78% | 3.7393 | 4.1329
Se, ppm | 2.3538 | 0.47776 | 1.3983 | 3.3093 | 0.92052 | 3.7871 | 20.3% | 40.59% | 60.89% | 2.2361 | 2.4715
Sm, ppm | 3.3267 | 0.11147 | 3.1037 | 3.5496 | 2.9922 | 3.6611 | 3.351% | 6.702% | 10.05% | 3.1603 | 3.493
Sn, ppm | 1.4607 | 0.11236 | 1.236 | 1.6854 | 1.1236 | 1.7978 | 7.692% | 15.38% | 23.08% | 1.3877 | 1.5337
Sr, ppm | 277.4968 | 13.2097 | 251.0774 | 303.9162 | 237.8677 | 317.1259 | 4.76% | 9.521% | 14.28% | 263.622 | 291.3717
Ta, ppm | 0.39859 | 0.02988 | 0.33883 | 0.45835 | 0.30895 | 0.48823 | 7.496% | 14.99% | 22.49% | 0.37866 | 0.41852
Tb, ppm | 0.4445 | 0.029883 | 0.38473 | 0.50427 | 0.35485 | 0.53415 | 6.723% | 13.45% | 20.17% | 0.42228 | 0.46673
Te, ppm | 0.2347 | 0.032021 | 0.17066 | 0.29874 | 0.13864 | 0.33076 | 13.64% | 27.29% | 40.93% | 0.22296 | 0.24643
Th, ppm | 5.3621 | 0.21995 | 4.9222 | 5.802 | 4.7022 | 6.0219 | 4.102% | 8.204% | 12.31% | 5.094 | 5.6302
Ti, wt.% | 0.16051 | 0.0071219 | 0.14627 | 0.17476 | 0.13915 | 0.18188 | 4.437% | 8.874% | 13.31% | 0.15249 | 0.16854
Tl, ppm | 28.7725 | 1.0586 | 26.6553 | 30.8897 | 25.5968 | 31.9482 | 3.679% | 7.358% | 11.04% | 27.3339 | 30.2111
Tm, ppm | 0.19773 | 0.0059709 | 0.18579 | 0.20968 | 0.17982 | 0.21565 | 3.02% | 6.039% | 9.059% | 0.18785 | 0.20762
U, ppm | 7.3108 | 0.35603 | 6.5987 | 8.0229 | 6.2427 | 8.3789 | 4.87% | 9.74% | 14.61% | 6.9453 | 7.6763
V, ppm | 105.3 | 5.3348 | 94.6304 | 115.9696 | 89.2956 | 121.3044 | 5.066% | 10.13% | 15.2% | 100.035 | 110.565
W, ppm | 25.9244 | 1.6163 | 22.6918 | 29.157 | 21.0755 | 30.7733 | 6.235% | 12.47% | 18.7% | 24.6282 | 27.2206
Y, ppm | 16.5977 | 0.79705 | 15.0036 | 18.1918 | 14.2065 | 18.9888 | 4.802% | 9.604% | 14.41% | 15.7678 | 17.4276
Yb, ppm | 1.3057 | 0.065345 | 1.175 | 1.4364 | 1.1097 | 1.5017 | 5.005% | 10.01% | 15.01% | 1.2404 | 1.371
Zn, ppm | 120.6753 | 8.3165 | 104.0424 | 137.3083 | 95.7259 | 145.6247 | 6.892% | 13.78% | 20.67% | 114.6416 | 126.7091
Zr, ppm | 57.8114 | 7.3734 | 43.0645 | 72.5583 | 35.6911 | 79.9317 | 12.75% | 25.51% | 38.26% | 54.9208 | 60.702
Aqua Regia Digestion
Ag, ppm | 0.57659 | 0.03176 | 0.51307 | 0.64011 | 0.48131 | 0.67187 | 5.508% | 11.02% | 16.52% | 0.54776 | 0.60542
Al, wt.% | 0.81479 | 0.073519 | 0.66775 | 0.96183 | 0.59423 | 1.0353 | 9.023% | 18.05% | 27.07% | 0.77405 | 0.85553
As, ppm | 2319.5166 | 130.0038 | 2059.509 | 2579.5241 | 1929.5053 | 2709.5279 | 5.605% | 11.21% | 16.81% | 2203.5407 | 2435.4924
Ba, ppm | 29.7765 | 2.3771 | 25.0223 | 34.5306 | 22.6453 | 36.9077 | 7.983% | 15.97% | 23.95% | 28.2877 | 31.2653
Be, ppm | 0.66693 | 0.043879 | 0.57917 | 0.75468 | 0.53529 | 0.79856 | 6.579% | 13.16% | 19.74% | 0.63358 | 0.70027
Bi, ppm | 0.90693 | 0.02628 | 0.85437 | 0.95949 | 0.82809 | 0.98577 | 2.898% | 5.795% | 8.693% | 0.86158 | 0.95227
Ca, wt.% | 16.3536 | 1.1262 | 14.1012 | 18.606 | 12.9751 | 19.7322 | 6.886% | 13.77% | 20.66% | 15.5359 | 17.1713
Cd, ppm | 1.0933 | 0.067347 | 0.95858 | 1.228 | 0.89123 | 1.2953 | 6.16% | 12.32% | 18.48% | 1.0386 | 1.1479
Ce, ppm | 33.8419 | 2.0756 | 29.6907 | 37.9932 | 27.6151 | 40.0688 | 6.133% | 12.27% | 18.4% | 32.1498 | 35.534
Co, ppm | 3.6111 | 0.38217 | 2.8468 | 4.3755 | 2.4646 | 4.7576 | 10.58% | 21.17% | 31.75% | 3.4306 | 3.7917
Cr, ppm | 23.7639 | 1.728 | 20.308 | 27.2199 | 18.5801 | 28.9478 | 7.271% | 14.54% | 21.81% | 22.5757 | 24.9521
Cs, ppm | 4.229 | 0.72564 | 2.7778 | 5.6803 | 2.0521 | 6.406 | 17.16% | 34.32% | 51.48% | 4.0176 | 4.4405
Cu, ppm | 18.5941 | 1.4612 | 15.6717 | 21.5166 | 14.2105 | 22.9778 | 7.859% | 15.72% | 23.58% | 17.6644 | 19.5238
Dy, ppm | 2.433 | 0.20256 | 2.0279 | 2.8381 | 1.8253 | 3.0407 | 8.325% | 16.65% | 24.98% | 2.3114 | 2.5546
Er, ppm | 1.2584 | 0.071638 | 1.1151 | 1.4016 | 1.0434 | 1.4733 | 5.693% | 11.39% | 17.08% | 1.1954 | 1.3213
Eu, ppm | 0.5985 | 0.053682 | 0.49114 | 0.70586 | 0.43745 | 0.75955 | 8.969% | 17.94% | 26.91% | 0.56858 | 0.62843
Fe, wt.% | 0.99785 | 0.041627 | 0.91459 | 1.0811 | 0.87296 | 1.1227 | 4.172% | 8.343% | 12.52% | 0.94795 | 1.0477
Ga, ppm | 2.9864 | 0.37193 | 2.2425 | 3.7302 | 1.8706 | 4.1022 | 12.45% | 24.91% | 37.36% | 2.8371 | 3.1357
Gd, ppm | 2.9224 | 0.22613 | 2.4701 | 3.3746 | 2.244 | 3.6007 | 7.738% | 15.48% | 23.21% | 2.7762 | 3.0685
Ge, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Hf, ppm | 0.35474 | 0.028997 | 0.29675 | 0.41274 | 0.26775 | 0.44174 | 8.174% | 16.35% | 24.52% | 0.33701 | 0.37248
Hg, ppm | 40.3023 | 2.0151 | 36.2721 | 44.3325 | 34.257 | 46.3476 | 5% | 10% | 15% | 38.2872 | 42.3174
Ho, ppm | 0.44986 | 0.024109 | 0.40164 | 0.49807 | 0.37753 | 0.52218 | 5.359% | 10.72% | 16.08% | 0.42736 | 0.47235
In, ppm | 0.055625 | 0.0056541 | 0.044317 | 0.066933 | 0.038663 | 0.072587 | 10.16% | 20.33% | 30.49% | 0.052844 | 0.058406
K, wt.% | 0.4121 | 0.024756 | 0.36258 | 0.46161 | 0.33783 | 0.48636 | 6.007% | 12.01% | 18.02% | 0.39149 | 0.4327
La, ppm | 18.2275 | 1.2522 | 15.723 | 20.732 | 14.4708 | 21.9843 | 6.87% | 13.74% | 20.61% | 17.3161 | 19.1389
Li, ppm | 12.2831 | 0.60643 | 11.0702 | 13.4959 | 10.4638 | 14.1023 | 4.937% | 9.874% | 14.81% | 11.6689 | 12.8972
Lu, ppm | 0.15352 | 0.011178 | 0.13117 | 0.17588 | 0.11999 | 0.18706 | 7.281% | 14.56% | 21.84% | 0.14585 | 0.1612
Mg, wt.% | 2.8034 | 0.069753 | 2.6639 | 2.9429 | 2.5942 | 3.0127 | 2.488% | 4.976% | 7.464% | 2.6632 | 2.9436
Mn, wt.% | 0.023668 | 0.001646 | 0.020377 | 0.02696 | 0.018731 | 0.028606 | 6.954% | 13.91% | 20.86% | 0.022485 | 0.024852
Mo, ppm | 13.0859 | 0.82998 | 11.4259 | 14.7458 | 10.5959 | 15.5758 | 6.343% | 12.69% | 19.03% | 12.4316 | 13.7402
Nb, ppm | 0.17204 | 0.051568 | 0.068905 | 0.27518 | 0.017337 | 0.32675 | 29.97% | 59.95% | 89.92% | 0.16344 | 0.18064
Nd, ppm | 16.245 | 1.614 | 13.0169 | 19.473 | 11.4029 | 21.087 | 9.935% | 19.87% | 29.81% | 15.4327 | 17.0572
Ni, ppm | 21.0368 | 1.8984 | 17.2399 | 24.8336 | 15.3415 | 26.7321 | 9.024% | 18.05% | 27.07% | 19.9849 | 22.0886
P, wt.% | 0.030541 | 0.0026292 | 0.025283 | 0.0358 | 0.022654 | 0.038429 | 8.609% | 17.22% | 25.83% | 0.029014 | 0.032068
Pb, ppm | 17.4767 | 1.0974 | 15.2819 | 19.6714 | 14.1846 | 20.7688 | 6.279% | 12.56% | 18.84% | 16.6028 | 18.3505
Pr, ppm | 4.2067 | 0.44854 | 3.3097 | 5.1038 | 2.8611 | 5.5524 | 10.66% | 21.32% | 31.99% | 3.9964 | 4.4171
Rb, ppm | 38.9758 | 3.8529 | 31.27 | 46.6817 | 27.4171 | 50.5346 | 9.885% | 19.77% | 29.66% | 37.0271 | 40.9246
Re, ppm | 0.023547 | 0.0022845 | 0.018978 | 0.028116 | 0.016694 | 0.030401 | 9.702% | 19.4% | 29.11% | 0.02237 | 0.024725
S, wt.% | 0.70784 | 0.05588 | 0.59608 | 0.8196 | 0.5402 | 0.87549 | 7.894% | 15.79% | 23.68% | 0.67245 | 0.74324
Sb, ppm | 10.4941 | 1.1793 | 8.1354 | 12.8527 | 6.9561 | 14.032 | 11.24% | 22.48% | 33.71% | 9.9694 | 11.0188
Sc, ppm | 3.523 | 0.32357 | 2.8759 | 4.1702 | 2.5523 | 4.4938 | 9.184% | 18.37% | 27.55% | 3.3469 | 3.6992
Se, ppm | 2.3454 | 0.40292 | 1.5395 | 3.1512 | 1.1366 | 3.5541 | 17.18% | 34.36% | 51.54% | 2.2281 | 2.4626
Sm, ppm | 3.1308 | 0.28369 | 2.5634 | 3.6982 | 2.2797 | 3.9819 | 9.061% | 18.12% | 27.18% | 2.9743 | 3.2873
Sn, ppm | 1.0244 | 0.093503 | 0.83735 | 1.2114 | 0.74385 | 1.3049 | 9.128% | 18.26% | 27.38% | 0.97314 | 1.0756
Sr, ppm | 217.5887 | 17.583 | 182.4228 | 252.7547 | 164.8398 | 270.3376 | 8.081% | 16.16% | 24.24% | 206.7093 | 228.4681
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.41656 | 0.031541 | 0.35347 | 0.47964 | 0.32193 | 0.51118 | 7.572% | 15.14% | 22.72% | 0.39573 | 0.43738
Te, ppm | 0.24286 | 0.029277 | 0.18431 | 0.30142 | 0.15503 | 0.33069 | 12.05% | 24.11% | 36.16% | 0.23072 | 0.25501
Th, ppm | 4.7125 | 0.23166 | 4.2492 | 5.1758 | 4.0175 | 5.4075 | 4.916% | 9.832% | 14.75% | 4.4769 | 4.9481
Ti, wt.% | 0.019677 | 0.0022118 | 0.015254 | 0.024101 | 0.013042 | 0.026313 | 11.24% | 22.48% | 33.72% | 0.018694 | 0.020661
Tl, ppm | 25.337 | 1.4632 | 22.4106 | 28.2634 | 20.9474 | 29.7266 | 5.775% | 11.55% | 17.32% | 24.0701 | 26.6039
Tm, ppm | 0.17192 | 0.013643 | 0.14463 | 0.1992 | 0.13099 | 0.21285 | 7.936% | 15.87% | 23.81% | 0.16332 | 0.18051
U, ppm | 6.1674 | 0.33461 | 5.4982 | 6.8366 | 5.1636 | 7.1712 | 5.425% | 10.85% | 16.28% | 5.859 | 6.4758
V, ppm | 60.3558 | 5.1288 | 50.0983 | 70.6134 | 44.9695 | 75.7422 | 8.498% | 17% | 25.49% | 57.338 | 63.3736
W, ppm | 4.4622 | 0.5535 | 3.3552 | 5.5692 | 2.8017 | 6.1227 | 12.4% | 24.81% | 37.21% | 4.2391 | 4.6853
Y, ppm | 14.9715 | 0.70237 | 13.5667 | 16.3762 | 12.8644 | 17.0786 | 4.691% | 9.383% | 14.07% | 14.2229 | 15.7201
Yb, ppm | 1.1081 | 0.05489 | 0.9983 | 1.2179 | 0.94341 | 1.2728 | 4.954% | 9.907% | 14.86% | 1.0527 | 1.1635
Zn, ppm | 112.0736 | 6.568 | 98.9376 | 125.2097 | 92.3696 | 131.7777 | 5.86% | 11.72% | 17.58% | 106.4699 | 117.6773
Zr, ppm | 12.0236 | 1.6111 | 8.8014 | 15.2457 | 7.1903 | 16.8568 | 13.4% | 26.8% | 40.2% | 11.4224 | 12.6247
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 282
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 13.7087 | 13.5986 | 13.8188 | 13.6796 | 13.7377
PhotonAssay
Au, Gold (ppm) | 14.0508 | 13.9334 | 14.1682 | 14.0421 | 14.0595
4-Acid Digestion
Ag, Silver (ppm) | 0.61106 | 0.5819 | 0.64022 | 0.58451 | 0.63761
Al, Aluminium (wt.%) | 2.9473 | 2.8725 | 3.0222 | 2.895 | 2.9997
As, Arsenic (ppm) | 2327.7175 | 2260.4199 | 2395.0152 | 2285.1633 | 2370.2718
Ba, Barium (ppm) | 99.7694 | 97.4215 | 102.1174 | 97.6893 | 101.8496
Be, Beryllium (ppm) | 0.92801 | 0.85642 | 0.9996 | 0.89474 | 0.96128
Bi, Bismuth (ppm) | 0.93214 | 0.89925 | 0.96503 | 0.90738 | 0.95689
Ca, Calcium (wt.%) | 17.4672 | 16.923 | 18.0114 | 17.0877 | 17.8467
Cd, Cadmium (ppm) | 1.0798 | 1.012 | 1.1477 | 1.0244 | 1.1353
Ce, Cerium (ppm) | 36.8927 | 35.4715 | 38.314 | 36.1138 | 37.6717
Co, Cobalt (ppm) | 3.7994 | 3.6127 | 3.9862 | 3.6688 | 3.9301
Cr, Chromium (ppm) | 32.5159 | 29.0608 | 35.971 | 30.4692 | 34.5625
Cs, Caesium (ppm) | 8.2399 | 7.9952 | 8.4846 | 8.0585 | 8.4212
Cu, Copper (ppm) | 19.1511 | 18.0422 | 20.2601 | 18.5612 | 19.7411
Dy, Dysprosium (ppm) | 2.6079 | 2.4665 | 2.7492 | 2.543 | 2.6727
Er, Erbium (ppm) | 1.461 | 1.3897 | 1.5322 | 1.4033 | 1.5186
Eu, Europium (ppm) | 0.65426 | 0.59918 | 0.70934 | 0.63123 | 0.67729
Fe, Iron (wt.%) | 1.0641 | 1.0391 | 1.0892 | 1.043 | 1.0853
Ga, Gallium (ppm) | 7.1461 | 6.8537 | 7.4385 | 6.9265 | 7.3657
Gd, Gadolinium (ppm) | 3.0464 | 2.8223 | 3.2705 | 2.9382 | 3.1545
Hf, Hafnium (ppm) | 1.6265 | 1.4741 | 1.7789 | 1.5159 | 1.737
Ho, Holmium (ppm) | 0.50778 | 0.47244 | 0.54312 | 0.48804 | 0.52751
In, Indium (ppm) | 0.059486 | 0.053117 | 0.065854 | 0.053096 | 0.065876
K, Potassium (wt.%) | 1.8512 | 1.7929 | 1.9096 | 1.8101 | 1.8923
La, Lanthanum (ppm) | 20.5254 | 19.7136 | 21.3372 | 20.0088 | 21.0421
Li, Lithium (ppm) | 21.5676 | 20.6855 | 22.4497 | 20.8972 | 22.2381
Lu, Lutetium (ppm) | 0.18917 | 0.16898 | 0.20935 | IND | IND
Mg, Magnesium (wt.%) | 3.3938 | 3.2974 | 3.4903 | 3.3332 | 3.4545
Mn, Manganese (wt.%) | 0.025112 | 0.024541 | 0.025682 | 0.024612 | 0.025611
Mo, Molybdenum (ppm) | 13.5416 | 13.0478 | 14.0354 | 13.2298 | 13.8534
Na, Sodium (wt.%) | 0.065404 | 0.062394 | 0.068415 | 0.062069 | 0.068739
Nb, Niobium (ppm) | 5.7269 | 5.481 | 5.9728 | 5.575 | 5.8788
Nd, Neodymium (ppm) | 17.4517 | 16.8389 | 18.0644 | 17.0289 | 17.8745
Ni, Nickel (ppm) | 21.8135 | 20.9625 | 22.6645 | 21.2628 | 22.3642
P, Phosphorus (wt.%) | 0.031805 | 0.030374 | 0.033236 | 0.030881 | 0.03273
Pb, Lead (ppm) | 18.7597 | 18.0258 | 19.4936 | 18.259 | 19.2604
Pr, Praseodymium (ppm) | 4.6041 | 4.4539 | 4.7544 | 4.4919 | 4.7164
Rb, Rubidium (ppm) | 126.2468 | 122.8084 | 129.6851 | 123.6223 | 128.8712
Re, Rhenium (ppm) | 0.022087 | 0.018163 | 0.02601 | 0.019629 | 0.024545
S, Sulphur (wt.%) | 0.71732 | 0.68632 | 0.74832 | 0.7023 | 0.73234
Sb, Antimony (ppm) | 15.2068 | 14.6159 | 15.7977 | 14.7814 | 15.6321
Sc, Scandium (ppm) | 3.9361 | 3.7285 | 4.1436 | 3.7907 | 4.0814
Se, Selenium (ppm) | 2.3538 | 1.7038 | 3.0038 | IND | IND
Sm, Samarium (ppm) | 3.3267 | 3.1933 | 3.46 | 3.1615 | 3.4919
Sn, Tin (ppm) | 1.4607 | 1.3618 | 1.5596 | IND | IND
Sr, Strontium (ppm) | 277.4968 | 268.6903 | 286.3033 | 271.9931 | 283.0006
Ta, Tantalum (ppm) | 0.39859 | 0.37058 | 0.4266 | 0.37736 | 0.41982
Tb, Terbium (ppm) | 0.4445 | 0.41081 | 0.47819 | 0.42492 | 0.46408
Te, Tellurium (ppm) | 0.2347 | 0.19115 | 0.27825 | 0.19883 | 0.27057
Th, Thorium (ppm) | 5.3621 | 5.172 | 5.5522 | 5.2307 | 5.4934
Ti, Titanium (wt.%) | 0.16051 | 0.15594 | 0.16508 | 0.15751 | 0.16351
Tl, Thallium (ppm) | 28.7725 | 27.991 | 29.554 | 28.0292 | 29.5158
Tm, Thulium (ppm) | 0.19773 | 0.18813 | 0.20734 | IND | IND
U, Uranium (ppm) | 7.3108 | 7.0578 | 7.5638 | 7.1588 | 7.4628
V, Vanadium (ppm) | 105.3 | 101.4974 | 109.1026 | 103.0742 | 107.5258
W, Tungsten (ppm) | 25.9244 | 24.8824 | 26.9664 | 25.204 | 26.6448
Y, Yttrium (ppm) | 16.5977 | 16.0388 | 17.1565 | 16.1866 | 17.0088
Yb, Ytterbium (ppm) | 1.3057 | 1.2165 | 1.3949 | 1.2487 | 1.3627
Zn, Zinc (ppm) | 120.6753 | 115.4854 | 125.8653 | 118.033 | 123.3177
Zr, Zirconium (ppm) | 57.8114 | 53.2276 | 62.3952 | 54.9399 | 60.6829
Aqua Regia Digestion
Ag, Silver (ppm) | 0.57659 | 0.54077 | 0.61241 | 0.55793 | 0.59525
Al, Aluminium (wt.%) | 0.81479 | 0.7752 | 0.85438 | 0.79763 | 0.83195
As, Arsenic (ppm) | 2319.5166 | 2246.3507 | 2392.6825 | 2269.299 | 2369.7341
Ba, Barium (ppm) | 29.7765 | 28.2431 | 31.3099 | 28.6478 | 30.9051
Be, Beryllium (ppm) | 0.66693 | 0.60747 | 0.72639 | 0.63798 | 0.69587
Bi, Bismuth (ppm) | 0.90693 | 0.88227 | 0.93158 | 0.87973 | 0.93412
Ca, Calcium (wt.%) | 16.3536 | 15.6302 | 17.077 | 16.0062 | 16.701
Cd, Cadmium (ppm) | 1.0933 | 1.0367 | 1.1498 | 1.042 | 1.1446
Ce, Cerium (ppm) | 33.8419 | 32.332 | 35.3518 | 32.9538 | 34.7301
Co, Cobalt (ppm) | 3.6111 | 3.3658 | 3.8565 | 3.4624 | 3.7598
Cr, Chromium (ppm) | 23.7639 | 22.6546 | 24.8733 | 22.9089 | 24.619
Cs, Caesium (ppm) | 4.229 | 3.824 | 4.6341 | 4.1036 | 4.3545
Cu, Copper (ppm) | 18.5941 | 17.5921 | 19.5961 | 17.9836 | 19.2046
Dy, Dysprosium (ppm) | 2.433 | 2.219 | 2.647 | 2.3148 | 2.5512
Er, Erbium (ppm) | 1.2584 | 1.1769 | 1.3398 | 1.2137 | 1.303
Eu, Europium (ppm) | 0.5985 | 0.52938 | 0.66762 | 0.58138 | 0.61562
Fe, Iron (wt.%) | 0.99785 | 0.97226 | 1.0234 | 0.97745 | 1.0182
Ga, Gallium (ppm) | 2.9864 | 2.7195 | 3.2532 | 2.8467 | 3.126
Gd, Gadolinium (ppm) | 2.9224 | 2.7115 | 3.1332 | 2.7947 | 3.05
Ge, Germanium (ppm) | < 0.1 | IND | IND | IND | IND
Hf, Hafnium (ppm) | 0.35474 | 0.32758 | 0.3819 | 0.33236 | 0.37713
Hg, Mercury (ppm) | 40.3023 | 38.9832 | 41.6214 | 39.3898 | 41.2148
Ho, Holmium (ppm) | 0.44986 | 0.4207 | 0.47902 | 0.43305 | 0.46667
In, Indium (ppm) | 0.055625 | 0.050098 | 0.061152 | 0.050176 | 0.061074
K, Potassium (wt.%) | 0.4121 | 0.39855 | 0.42564 | 0.39999 | 0.42421
La, Lanthanum (ppm) | 18.2275 | 17.3173 | 19.1377 | 17.7478 | 18.7072
Li, Lithium (ppm) | 12.2831 | 11.6288 | 12.9373 | 11.873 | 12.6931
Lu, Lutetium (ppm) | 0.15352 | 0.13915 | 0.16789 | 0.14489 | 0.16215
Mg, Magnesium (wt.%) | 2.8034 | 2.7422 | 2.8647 | 2.7484 | 2.8584
Mn, Manganese (wt.%) | 0.023668 | 0.022799 | 0.024538 | 0.023095 | 0.024242
Mo, Molybdenum (ppm) | 13.0859 | 12.6117 | 13.5601 | 12.7725 | 13.3993
Nb, Niobium (ppm) | 0.17204 | 0.1323 | 0.21179 | 0.13851 | 0.20558
Nd, Neodymium (ppm) | 16.245 | 14.6395 | 17.8505 | 15.6361 | 16.8538
Ni, Nickel (ppm) | 21.0368 | 19.7788 | 22.2948 | 20.4045 | 21.669
P, Phosphorus (wt.%) | 0.030541 | 0.028859 | 0.032224 | 0.029462 | 0.031621
Pb, Lead (ppm) | 17.4767 | 16.5845 | 18.3688 | 16.9421 | 18.0112
Pr, Praseodymium (ppm) | 4.2067 | 3.7902 | 4.6233 | 4.0549 | 4.3586
Rb, Rubidium (ppm) | 38.9758 | 36.459 | 41.4927 | 37.8598 | 40.0919
Re, Rhenium (ppm) | 0.023547 | 0.021026 | 0.026069 | 0.021412 | 0.025682
S, Sulphur (wt.%) | 0.70784 | 0.67621 | 0.73948 | 0.69155 | 0.72414
Sb, Antimony (ppm) | 10.4941 | 9.7743 | 11.2138 | 10.1887 | 10.7994
Sc, Scandium (ppm) | 3.523 | 3.3054 | 3.7406 | 3.4001 | 3.646
Se, Selenium (ppm) | 2.3454 | 1.9708 | 2.7199 | 2.1317 | 2.559
Sm, Samarium (ppm) | 3.1308 | 2.8622 | 3.3994 | 3.0135 | 3.2481
Sn, Tin (ppm) | 1.0244 | 0.93656 | 1.1122 | 0.96537 | 1.0833
Sr, Strontium (ppm) | 217.5887 | 208.4273 | 226.7501 | 212.991 | 222.1864
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.41656 | 0.39163 | 0.44148 | 0.40614 | 0.42697
Te, Tellurium (ppm) | 0.24286 | 0.21656 | 0.26916 | 0.22451 | 0.26122
Th, Thorium (ppm) | 4.7125 | 4.5291 | 4.8959 | 4.5285 | 4.8964
Ti, Titanium (wt.%) | 0.019677 | 0.01826 | 0.021095 | 0.019031 | 0.020324
Tl, Thallium (ppm) | 25.337 | 24.4008 | 26.2732 | 24.6858 | 25.9882
Tm, Thulium (ppm) | 0.17192 | 0.15759 | 0.18625 | 0.16111 | 0.18272
U, Uranium (ppm) | 6.1674 | 5.9078 | 6.427 | 6.0218 | 6.313
V, Vanadium (ppm) | 60.3558 | 57.4175 | 63.2942 | 58.389 | 62.3227
W, Tungsten (ppm) | 4.4622 | 4.1397 | 4.7847 | 4.3236 | 4.6008
Y, Yttrium (ppm) | 14.9715 | 14.4667 | 15.4763 | 14.5855 | 15.3575
Yb, Ytterbium (ppm) | 1.1081 | 1.0425 | 1.1737 | 1.0637 | 1.1524
Zn, Zinc (ppm) | 112.0736 | 108.3759 | 115.7713 | 109.3045 | 114.8428
Zr, Zirconium (ppm) | 12.0236 | 10.9587 | 13.0884 | 11.5507 | 12.4964
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 282
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | < 5 | Pt | ppb | < 5
4-Acid Digestion
B | ppm | 653.8333 | Ge | ppm | 0.19029 | Hg | ppm | 21.6667
Aqua Regia Digestion (sample weights 10-50g)
Au | ppm | 10.8661
Aqua Regia Digestion
B | ppm | 10 | Pd | ppb | < 10 | Si | wt.% | 0.014333
Na | wt.% | 0.020126 | Pt | ppb | < 5
Borate Fusion XRF
Al2O3 | wt.% | 5.58 | MgO | wt.% | 5.825 | SiO2 | wt.% | 32.955
CaO | wt.% | 25.645 | MnO | wt.% | 0.03 | SO3 | wt.% | 1.679
Fe2O3 | wt.% | 1.535 | Na2O | wt.% | 0.07 | TiO2 | wt.% | 0.27
K2O | wt.% | 2.175 | P2O5 | wt.% | 0.0735
Thermogravimetry
LOI1000 | wt.% | 24.585
Infrared Combustion
C | wt.% | 6.755 | S | wt.% | 0.32
Laser Ablation ICP-MS
Ag | ppm | 0.95 | Hf | ppm | 4.22 | Sm | ppm | 3.545
As | ppm | 2760 | Ho | ppm | 0.565 | Sn | ppm | 1.8
Ba | ppm | 105 | In | ppm | 0.0625 | Sr | ppm | 283
Be | ppm | 1 | La | ppm | 20.95 | Ta | ppm | 0.435
Bi | ppm | 1.2 | Lu | ppm | 0.205 | Tb | ppm | 0.46
Cd | ppm | 1.45 | Mn | wt.% | 0.02735 | Te | ppm | < 0.2
Ce | ppm | 37.3 | Mo | ppm | 13.6 | Th | ppm | 5.625
Co | ppm | 4.9 | Nb | ppm | 5.9 | Ti | wt.% | 0.163
Cr | ppm | 46 | Nd | ppm | 18.8 | Tl | ppm | 25.1
Cs | ppm | 8.98 | Ni | ppm | 34 | Tm | ppm | 0.21
Cu | ppm | 30 | Pb | ppm | 24 | U | ppm | 7.875
Dy | ppm | 2.69 | Pr | ppm | 4.885 | V | ppm | 116
Er | ppm | 1.535 | Rb | ppm | 135.5 | W | ppm | 27
Eu | ppm | 0.7 | Re | ppm | 0.035 | Y | ppm | 17.55
Ga | ppm | 8.75 | Sb | ppm | 17.7 | Yb | ppm | 1.51
Gd | ppm | 3.02 | Sc | ppm | 4.15 | Zn | ppm | 140
Ge | ppm | 0.55 | Se | ppm | < 5 | Zr | ppm | 153.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 282
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*GRAV | fire assay with gravimetric finish
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | Raw ~350g sample packed into 300cc jar analysed by PhotonAssay (high-power X-ray activation)
Text Values:
I/S | Insufficient Sample (Lab 2.10)
N.A. | Not Applicable (Lab 2.08)
N.A. | Not Applicable (Lab 2.27)
NA | Not Applicable (Lab 2.21)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 282
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Canning Vale, WA, Australia
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Reno, Nevada, USA
ALS, Vancouver, BC, Canada
ANSTO, Lucas Heights, NSW, Australia
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities and Trade, Inc., Sparks, Nevada, USA
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MinAnalytical Services, Perth, WA, Australia
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Val-d'Or, Quebec, Canada
MSALABS, Vancouver, BC, Canada
Nagrom, Perth, WA, Australia
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Kalgoorlie, WA, Australia
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Tarkwa, Tarkwa, Western Region, Ghana
Skyline Assayers & Laboratories, Tucson, Arizona, USA

## Upscaled Metrics
ORE - Lab-Upscaled RSD Results for CRM: OREAS 282 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response | Upscaled Value
1 | 1 | ANSLu | 1 | 10 | 10 | 202446 | 0.081805 | 13.9041 | 14.0022
1 | 1 | ANSLu | 1 | 1 | 1 | 202447 | 0.082408 | 13.7119 | 13.992
1 | 1 | ANSLu | 1 | 15 | 15 | 202448 | 0.082858 | 13.7712 | 13.9952
1 | 1 | ANSLu | 1 | 11 | 11 | 202449 | 0.086988 | 13.6999 | 13.9914
1 | 1 | ANSLu | 1 | 16 | 16 | 202450 | 0.083172 | 13.7865 | 13.996
1 | 1 | ANSLu | 1 | 12 | 12 | 202451 | 0.085564 | 13.9543 | 14.0049
1 | 1 | ANSLu | 1 | 6 | 6 | 202452 | 0.086566 | 13.8885 | 14.0014
1 | 1 | ANSLu | 1 | 14 | 14 | 202453 | 0.088083 | 14.002 | 14.0075
1 | 1 | ANSLu | 1 | 2 | 2 | 202454 | 0.087064 | 14.1603 | 14.0159
1 | 1 | ANSLu | 1 | 4 | 4 | 202455 | 0.084169 | 14.2924 | 14.0229
1 | 1 | ANSLu | 1 | 19 | 19 | 202456 | 0.085731 | 14.0851 | 14.0119
1 | 1 | ANSLu | 1 | 8 | 8 | 202457 | 0.085282 | 14.021 | 14.0085
1 | 1 | ANSLu | 1 | 7 | 7 | 202458 | 0.084687 | 14.0546 | 14.0103
1 | 1 | ANSLu | 1 | 18 | 18 | 202459 | 0.083875 | 14.0894 | 14.0121
1 | 1 | ANSLu | 1 | 3 | 3 | 202460 | 0.084051 | 14.2625 | 14.0213
1 | 1 | ANSLu | 1 | 5 | 5 | 202461 | 0.088259 | 13.9612 | 14.0053
1 | 1 | ANSLu | 1 | 13 | 13 | 202462 | 0.08218 | 14.0465 | 14.0098
1 | 1 | ANSLu | 1 | 20 | 20 | 202463 | 0.083714 | 14.0355 | 14.0092
1 | 1 | ANSLu | 1 | 9 | 9 | 202464 | 0.082634 | 14.1988 | 14.0179
1 | 1 | ANSLu | 1 | 17 | 17 | 202465 | 0.087698 | 14.2296 | 14.0196
Mean | [6 empty] | 0.084839 | 14.0078 | 14.0078
Median | [7 empty] | 14.0283 | 14.0089
SD | [7 empty] | 0.17613 | 0.0093664
RSD | [7 empty] | 1.257% | 0.06687%
RSD Upscale Factor | [7 empty] | 0.053179
Parameters | :
Nominal Mass Value |  | 30 | gm
Note: Upscaled Metrics are also available in this CSG for analytes with INAA results for Au

## Fire Assay
| Analytical results for Au in OREAS 282 (Certified Value 13.71 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | INAA | FA*GRAV | FA*OES | FA*AAS | FA*GRAV | FA*OES | FA*GRAV | FA*OES | FA*GRAV | FA*GRAV | FA*GRAV | FA*AAS | FA*AAS | FA*GRAV | FA*MS | FA*OES | FA*GRAV | FA*AAS | FA*GRAV | FA*OES | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*AAS | FA*GRAV | FA*AAS | FA*GRAV | FA*GRAV | FA*GRAV | 2
|  |  | 0.085g | 30g | 40g | 40g | 30g | 15g | 30g | 30g | 50g | 30g | 25g | 50g | 25g | 30g | 25g | 25g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 13.7119 | 13.4 | 13.7 | 13.61 | 13.88 | 12.8 | 13.5 | 13.811 | 13.22 | 13 | 14.27 | 13.1575 | 13.37 | 13.68 | 16.77 | 14.094 | 13.5 | 11 | 13.6 |  | 13.7 | 13 | 13.5 | 13.8 | 13.8 | 14.1 | 13.5 | 14.1 | 13.75 | 14.4 | 14 | 13.65 | 1
| 1 | 2 | 14.1603 | 13.75 | 13.5 | 13.53 | 14.15 | 12.9 | 13.5 | 13.851 | 13.76 | 12.8 | I/S | 13.5019 | 12.91 | 13.65 | 16.87 | 14.174 | 13.8 | 12.6 | 13.8 |  | 13.75 | 14 | 13.6 | 13.7 | 13.8 | 13.8 | 13.4 | 14 | 13.97 | 14.7 | 13.9 | 13.7 | 5
| 1 | 3 | 14.2625 | 13.5 | 13.5 | 13.69 | 14.07 | 12.8 | 12.9 | 13.251 | 13.45 | 12.5 | 13.84 | 13.6267 | 13.26 | 13.73 | 16.84 | 14.024 | 13.5 | 12.9 | 13.7 |  | 13.65 | 14 | 13.65 | 13.3 | 13.8 | 14.1 | 13.6 | 14 | 14.29 | 14.2 | 13.8 | 13.8 | 16
| 1 | 4 | 14.2924 | 13.3 | 13.5 | 13.68 | 13.97 | 13 | 13.4 | 13.254 | 13.21 | 13 | 13.84 | 13.5738 | 13.05 | 13.54 | 17.21 | 14.122 | 13.6 | 12.3 | 14 |  | 13.9 | 14 | 13.65 | 13.8 | 13.8 | 13.9 | 13.8 | 14 | 13.97 | 14.1 | 13.9 | 13.95 | 13.7087
| 1 | 5 | 13.9612 | 12.8 | 13.5 | 13.67 | 14.09 | 12.7 | 13 | 13.596 | 13.31 | 12.8 | 14.31 | 13.5978 | 12.95 | 13.52 | 16.63 | 14.088 | 13.3 | 12.8 | 13.5 |  | 13.7 | 14 | 13.8 | 13.8 | 13.7 | 14.1 | 13.6 | 13.8 | 14.23 | 13.7 | 13.9 | 13.95 | 7
| 1 | 6 | 13.8885 | 13.6 | 13.3 | 13.7 | 14.1 | 12.8 | 13.4 | 13.71 | 13.73 | 12.9 | 13.24 | 13.0979 | 14.66 | 13.55 | 17.46 | 13.807 | 13.4 | 12.7 | 14.2 |  | 13.35 | 14 | 13.55 | 13.8 | 13.6 | 14 | 13.7 | 13.9 | 14.32 | 15.1 | 13.7 | 13.25
|  | 7 | 14.0546
|  | 8 | 14.021
|  | 9 | 14.1988
|  | 10 | 13.9041
|  | 11 | 13.6999
|  | 12 | 13.9543
|  | 13 | 14.0465
|  | 14 | 14.002
|  | 15 | 13.7712
|  | 16 | 13.7865
|  | 17 | 14.2296
|  | 18 | 14.0894
|  | 19 | 14.0851
|  | 20 | 14.0355
| Mean |  | 14.0078 | 13.3917 | 13.5 | 13.6467 | 14.0433 | 12.8333 | 13.2833 | 13.5788 | 13.4467 | 12.8333 | 13.9 | 13.4259 | 13.3667 | 13.6117 | 16.9633 | 14.0515 | 13.5167 | 12.3833 | 13.8 |  | 13.675 | 13.8333 | 13.625 | 13.7 | 13.75 | 14 | 13.6 | 13.9667 | 14.0883 | 14.3667 | 13.8667 | 13.7167
| Median |  | 14.0283 | 13.45 | 13.5 | 13.675 | 14.08 | 12.8 | 13.4 | 13.653 | 13.38 | 12.85 | 13.84 | 13.5379 | 13.155 | 13.6 | 16.855 | 14.091 | 13.5 | 12.65 | 13.75 |  | 13.7 | 14 | 13.625 | 13.8 | 13.8 | 14.05 | 13.6 | 14 | 14.1 | 14.3 | 13.9 | 13.75
| Std Dev. |  | 0.17613 | 0.32927 | 0.12649 | 0.06532 | 0.099532 | 0.10328 | 0.26394 | 0.26776 | 0.24679 | 0.18619 | 0.43238 | 0.23545 | 0.65826 | 0.086583 | 0.30982 | 0.12935 | 0.17224 | 0.70828 | 0.26077 |  | 0.18097 | 0.40825 | 0.10368 | 0.2 | 0.083666 | 0.12649 | 0.14142 | 0.10328 | 0.22666 | 0.48854 | 0.10328 | 0.26013
| Rel.Std.Dev. |  | 1.257% | 2.459% | 0.937% | 0.4786% | 0.7088% | 0.8048% | 1.987% | 1.972% | 1.835% | 1.451% | 3.111% | 1.754% | 4.925% | 0.6361% | 1.826% | 0.9205% | 1.274% | 5.72% | 1.89% |  | 1.323% | 2.951% | 0.761% | 1.46% | 0.6085% | 0.9035% | 1.04% | 0.7395% | 1.609% | 3.4% | 0.7448% | 1.896%
| PDM3 |  | 2.182% | -2.313% | -1.522% | -0.4524% | 2.441% | -6.385% | -3.103% | -0.9473% | -1.911% | -6.385% | 1.396% | -2.063% | -2.495% | -0.7077% | 23.74% | 2.501% | -1.401% | -9.668% | 0.6661% |  | -0.2458% | 0.9092% | -0.6105% | -0.06339% | 0.3013% | 2.125% | -0.7929% | 1.882% | 2.769% | 4.8% | 1.152% | 0.05819%
| Z-Score (Absolute) |  | NA | 0.81 | 0.48 | 0.04 | 1.15 | 2.49 | 1.13 | 0.25 | 0.64 | 2.49 | 0.72 | 0.71 | 0.88 | 0.15 | 9.92 | 1.17 | 0.43 | 3.84 | 0.42 | NA | 0.04 | 0.52 | 0.11 | 0.12 | 0.27 | 1.02 | 0.18 | 0.92 | 1.28 | 2.12 | 0.62 | 0.17
| Analytical results for Pd in OREAS 282 (Indicative Value < 5 ppb) | [33 empty] | Indicative
Pd | Round | Replicate | Lab | [31 empty] | 1
| No | No | 2.02 | [31 empty] | ppb
|  |  | FA*OES | [31 empty] | 2
|  |  | 40g | [31 empty] | 2
| 1 | 1 | 5 | [31 empty] | 1
| 1 | 2 | < 5 | [31 empty] | 1
| 1 | 3 | < 5 | [31 empty] | 16
| 1 | 4 | < 5 | [31 empty] | < 5
| 1 | 5 | < 5 | [31 empty] | 7
| 1 | 6 | < 5
| Mean |  | 5
| Median |  | 5
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 282 (Indicative Value < 5 ppb) | [33 empty] | Indicative
Pt | Round | Replicate | Lab | [31 empty] | 1
| No | No | 2.02 | [31 empty] | ppb
|  |  | FA*OES | [31 empty] | 2
|  |  | 40g | [31 empty] | 2
| 1 | 1 | < 5 | [31 empty] | 1
| 1 | 2 | < 5 | [31 empty] | 1
| 1 | 3 | < 5 | [31 empty] | 16
| 1 | 4 | < 5 | [31 empty] | < 5
| 1 | 5 | < 5 | [31 empty] | 7
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 282 (Certified Value 14.05 ppm) | [9 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 3
| 1 | 1 | 13.9612 | 14.0668 | 14.3405 | 14.3405 | 13.8022 | 14.1805 | 14.0268 | 13.9527 | 1
| 1 | 2 | 14.0851 | 13.8842 | 14.0524 | 14.0524 | 13.8255 | 14.3877 | 14.5217 | 13.9674 | 5
| 1 | 3 | 14.0546 | 14.5778 | 13.9467 | 13.9467 | 13.6502 | 14.3426 | 13.8899 | 13.9695 | 16
| 1 | 4 | 13.9041 | 14.2789 | 14.3367 | 14.3367 | 13.9215 | 13.6311 | 14.5947 | 13.9784 | 14.0508
| 1 | 5 | 14.1988 | 14.2286 | 14.4485 | 14.4485 | 13.7967 | 14.3001 | 14.2679 | 14.0189 | 9
| 1 | 6 | 14.0465 | 13.9442 | 13.8231 | 13.8231 | 14.1091 | 13.9155 | 14.1398 | 13.8474
| 2 | 7 | 14.0894 | 14.0717 | 14.4952 | 14.4952 | 13.8244 | 14.0627 | 13.9533 | 14.639
| 2 | 8 | 13.6999 | 13.9224 | 14.2047 | 14.2047 | 13.6156 | 14.2333 | 14.0846 | 14.3339
| 2 | 9 | 14.2296 | 13.9406 | 14.2224 | 14.2224 | 13.8322 | 13.932 | 14.1979 | 14.6455
| 2 | 10 | 14.2924 | 13.7611 | 13.8274 | 13.8274 | 13.754 | 13.7338 | 14.0594 | 13.9177
| 2 | 11 | 13.7865 | 13.9233 | 14.4474 | 14.4474 | 13.8906 | 13.7114 | 14.2527 | 14.1461
| 2 | 12 | 13.9543 | 14.042 | 14.0139 | 14.0139 | 13.6552 | 14.1248 | 14.1993 | 14.3831
| 3 | 13 | 14.021 | 14.26 | 14.0073 | 14.0073
| 3 | 14 | 14.1603 | 14.05 | 14.0405 | 14.0405
| 3 | 15 | 14.0355 | 14.21 | 14.1252 | 14.1252
| 3 | 16 | 13.7119 | 14.17 | 14.0903 | 14.0903
| 3 | 17 | 14.002 | 13.78 | 13.8711 | 13.8711
| 3 | 18 | 14.2625 | 14.25 | 14.3153 | 14.3153
| 4 | 19 | 13.7712 | 13.77 | 14.09 | 14.4128
| 4 | 20 | 13.8885 | 13.79 | 14.13 | 14.0917
| 4 | 21 |  | 13.94 | 13.75 | 14.0278
| 4 | 22 |  | 13.49 | 13.78 | 14.6507
| 4 | 23 |  | 13.93 | 13.98 | 14.7267
| 4 | 24 |  | 13.69 | 13.71 | 14.139
| 5 | 25 |  |  | 14.13 | 14.0484
| 5 | 26 |  |  | 14.14 | 14.3876
| 5 | 27 |  |  | 13.66 | 13.8672
| 5 | 28 |  |  | 14.38 | 14.267
| 5 | 29 |  |  | 13.79 | 14.1861
| 5 | 30 |  |  | 13.47 | 14.0815
| 6 | 31 |  |  | 13.83 | 14.0172
| 6 | 32 |  |  | 13.85 | 13.7661
| 6 | 33 |  |  | 14.24 | 13.5861
| 6 | 34 |  |  | 14.01 | 13.6364
| 6 | 35 |  |  | 13.82 | 14.0268
| 6 | 36 |  |  | 14.04 | 14.2401
| Mean |  | 14.0078 | 13.9988 | 14.0391 | 14.1324 | 13.8064 | 14.0463 | 14.1823 | 14.15
| Median |  | 14.0283 | 13.9424 | 14.0403 | 14.091 | 13.8133 | 14.0937 | 14.1688 | 13.9986
| Std Dev. |  | 0.17613 | 0.2359 | 0.24528 | 0.26156 | 0.13479 | 0.25947 | 0.21034 | 0.2812
| Rel.Std.Dev. |  | 1.257% | 1.685% | 1.747% | 1.851% | 0.9763% | 1.847% | 1.483% | 1.987%
| PDM3 |  | -0.3061% | -0.3698% | -0.08287% | 0.5813% | -1.739% | -0.03187% | 0.9363% | 0.7059%
| Z-Score (Absolute) |  | NA | 0.37 | 0.06 | 0.67 | 1.88 | 0 | 1.06 | 0.81
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 282 (Certified Value 0.611 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.59 | 0.5 | 0.7 | 0.6 | 0.74 | 0.3 | < 1 | 0.8 | 0.6 | 0.63 | 0.6 | 0.5 | 0.62 | 0.63 | 0.64 | 0.65 | 0.62 | 0.6 | 0.598 | < 2 | 0.58 | 0.8 | 0.6 | 0.6 | 0.61 | 1
| 1 | 2 | 0.62 | 0.5 | 0.7 | 0.6 | 0.76 | 0.4 | < 1 | 0.8 | 0.6 | 0.6 | 0.64 | 0.5 | 0.62 | 0.59 | 0.63 | 0.56 | 0.61 | 0.6 | 0.636 | < 2 | 0.58 | 0.5 | 0.57 | 0.62 | 0.62 | 19
| 1 | 3 | 0.63 | 0.5 | 0.7 | 0.63 | 0.78 | 0.4 | < 1 | 0.8 | 0.6 | 0.64 | 0.6 | 0.5 | 0.64 | 0.6 | 0.67 | 0.54 | 0.61 | 0.6 | 0.604 | < 2 | 0.59 | 0.7 | 0.58 | 0.61 | 0.62 | 16
| 1 | 4 | 0.61 | 0.5 | 0.7 | 0.66 | 0.74 | 0.5 | < 1 | 0.8 | 0.7 | 0.62 | 0.59 | 0.5 | 0.61 | 0.61 | 0.62 | 0.57 | 0.6 | 0.6 | 0.608 | < 2 | 0.61 | 0.8 | 0.59 | 0.61 | 0.58 | 0.61106
| 1 | 5 | 0.63 | 0.5 | 0.7 | 0.6 | 0.78 | 0.3 | < 1 | 0.8 | 0.6 | 0.61 | 0.6 | 0.5 | 0.63 | 0.62 | 0.65 | 0.6 | 0.63 | 0.6 | 0.588 | < 2 | 0.6 | < 0.5 | 0.57 | 0.61 | 0.64 | 11
| 1 | 6 | 0.62 | 0.5 | 0.7 | 0.61 | 0.77 | 0.3 | < 1 | 0.8 | 0.7 | 0.64 | 0.61 | 0.5 | 0.66 | 0.6 | 0.62 | 0.63 | 0.61 | 0.6 | 0.585 | < 2 | 0.6 | 0.9 | 0.63 | 0.6 | 0.65
| Mean |  | 0.61667 | 0.5 | 0.7 | 0.61667 | 0.76167 | 0.36667 |  | 0.8 | 0.63333 | 0.62333 | 0.60667 | 0.5 | 0.63 | 0.60833 | 0.63833 | 0.59167 | 0.61333 | 0.6 | 0.60317 |  | 0.59333 | 0.74 | 0.59 | 0.60833 | 0.62
| Median |  | 0.62 | 0.5 | 0.7 | 0.605 | 0.765 | 0.35 |  | 0.8 | 0.6 | 0.625 | 0.6 | 0.5 | 0.625 | 0.605 | 0.635 | 0.585 | 0.61 | 0.6 | 0.601 |  | 0.595 | 0.8 | 0.585 | 0.61 | 0.62
| Std Dev. |  | 0.015055 | 0 | 1.2162e-16 | 0.024221 | 0.018348 | 0.08165 |  | 1.2162e-16 | 0.05164 | 0.01633 | 0.017512 | 0 | 0.017889 | 0.01472 | 0.019408 | 0.042622 | 0.010328 | 0 | 0.018378 |  | 0.012111 | 0.15166 | 0.022804 | 0.0075277 | 0.024495
| Rel.Std.Dev. |  | 2.441% | 0% | 1.737e-14% | 3.928% | 2.409% | 22.27% |  | 1.52e-14% | 8.154% | 2.62% | 2.887% | 0% | 2.839% | 2.42% | 3.04% | 7.204% | 1.684% | 0% | 3.047% |  | 2.041% | 20.49% | 3.865% | 1.237% | 3.951%
| PDM3 |  | 0.9176% | -18.17% | 14.56% | 0.9176% | 24.65% | -39.99% |  | 30.92% | 3.645% | 2.009% | -0.7189% | -18.17% | 3.1% | -0.4461% | 4.463% | -3.174% | 0.3721% | -1.81% | -1.292% |  | -2.901% | 21.1% | -3.446% | -0.4461% | 1.463%
| Z-Score (Absolute) |  | 0.13 | NA | NA | 0.13 | 6 | 9.98 | 4.59 | NA | NA | 0.4 | 0.27 | NA | 0.67 | 0.2 | 1.01 | 0.88 | 0 | NA | 0.41 | 15.64 | 0.81 | 1.82 | 0.94 | 0.2 | 0.27
Results from laboratories 2, 3, 8, 9, 13 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 282 (Certified Value 2.95 wt.%) | [26 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 2.98 | 2.94 | 2.68 | 3.24 | 3.107 | 3.09 | 2.856 | 2.9 | 3.44 | 3.03 | 2.91 | 2.867 | 2.98 | 2.7 | 3.02 | 3.3 | 2.93 | 2.87 | 2.98 | 2.98 | 2.842 | 2.91 | 3.02 | 3.06 | 3.04 | 1
| 1 | 2 | 2.99 | 3.05 | 2.68 | 3.38 | 3.123 | 3.08 | 2.806 | 2.99 | 3.53 | 2.935 | 3.1 | 2.885 | 2.99 | 2.58 | 3.08 | 3.36 | 2.91 | 2.63 | 2.99 | 3 | 2.7832 | 2.95 | 2.97 | 3.04 | 2.96 | #N/A
| 1 | 3 | 2.99 | 3.06 | 2.73 | 3.33 | 3.085 | 3.17 | 2.887 | 3.02 | 3.5 | 2.9843 | 2.92 | 2.886 | 3.08 | 2.64 | 3.04 | 3.3 | 2.92 | 2.74 | 2.98 | 2.92 | 2.9008 | 2.91 | 2.95 | 3.04 | 2.91 | 16
| 1 | 4 | 3.07 | 2.93 | 2.61 | 3.37 | 3.133 | 3.06 | 2.838 | 3 | 3.54 | 2.9767 | 2.84 | 2.864 | 3.02 | 2.68 | 3.04 | 3.35 | 2.91 | 2.73 | 2.98 | 2.91 | 2.9106 | 2.89 | 3 | 3.03 | 2.77 | 2.9473
| 1 | 5 | 3 | 2.99 | 2.72 | 3.28 | 3.091 | 3.14 | 2.861 | 2.94 | 3.53 | 2.9942 | 2.71 | 2.871 | 2.97 | 2.62 | 3.11 | 3.31 | 2.89 | 2.72 | 3.01 | 2.97 | 2.8812 | 2.9 | 2.99 | 3.04 | 2.98 | 12
| 1 | 6 | 3.04 | 3.05 | 2.66 | 3.27 | 3.128 | 3 | 2.846 | 2.97 | 3.51 | 2.987 | 2.89 | 2.886 | 3.03 | 2.67 | 3.09 | 3.36 | 2.92 | 2.74 | 2.91 | 2.94 | 2.9478 | 2.92 | 3.01 | 3.03 | 3.02
| Mean |  | 3.0117 | 3.0033 | 2.68 | 3.3117 | 3.1112 | 3.09 | 2.849 | 2.97 | 3.5083 | 2.9845 | 2.895 | 2.8765 | 3.0117 | 2.6483 | 3.0633 | 3.33 | 2.9133 | 2.7383 | 2.975 | 2.9533 | 2.8776 | 2.9133 | 2.99 | 3.04 | 2.9467
| Median |  | 2.995 | 3.02 | 2.68 | 3.305 | 3.115 | 3.085 | 2.851 | 2.98 | 3.52 | 2.9857 | 2.9 | 2.878 | 3.005 | 2.655 | 3.06 | 3.33 | 2.915 | 2.735 | 2.98 | 2.955 | 2.891 | 2.91 | 2.995 | 3.04 | 2.97
| Std Dev. |  | 0.035449 | 0.058538 | 0.043359 | 0.057067 | 0.020044 | 0.06 | 0.026892 | 0.043818 | 0.03656 | 0.0306 | 0.12661 | 0.010291 | 0.040702 | 0.044008 | 0.035024 | 0.029665 | 0.013663 | 0.07679 | 0.033912 | 0.03559 | 0.057874 | 0.020656 | 0.026077 | 0.010954 | 0.097912
| Rel.Std.Dev. |  | 1.177% | 1.949% | 1.618% | 1.723% | 0.6443% | 1.942% | 0.9439% | 1.475% | 1.042% | 1.025% | 4.373% | 0.3578% | 1.351% | 1.662% | 1.143% | 0.8908% | 0.469% | 2.804% | 1.14% | 1.205% | 2.011% | 0.709% | 0.8721% | 0.3603% | 3.323%
| PDM3 |  | 2.183% | 1.901% | -9.07% | 12.36% | 5.559% | 4.841% | -3.336% | 0.7696% | 19.03% | 1.263% | -1.775% | -2.403% | 2.183% | -10.14% | 3.936% | 12.98% | -1.153% | -7.091% | 0.9392% | 0.2041% | -2.365% | -1.153% | 1.448% | 3.145% | -0.02211%
| Z-Score (Absolute) |  | 0.31 | 0.24 | 2.49 | 2.84 | 1.15 | 0.97 | 1.06 | 0.04 | 4.5 | 0.08 | 0.67 | 0.83 | 0.31 | 2.75 | 0.74 | 2.99 | 0.52 | 1.99 | 0 | 0.18 | 0.82 | 0.52 | 0.13 | 0.55 | 0.24
| Analytical results for As in OREAS 282 (Certified Value 2328 ppm) | [26 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 2130 | 2330 | 2327 | 2499 | 2060 | 1412 | 2450 | > 1000 | 2439.9 | 2327.9 | 2258 | 2443 | 2280 | 2127 | 2340 | 2380 | 2330 | 1856.2 | 2158.2 | 2333 | 2290 | 2395 | 2430 |  |  | 1
| 1 | 2 | 2180 | 2480 | 2369 | 2507 | 1950 | 1368 | 2450 | > 1000 | 2363.1 | 2377.4 | 2255 | 2446.4 | 2240 | 2160 | 2340 | 2360 | 2310 | 1793.3 | 2128.5 | 2356 | 2240 | 2358.9 | 2370 |  |  | 20
| 1 | 3 | 2160 | 2490 | 2388 | 2444 | 1970 | 1394 | 2380 | > 1000 | 2380.3 | 2322 | 2207 | 2444.7 | 2290 | 2141 | 2370 | 2360 | 2250 | 1742 | 2187.9 | 2333 | 2270 | 2362 | 2280 |  |  | 16
| 1 | 4 | 2200 | 2330 | 2313 | 2534 | 1950 | 1390 | 2450 | > 1000 | 2406.7 | 2291.7 | 2227 | 2460 | 2270 | 2108 | 2320 | 2350 | 2320 | 1718.6 | 2158.2 | 2342 | 2270 | 2360.4 | 2240 |  |  | 2327.7175
| 1 | 5 | 2190 | 2450 | 2412 | 2554 | 1950 | 1369 | 2450 | > 1000 | 2394.1 | 2320.1 | 2190 | 2410.3 | 2210 | 2150 | 2350 | 2350 | 2270 | 1790.9 | 2178 | 2324 | 2280 | 2378.1 | 2390 |  |  | 13
| 1 | 6 | 2170 | 2490 | 2407 | 2483 | 1950 | 1378 | 2390 | > 1000 | 2406.2 | 2367.1 | 2197 | 2462 | 2280 | 2168 | 2390 | 2380 | 2310 | 1726.4 | 2242.2 | 2327 | 2290 | 2369.5 | 2410
| Mean |  | 2171.6667 | 2428.3333 | 2369.3333 | 2503.5 | 1971.6667 | 1385.1667 | 2428.3333 |  | 2398.3833 | 2334.3667 | 2222.3333 | 2444.4 | 2261.6667 | 2142.3333 | 2351.6667 | 2363.3333 | 2298.3333 | 1771.2333 | 2175.5 | 2335.8333 | 2273.3333 | 2370.65 | 2353.3333
| Median |  | 2175 | 2465 | 2378.5 | 2503 | 1950 | 1384 | 2450 |  | 2400.15 | 2324.95 | 2217 | 2445.55 | 2275 | 2145.5 | 2345 | 2360 | 2310 | 1766.45 | 2168.1 | 2333 | 2275 | 2365.75 | 2380
| Std Dev. |  | 24.8328 | 77.5672 | 41.3602 | 38.6614 | 44.0076 | 16.9046 | 33.7145 |  | 26.2456 | 32.0689 | 29.2689 | 18.5579 | 30.605 | 22.1149 | 24.8328 | 13.6626 | 31.2517 | 52.3983 | 38.5136 | 11.6519 | 18.619 | 13.9118 | 76.0701
| Rel.Std.Dev. |  | 1.143% | 3.194% | 1.746% | 1.544% | 2.232% | 1.22% | 1.388% |  | 1.094% | 1.374% | 1.317% | 0.7592% | 1.353% | 1.032% | 1.056% | 0.5781% | 1.36% | 2.958% | 1.77% | 0.4988% | 0.819% | 0.5868% | 3.232%
| PDM3 |  | -6.704% | 4.323% | 1.788% | 7.552% | -15.3% | -40.49% | 4.323% |  | 3.036% | 0.2856% | -4.527% | 5.013% | -2.838% | -7.964% | 1.029% | 1.53% | -1.262% | -23.91% | -6.539% | 0.3487% | -2.336% | 1.844% | 1.1%
| Z-Score (Absolute) |  | 1.32 | 0.75 | 0.28 | 1.36 | 2.94 | 7.69 | 0.75 | NA | 0.51 | 0.01 | 0.91 | 0.88 | 0.59 | 1.56 | 0.13 | 0.23 | 0.3 | 4.56 | 1.29 | 0.01 | 0.5 | 0.29 | 0.15
| Analytical results for B in OREAS 282 (Indicative Value 654 ppm) | [26 empty] | Indicative
B | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.03 | [24 empty] | ppm
|  |  | 4A*OES | [24 empty] | 0
0
| 1 | 1 | 689 | [24 empty] | 1
| 1 | 2 | 654 | [24 empty] | 3
| 1 | 3 | 614 | [24 empty] | 16
| 1 | 4 | 643 | [24 empty] | 653.8333
| 1 | 5 | 669 | [24 empty] | 9
| 1 | 6 | 654
| Mean |  | 653.8333
| Median |  | 654
| Std Dev. |  | 25.1827
| Rel.Std.Dev. |  | 3.852%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 282 (Certified Value 100 ppm) | [26 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 100 | 101 | 312 | 108 | 98 | 100 | 95 | 97.5 | 102 | 99.7 | 103 | 100 | 98.3 | 105 | 95 | 107 | 100 | 102 | 92 | 101 | 101 | 100 | 110 | 96 | 110 | 1
| 1 | 2 | 110 | 102 | 302 | 109 | 97 | 103 | 95 | 101 | 98 | 98.9 | 99 | 100 | 96.6 | 101 | 95 | 102 | 100 | 100 | 95 | 102 | 99 | 100 | 130 | 95 | 100 | 21
| 1 | 3 | 100 | 103 | 310 | 106 | 100 | 104 | 95 | 98 | 99 | 100.3 | 102 | 100 | 98.5 | 108 | 97 | 105 | 130 | 101 | 93 | 100 | 102 | 101 | 100 | 96 | 100 | 16
| 1 | 4 | 110 | 101 | 292 | 111 | 99 | 98.5 | 95 | 99.5 | 100 | 99.5 | 99 | 99 | 96.3 | 107 | 94 | 105 | 190 | 103 | 95 | 100 | 111 | 93 | 100 | 95 | 100 | 99.7694
| 1 | 5 | 100 | 101 | 301 | 108 | 101 | 102 | 95 | 98.5 | 98 | 100 | 101 | 100 | 98.3 | 103 | 96 | 102 | 130 | 100 | 94 | 101 | 101 | 99 | 100 | 95 | 100 | 14
| 1 | 6 | 110 | 102 | 297 | 102 | 100 | 98.6 | 95 | 98 | 101 | 100 | 103 | 100 | 98.7 | 108 | 95 | 101 | 100 | 102 | 93 | 100 | 100 | 99 | 100 | 95 | 110
| Mean |  | 105 | 101.6667 | 302.3333 | 107.3333 | 99.1667 | 101.0167 | 95 | 98.75 | 99.6667 | 99.7333 | 101.1667 | 99.8333 | 97.7833 | 105.3333 | 95.3333 | 103.6667 | 125 | 101.3333 | 93.6667 | 100.6667 | 102.3333 | 98.6667 | 106.6667 | 95.3333 | 103.3333
| Median |  | 105 | 101.5 | 301.5 | 108 | 99.5 | 101 | 95 | 98.25 | 99.5 | 99.85 | 101.5 | 100 | 98.3 | 106 | 95 | 103.5 | 115 | 101.5 | 93.5 | 100.5 | 101 | 99.5 | 100 | 95 | 100
| Std Dev. |  | 5.4772 | 0.8165 | 7.607 | 3.0768 | 1.472 | 2.3241 | 0 | 1.2942 | 1.633 | 0.49261 | 1.8348 | 0.40825 | 1.0477 | 2.8752 | 1.0328 | 2.3381 | 35.0714 | 1.2111 | 1.2111 | 0.8165 | 4.3665 | 2.8752 | 12.1106 | 0.5164 | 5.164
| Rel.Std.Dev. |  | 5.216% | 0.8031% | 2.516% | 2.867% | 1.484% | 2.301% | 0% | 1.311% | 1.638% | 0.4939% | 1.814% | 0.4089% | 1.071% | 2.73% | 1.083% | 2.255% | 28.06% | 1.195% | 1.293% | 0.8111% | 4.267% | 2.914% | 11.35% | 0.5417% | 4.997%
| PDM3 |  | 5.243% | 1.902% | 203% | 7.581% | -0.6042% | 1.25% | -4.78% | -1.022% | -0.103% | -0.03619% | 1.4% | 0.06404% | -1.991% | 5.577% | -4.446% | 3.906% | 25.29% | 1.568% | -6.117% | 0.8993% | 2.57% | -1.105% | 6.913% | -4.446% | 3.572%
| Z-Score (Absolute) |  | NA | 0.67 | 74.48 | 2.76 | 0.25 | 0.44 | 1.78 | 0.4 | 0.06 | 0.04 | 0.49 | 0 | 0.75 | 2.02 | 1.66 | 1.41 | NA | 0.55 | 2.27 | 0.31 | 0.92 | 0.43 | NA | 1.66 | NA
Results from laboratories 1, 20, 29 and 31 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 282 (Certified Value 0.93 ppm) | [26 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 0.9 | 1 | < 5 | 0.9 | 0.9 | 0.7 | 1 | 1 | < 1 | 0.97 | 0.83 | 0.9 | 0.96 | 1 | 0.8 | 0.6 | 0.85 | 0.8 | < 1 | 0.95 | < 1 | 0.99 | 0.86 | 0.89 | 0.94 | 1
| 1 | 2 | 0.87 | 1 | < 5 | 1 | 0.9 | 0.7 | 0.5 | 1 | < 1 | 0.97 | 1.21 | 1 | 0.96 | 0.9 | 0.9 | 0.8 | 0.91 | 0.9 | < 1 | 0.92 | < 1 | 0.86 | 0.95 | 0.9 | 0.96 | 22
| 1 | 3 | 0.87 | 1 | < 5 | 0.9 | 0.9 | 0.8 | 1 | 1 | 1 | 1.01 | 1.14 | 1 | 1.04 | 0.9 | 0.9 | 0.6 | 0.94 | 0.8 | < 1 | 0.91 | < 1 | 1 | 0.89 | 0.89 | 0.95 | 16
| 1 | 4 | 0.9 | 1 | < 5 | 1 | 0.9 | 0.7 | 1 | 1 | < 1 | 0.99 | 0.96 | 1 | 1.04 | 0.9 | 0.9 | 0.7 | 0.88 | 0.9 | < 1 | 0.91 | < 1 | 0.87 | 0.84 | 0.89 | 0.85 | 0.92801
| 1 | 5 | 0.88 | 1 | < 5 | 1 | 0.9 | 0.8 | 1 | 1 | < 1 | 1.01 | 0.89 | 0.9 | 0.93 | 0.9 | 0.9 | 0.6 | 0.9 | 0.9 | < 1 | 0.9 | < 1 | 1.07 | 0.93 | 0.89 | 0.98 | 15
| 1 | 6 | 0.89 | 1 | < 5 | 0.9 | 0.9 | 0.7 | 1 | 1 | < 1 | 0.97 | 1.05 | 1 | 0.96 | 0.9 | 0.9 | 0.6 | 0.94 | 0.9 | < 1 | 0.91 | < 1 | 0.95 | 0.9 | 0.93 | 0.92
| Mean |  | 0.885 | 1 |  | 0.95 | 0.9 | 0.73333 | 0.91667 | 1 | 1 | 0.98667 | 1.0133 | 0.96667 | 0.98167 | 0.91667 | 0.88333 | 0.65 | 0.90333 | 0.86667 |  | 0.91667 |  | 0.95667 | 0.895 | 0.89833 | 0.93333
| Median |  | 0.885 | 1 |  | 0.95 | 0.9 | 0.7 | 1 | 1 | 1 | 0.98 | 1.005 | 1 | 0.96 | 0.9 | 0.9 | 0.6 | 0.905 | 0.9 |  | 0.91 |  | 0.97 | 0.895 | 0.89 | 0.945
| Std Dev. |  | 0.013784 | 0 |  | 0.054772 | 0 | 0.05164 | 0.20412 | 0 |  | 0.019664 | 0.14679 | 0.05164 | 0.046655 | 0.040825 | 0.040825 | 0.083666 | 0.035024 | 0.05164 |  | 0.017512 |  | 0.080911 | 0.041352 | 0.016021 | 0.045461
| Rel.Std.Dev. |  | 1.558% | 0% |  | 5.766% | 0% | 7.042% | 22.27% | 0% |  | 1.993% | 14.49% | 5.342% | 4.753% | 4.454% | 4.622% | 12.87% | 3.877% | 5.958% |  | 1.91% |  | 8.458% | 4.62% | 1.783% | 4.871%
| PDM3 |  | -4.635% | 7.757% |  | 2.37% | -3.018% | -20.98% | -1.222% | 7.757% | 7.757% | 6.321% | 9.194% | 4.166% | 5.782% | -1.222% | -4.814% | -29.96% | -2.659% | -6.61% |  | -1.222% |  | 3.088% | -3.557% | -3.198% | 0.5737%
| Z-Score (Absolute) |  | 0.4 | NA | 25.73 | 0.65 | 0.16 | 2.86 | 0.11 | NA | NA | 1.24 | 1.67 | 0.92 | 1.16 | 0.11 | 0.43 | 4.21 | 0.11 | 0.7 | 6.64 | 0.11 | 6.64 | 0.76 | 0.24 | 0.19 | 0.38
Results from laboratories 2, 8 and 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 282 (Certified Value 0.93 ppm) | [26 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.84 | 0.9 | < 5 | 0.94 | 1.1 | 0.9 | 0.9 | 0.9 | 0.97 | 0.99 | 0.93 | 0.92 | 0.95 | 0.87 | 0.8 | 0.95 | 0.96 | 0.87 | 0.77 | 0.88 | 1.08 | 0.93 | 0.95 | 0.97 |  | 1
| 1 | 2 | 0.85 | 0.9 | < 5 | 0.96 | 1.1 | 0.9 | 0.9 | 0.9 | 0.93 | 0.97 | 0.95 | 0.93 | 0.92 | 0.89 | 0.9 | 0.95 | 0.98 | 0.88 | 0.78 | 0.88 | 1.07 | 0.9 | 0.94 | 0.93 |  | 23
| 1 | 3 | 0.85 | 1 | < 5 | 0.94 | 1.1 | 0.9 | 1 | 0.9 | 0.97 | 0.96 | 0.94 | 0.94 | 0.94 | 0.9 | 0.9 | 0.96 | 0.98 | 0.88 | 0.76 | 0.92 | 1.04 | 0.87 | 0.93 | 0.96 |  | 16
| 1 | 4 | 0.83 | 1 | < 5 | 0.96 | 1.1 | 0.9 | 0.9 | 0.9 | 0.95 | 0.96 | 0.95 | 0.93 | 0.94 | 0.87 | 0.8 | 0.95 | 1 | 0.9 | 0.76 | 0.89 | 0.98 | 0.94 | 0.94 | 0.92 |  | 0.93214
| 1 | 5 | 0.87 | 0.9 | < 5 | 0.97 | 1.1 | 0.9 | 0.9 | 0.9 | 0.94 | 0.98 | 0.92 | 0.93 | 0.88 | 0.87 | 0.9 | 0.96 | 0.98 | 0.89 | 0.76 | 0.9 | 1.05 | 0.96 | 0.94 | 0.96 |  | 16
| 1 | 6 | 0.86 | 1 | < 5 | 0.94 | 1.1 | 0.9 | 1 | 1 | 0.95 | 0.97 | 0.92 | 0.93 | 0.94 | 0.91 | 0.9 | 0.98 | 0.99 | 0.87 | 0.75 | 0.91 | 1.02 | 0.91 | 0.95 | 0.95
| Mean |  | 0.85 | 0.95 |  | 0.95167 | 1.1 | 0.9 | 0.93333 | 0.91667 | 0.95167 | 0.97167 | 0.935 | 0.93 | 0.92833 | 0.885 | 0.86667 | 0.95833 | 0.98167 | 0.88167 | 0.76333 | 0.89667 | 1.04 | 0.91833 | 0.94167 | 0.94833
| Median |  | 0.85 | 0.95 |  | 0.95 | 1.1 | 0.9 | 0.9 | 0.9 | 0.95 | 0.97 | 0.935 | 0.93 | 0.94 | 0.88 | 0.9 | 0.955 | 0.98 | 0.88 | 0.76 | 0.895 | 1.045 | 0.92 | 0.94 | 0.955
| Std Dev. |  | 0.014142 | 0.054772 |  | 0.013292 | 2.4324e-16 | 0 | 0.05164 | 0.040825 | 0.016021 | 0.01169 | 0.013784 | 0.0063246 | 0.025626 | 0.017607 | 0.05164 | 0.01169 | 0.013292 | 0.01169 | 0.010328 | 0.01633 | 0.036332 | 0.031885 | 0.0075277 | 0.019408
| Rel.Std.Dev. |  | 1.664% | 5.766% |  | 1.397% | 2.211e-14% | 0% | 5.533% | 4.454% | 1.683% | 1.203% | 1.474% | 0.6801% | 2.76% | 1.989% | 5.958% | 1.22% | 1.354% | 1.326% | 1.353% | 1.821% | 3.493% | 3.472% | 0.7994% | 2.047%
| PDM3 |  | -8.812% | 1.916% |  | 2.095% | 18.01% | -3.448% | 0.1283% | -1.66% | 2.095% | 4.241% | 0.3071% | -0.2293% | -0.4081% | -5.057% | -7.024% | 2.81% | 5.314% | -5.415% | -18.11% | -3.805% | 11.57% | -1.481% | 1.022% | 1.738%
| Z-Score (Absolute) |  | 1.56 | NA | 28.78 | 0.31 | NA | NA | NA | NA | 0.31 | 0.67 | 0 | 0.09 | 0.12 | 0.92 | 1.26 | 0.43 | 0.86 | 0.98 | 3.16 | 0.7 | 1.93 | 0.31 | 0.12 | 0.25
Results from laboratories 2, 5, 7, 8 and 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 282 (Certified Value 17.47 wt.%) | [26 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 18.05 | 18.5 | 16.63 | > 15 | 18.225 | 16.3845 | 18.1 | 19.8 | 18.2403 | 17.97 | 16.103 | 18.67 | 16.3 | > 15 | 18.3 | 17.6 | 17.6 | 17.25 | 16.4 | 16.665 | > 10 | 18.5 | 17.24 | 18.5 |  | 1
| 1 | 2 | 18 | 18.5 | 17.16 | > 15 | 18.511 | 16.091 | 18.5 | 20.28 | 17.6745 | 19.02 | 16.156 | 18.71 | 15.7 | > 15 | 18.4 | 17.45 | 17.2 | 17.32 | 16.3 | 16.16 | > 10 | 18.25 | 17.4 | 17.85 |  | #N/A
| 1 | 3 | 17.95 | 18.4 | 17.01 | > 15 | 18.368 | 16.56 | 18.2 | 20.38 | 17.795 | 18.04 | 16.427 | 18.93 | 16.2 | > 15 | 18.7 | 17.5 | 17.7 | 17.06 | 15.6 | 16.766 | > 10 | 17.95 | 17.17 | 17.5 |  | 16
| 1 | 4 | 18.5 | 18.4 | 16.41 | > 15 | 18.296 | 16.314 | 18.2 | 20.56 | 17.6944 | 17.32 | 16.309 | 18.3 | 16.1 | > 15 | 18 | 17.55 | 17.2 | 17.34 | 16 | 16.867 | > 10 | 18.1 | 16.82 | 16.85 |  | 17.4672
| 1 | 5 | 18 | 19 | 17.12 | > 15 | 18.296 | 16.2685 | 18 | 20.57 | 17.9975 | 17.15 | 16.383 | 18.3 | 15.7 | > 15 | 18.1 | 17.35 | 17.5 | 17.21 | 15.9 | 16.463 | > 10 | 18.15 | 16.92 | 18.15 |  | 17
| 1 | 6 | 18.25 | 18.3 | 18.63 | > 15 | 18.439 | 16.329 | 18.5 | 20.3 | 17.7799 | 17.96 | 16.004 | 18.45 | 16 | > 15 | 19.3 | 17.5 | 17.2 | 16.61 | 15.9 | 16.362 | > 10 | 18.4 | 17.05 | 18.3
| Mean |  | 18.125 | 18.5167 | 17.16 |  | 18.3558 | 16.3245 | 18.25 | 20.315 | 17.8636 | 17.91 | 16.2303 | 18.56 | 16 |  | 18.4667 | 17.4917 | 17.4 | 17.1317 | 16.0167 | 16.5472 |  | 18.225 | 17.1 | 17.8583
| Median |  | 18.025 | 18.45 | 17.065 |  | 18.332 | 16.3215 | 18.2 | 20.34 | 17.7874 | 17.965 | 16.2325 | 18.56 | 16.05 |  | 18.35 | 17.5 | 17.35 | 17.23 | 15.95 | 16.564 |  | 18.2 | 17.11 | 18
| Std Dev. |  | 0.21154 | 0.24833 | 0.77825 |  | 0.10528 | 0.15287 | 0.20736 | 0.28141 | 0.21722 | 0.66042 | 0.16801 | 0.25251 | 0.25298 |  | 0.4761 | 0.08612 | 0.22804 | 0.27433 | 0.29269 | 0.26658 |  | 0.20187 | 0.21345 | 0.60615
| Rel.Std.Dev. |  | 1.167% | 1.341% | 4.535% |  | 0.5735% | 0.9365% | 1.136% | 1.385% | 1.216% | 3.687% | 1.035% | 1.36% | 1.581% |  | 2.578% | 0.4923% | 1.311% | 1.601% | 1.827% | 1.611% |  | 1.108% | 1.248% | 3.394%
| PDM3 |  | 3.766% | 6.008% | -1.759% |  | 5.088% | -6.542% | 4.482% | 16.3% | 2.27% | 2.535% | -7.081% | 6.257% | -8.4% |  | 5.722% | 0.1403% | -0.3845% | -1.921% | -8.304% | -5.267% |  | 4.339% | -2.102% | 2.239%
| Z-Score (Absolute) |  | 0.27 | 0.67 | 0.72 | NA | 0.51 | 1.57 | 0.4 | 2.52 | 0.01 | 0.05 | 1.67 | 0.72 | 1.9 | NA | 0.62 | 0.38 | 0.47 | 0.74 | 1.89 | 1.34 | NA | 0.38 | 0.78 | 0
| Analytical results for Cd in OREAS 282 (Certified Value 1.08 ppm) | [26 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.07 | 1 | 16 | 1.05 | 1.3 | 1 | 1 | 1.3 | 1.14 | 1.15 | 1.07 | 1.02 | 1 | 1.08 | 1.1 | 1.09 | 1.13 | 0.99 | 1.29 | 1.04 | 1.03 | 1.15 | 1.06 | 1.08 |  | 1
| 1 | 2 | 1.11 | 1 | 16 | 0.99 | 1.3 | 1 | 1 | 1.1 | 1.12 | 1.08 | 1.09 | 0.99 | 1.1 | 1.13 | 1.1 | 1.11 | 1.16 | 0.99 | 1.26 | 1.06 | 1.04 | 1.05 | 1.09 | 1.13 |  | 24
| 1 | 3 | 1.14 | 1 | 15 | 1.08 | 1.3 | 1 | 1 | 1.2 | 1.04 | 1.3 | 1.07 | 0.98 | 0.9 | 1.09 | 1.1 | 1.08 | 1.14 | 0.97 | 1.24 | 1.05 | 1 | 1.07 | 1.1 | 1.15 |  | 16
| 1 | 4 | 1.1 | 1 | 14 | 1.02 | 1.3 | 1 | 1 | 1.2 | 1.06 | 1.04 | 1.1 | 1.06 | 0.9 | 1.06 | 1 | 1.12 | 1.18 | 1 | 1.22 | 1.03 | 0.96 | 1.1 | 1.06 | 1.03 |  | 1.0798
| 1 | 5 | 1.12 | 1 | 16 | 1 | 1.3 | 1 | 1 | 1.2 | 1.13 | 1.23 | 1.13 | 1.03 | 0.9 | 1.1 | 1.1 | 1.14 | 1.09 | 0.99 | 1.2 | 1.04 | 1.06 | 1.09 | 1.06 | 1.14 |  | 18
| 1 | 6 | 1.11 | 1 | 16 | 1.04 | 1.3 | 1 | 1 | 1.2 | 1.09 | 0.98 | 1.14 | 1 | 0.9 | 1.09 | 1.1 | 1.14 | 1.04 | 0.98 | 1.18 | 1.08 | 1.02 | 1.06 | 1.16 | 1.08
| Mean |  | 1.1083 | 1 | 15.5 | 1.03 | 1.3 | 1 | 1 | 1.2 | 1.0967 | 1.13 | 1.1 | 1.0133 | 0.95 | 1.0917 | 1.0833 | 1.1133 | 1.1233 | 0.98667 | 1.2317 | 1.05 | 1.0183 | 1.0867 | 1.0883 | 1.1017
| Median |  | 1.11 | 1 | 16 | 1.03 | 1.3 | 1 | 1 | 1.2 | 1.105 | 1.115 | 1.095 | 1.01 | 0.9 | 1.09 | 1.1 | 1.115 | 1.135 | 0.99 | 1.23 | 1.045 | 1.025 | 1.08 | 1.075 | 1.105
| Std Dev. |  | 0.023166 | 0 | 0.83666 | 0.033466 | 0 | 0 | 0 | 0.063246 | 0.040332 | 0.12033 | 0.029665 | 0.029439 | 0.083666 | 0.023166 | 0.040825 | 0.025033 | 0.050859 | 0.010328 | 0.040208 | 0.017889 | 0.034881 | 0.036148 | 0.0392 | 0.046224
| Rel.Std.Dev. |  | 2.09% | 0% | 5.398% | 3.249% | 0% | 0% | 0% | 5.27% | 3.678% | 10.65% | 2.697% | 2.905% | 8.807% | 2.122% | 3.768% | 2.249% | 4.528% | 1.047% | 3.265% | 1.704% | 3.425% | 3.326% | 3.602% | 4.196%
| PDM3 |  | 2.639% | -7.393% | 1335% | -4.615% | 20.39% | -7.393% | -7.393% | 11.13% | 1.559% | 4.646% | 1.868% | -6.158% | -12.02% | 1.096% | 0.3241% | 3.102% | 4.028% | -8.628% | 14.06% | -2.763% | -5.695% | 0.6328% | 0.7872% | 2.022%
| Z-Score (Absolute) |  | 0.4 | NA | NA | 1.79 | NA | NA | NA | NA | 0.07 | 1 | 0.16 | 2.26 | NA | 0.07 | NA | 0.53 | 0.81 | 3 | 3.84 | 1.23 | 2.12 | 0.21 | 0.16 | 0.21
Results from laboratories 2, 3, 5, 7, 8, 9, 16 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 282 (Certified Value 36.9 ppm) | [26 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 39.7 | 37.7 | 38.42 | 31 | 36 | 35.7 | 37 | 36.88 | 35.23 | 36.63 | 34.4 | 35.56 | 33.6 | 37.8 | 34.36 | 42.8 | 35.24 | 36.42 | 41.2 | 41.86 | 39.3 |  |  |  |  | 1
| 1 | 2 | 42 | 37.9 | 40.91 | 31 | 35.5 | 36 | 37 | 36.08 | 34.26 | 36.24 | 33.4 | 35.8 | 34.8 | 37.6 | 36.32 | 42.1 | 35.44 | 36.26 | 39.4 | 41.33 | 38 |  |  |  |  | 25
| 1 | 3 | 42.7 | 38.6 | 37.49 | 32 | 36.5 | 35.8 | 37 | 36.77 | 36.24 | 37.12 | 35.5 | 35.56 | 34.9 | 38 | 34.87 | 42.6 | 37.32 | 36.28 | 37.8 | 40.95 | 38.9 |  |  |  |  | 16
| 1 | 4 | 41.9 | 37.3 | 40.24 | 31 | 36 | 36.2 | 37 | 36.49 | 34.28 | 36.92 | 34.5 | 37.03 | 33.7 | 37.6 | 36.03 | 41.7 | 36.43 | 33.92 | 40.2 | 40.41 | 36.6 |  |  |  |  | 36.8927
| 1 | 5 | 41.4 | 37.4 | 39.22 | 32 | 36 | 35.9 | 37 | 36.4 | 35.61 | 36.31 | 33.3 | 36.14 | 34.3 | 37.5 | 36.7 | 42.5 | 36.33 | 36.2 | 39.9 | 40.34 | 39.4 |  |  |  |  | 19
| 1 | 6 | 39.8 | 38.2 | 37.13 | 32 | 37 | 35.8 | 37 | 36.63 | 36.08 | 37.33 | 34.1 | 36.67 | 33.3 | 37.1 | 35.93 | 41.2 | 36.2 | 36.17 | 39.8 | 40.73 | 38.7
| Mean |  | 41.25 | 37.85 | 38.9017 | 31.5 | 36.1667 | 35.9 | 37 | 36.5417 | 35.2833 | 36.7583 | 34.2 | 36.1267 | 34.1 | 37.6 | 35.7017 | 42.15 | 36.16 | 35.875 | 39.7167 | 40.9367 | 38.4833
| Median |  | 41.65 | 37.8 | 38.82 | 31.5 | 36 | 35.85 | 37 | 36.56 | 35.42 | 36.775 | 34.25 | 35.97 | 34 | 37.6 | 35.98 | 42.3 | 36.265 | 36.23 | 39.85 | 40.84 | 38.8
| Std Dev. |  | 1.2341 | 0.49295 | 1.5026 | 0.54772 | 0.5164 | 0.17889 | 0 | 0.28646 | 0.86178 | 0.44043 | 0.80994 | 0.60997 | 0.66633 | 0.30332 | 0.89785 | 0.6091 | 0.75017 | 0.96166 | 1.1179 | 0.58026 | 1.0496
| Rel.Std.Dev. |  | 2.992% | 1.302% | 3.863% | 1.739% | 1.428% | 0.4983% | 0% | 0.7839% | 2.442% | 1.198% | 2.368% | 1.688% | 1.954% | 0.8067% | 2.515% | 1.445% | 2.075% | 2.681% | 2.815% | 1.417% | 2.727%
| PDM3 |  | 11.81% | 2.595% | 5.445% | -14.62% | -1.968% | -2.691% | 0.2907% | -0.9516% | -4.362% | -0.3643% | -7.299% | -2.076% | -7.57% | 1.917% | -3.228% | 14.25% | -1.986% | -2.759% | 7.654% | 10.96% | 4.311%
| Z-Score (Absolute) |  | 2.52 | 0.7 | 1.26 | 2.7 | 0.2 | 0.34 | 0.25 | 0 | 0.67 | 0.12 | 1.25 | 0.22 | 1.31 | 0.57 | 0.45 | 3.01 | 0.2 | 0.36 | 1.7 | 2.36 | 1.04
| Analytical results for Co in OREAS 282 (Certified Value 3.8 ppm) | [26 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.5 | < 5 | 4 | 4.2 | 3.97 | 3.7 | < 5 | 3 | 3.6 | 3.7 | 3.9 | 3.5 | 3.6 | 3.9 | 3.9 | 3.8 | 4.2 | 4 | 3.8 | 3.5 | 3.5 | 4.6 | 3.7 | 4.2 |  | 1
| 1 | 2 | 3.6 | < 5 | 5 | 4.4 | 3.89 | 3.9 | < 5 | 3 | 3.6 | 3.7 | 3.8 | 3.4 | 3.6 | 3.5 | 4 | 4 | 4.2 | 4 | 3.9 | 3.4 | 3.7 | 4.7 | 3.6 | 4.1 |  | 26
| 1 | 3 | 3.7 | 5 | 5 | 4.3 | 3.94 | 4 | < 5 | 3 | 3.7 | 3.7 | 3.8 | 3.6 | 3.7 | 3.7 | 4 | 3.9 | 4.2 | 4 | 3.8 | 3.6 | 3.7 | 4.4 | 3.6 | 4.2 |  | 16
| 1 | 4 | 3.6 | < 5 | 4 | 4.4 | 3.87 | 3.9 | < 5 | 3 | 3.6 | 3.6 | 3.7 | 3.6 | 3.7 | 3.8 | 3.9 | 3.8 | 4.1 | 4 | 3.8 | 3.6 | 3.2 | 4.6 | 3.6 | 3.9 |  | 3.7994
| 1 | 5 | 3.7 | < 5 | 5 | 4.4 | 3.88 | 4 | < 5 | 3 | 3.8 | 3.7 | 3.9 | 3.4 | 3.6 | 3.5 | 4.1 | 3.9 | 4 | 4 | 3.8 | 3.6 | 3.7 | 4.8 | 3.6 | 4.2 |  | 20
| 1 | 6 | 3.6 | 5 | 5 | 4.2 | 3.99 | 3.9 | < 5 | 3 | 3.6 | 3.7 | 3.8 | 3.4 | 3.6 | 3.7 | 4 | 3.8 | 4.1 | 4 | 3.7 | 3.5 | 3.5 | 4.5 | 3.6 | 4.3
| Mean |  | 3.6167 | 5 | 4.6667 | 4.3167 | 3.9233 | 3.9 |  | 3 | 3.65 | 3.6833 | 3.8167 | 3.4833 | 3.6333 | 3.6833 | 3.9833 | 3.8667 | 4.1333 | 4 | 3.8 | 3.5333 | 3.55 | 4.6 | 3.6167 | 4.15
| Median |  | 3.6 | 5 | 5 | 4.35 | 3.915 | 3.9 |  | 3 | 3.6 | 3.7 | 3.8 | 3.45 | 3.6 | 3.7 | 4 | 3.85 | 4.15 | 4 | 3.8 | 3.55 | 3.6 | 4.6 | 3.6 | 4.2
| Std Dev. |  | 0.075277 | 0 | 0.5164 | 0.098319 | 0.050465 | 0.10954 |  | 0 | 0.083666 | 0.040825 | 0.075277 | 0.098319 | 0.05164 | 0.16021 | 0.075277 | 0.08165 | 0.08165 | 0 | 0.063246 | 0.08165 | 0.19748 | 0.14142 | 0.040825 | 0.13784
| Rel.Std.Dev. |  | 2.081% | 0% | 11.07% | 2.278% | 1.286% | 2.809% |  | 0% | 2.292% | 1.108% | 1.972% | 2.823% | 1.421% | 4.35% | 1.89% | 2.112% | 1.975% | 0% | 1.664% | 2.311% | 5.563% | 3.074% | 1.129% | 3.321%
| PDM3 |  | -4.811% | 31.6% | 22.82% | 13.61% | 3.261% | 2.647% |  | -21.04% | -3.933% | -3.056% | 0.4533% | -8.32% | -4.372% | -3.056% | 4.84% | 1.769% | 8.788% | 5.279% | 0.01462% | -7.004% | -6.565% | 21.07% | -4.811% | 9.226%
| Z-Score (Absolute) |  | 0.5 | NA | NA | 2.28 | 0.72 | 0.63 | 4.93 | NA | 0.36 | 0.23 | 0.3 | 1.02 | 0.43 | 0.23 | 0.96 | 0.5 | 1.55 | NA | 0.23 | 0.83 | 0.76 | 3.4 | 0.5 | 1.62
Results from laboratories 2, 3, 8 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 282 (Certified Value 32.5 ppm) | [26 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
1
| 1 | 1 | 34 | 30 | 30 | 36 | 41 | 25 | 30 | 36 | 41 | 36 | 37 | 23 | 39 | 24 | 35 | 28 | 30 | 28 | 45 | 18 | 32 | 28 | 28 |  |  | 1
| 1 | 2 | 33 | 20 | 28 | 39 | 41 | 40 | 30 | 36 | 39 | 38 | 37 | 23 | 28 | 30 | 30 | 30 | 30 | 26 | 44 | 23 | 32 | 26 | 26 |  |  | 27
| 1 | 3 | 34 | 30 | 27 | 36 | 40 | 30 | 30 | 36 | 40 | 34 | 40 | 24 | 29 | 24 | 30 | 29 | 30 | 26 | 46 | 29 | 31 | 26 | 26 |  |  | 16
| 1 | 4 | 33 | 30 | 28 | 38 | 41 | 30 | 30 | 37 | 40 | 34 | 40 | 22 | 27 | 27 | 35 | 31 | 30 | 26 | 44 | 39 | 30 | 28 | 26 |  |  | 32.5159
| 1 | 5 | 34 | 20 | 28 | 38 | 41 | 35 | 30 | 37 | 40 | 38 | 37 | 19 | 27 | 24 | 30 | 29 | 30 | 27 | 46 | 39 | 33 | 26 | 26 |  |  | 21
| 1 | 6 | 33 | 30 | 29 | 38 | 42 | 40 | 30 | 35 | 40 | 41 | 37 | 19 | 37 | 26 | 35 | 30 | 30 | 26 | 44 | 38 | 33 | 26 | 28
| Mean |  | 33.5 | 26.6667 | 28.3333 | 37.5 | 41 | 33.3333 | 30 | 36.1667 | 40 | 36.8333 | 38 | 21.6667 | 31.1667 | 25.8333 | 32.5 | 29.5 | 30 | 26.5 | 44.8333 | 31 | 31.8333 | 26.6667 | 26.6667
| Median |  | 33.5 | 30 | 28 | 38 | 41 | 32.5 | 30 | 36 | 40 | 37 | 37 | 22.5 | 28.5 | 25 | 32.5 | 29.5 | 30 | 26 | 44.5 | 33.5 | 32 | 26 | 26
| Std Dev. |  | 0.54772 | 5.164 | 1.0328 | 1.2247 | 0.63246 | 6.0553 | 0 | 0.75277 | 0.63246 | 2.7142 | 1.5492 | 2.1602 | 5.3821 | 2.4014 | 2.7386 | 1.0488 | 0 | 0.83666 | 0.98319 | 9.0995 | 1.169 | 1.0328 | 1.0328
| Rel.Std.Dev. |  | 1.635% | 19.36% | 3.645% | 3.266% | 1.543% | 18.17% | 0% | 2.081% | 1.581% | 7.369% | 4.077% | 9.97% | 17.27% | 9.296% | 8.427% | 3.555% | 0% | 3.157% | 2.193% | 29.35% | 3.672% | 3.873% | 3.873%
| PDM3 |  | 3.027% | -17.99% | -12.86% | 15.33% | 26.09% | 2.514% | -7.737% | 11.23% | 23.02% | 13.28% | 16.87% | -33.37% | -4.149% | -20.55% | -0.04882% | -9.275% | -7.737% | -18.5% | 37.88% | -4.662% | -2.099% | -17.99% | -17.99%
| Z-Score (Absolute) |  | 0.22 | NA | 0.47 | 0.76 | 1.24 | 0.2 | NA | 0.58 | 1.1 | 0.67 | 0.83 | 1.37 | 0.09 | 0.81 | 0.09 | 0.31 | 0.25 | 0.72 | 1.75 | 0.11 | 0 | 0.7 | 0.7
Results from laboratories 2 and 8 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 282 (Certified Value 8.24 ppm) | [26 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 8.55 | 8.2 | 8.21 | 8.5 | 7.5 | 8.5 | 8.7 | 7.95 | 8.7 | 8 | 8.56 | 7.62 | 9 | 8.14 | 8.94 | 7.8 | 7.6 | 8.3 | 8.52 | 8.51 | 8.57 | 8.65 |  |  |  | 1
| 1 | 2 | 8.69 | 8.1 | 8.14 | 8.4 | 7.5 | 8.5 | 8.1 | 7.73 | 8.71 | 8 | 8.39 | 7.19 | 9 | 7.49 | 8.85 | 7.5 | 8 | 8.2 | 8.44 | 8.45 | 8.58 | 8.54 |  |  |  | 28
| 1 | 3 | 8.66 | 8.2 | 8.26 | 8.5 | 8 | 8.4 | 8.4 | 7.9 | 8.48 | 7.7 | 8.52 | 7.56 | 9 | 8.39 | 8.59 | 7.8 | 7.7 | 8.3 | 8.46 | 8.15 | 8.51 | 8.64 |  |  |  | 16
| 1 | 4 | 8.82 | 8 | 8.21 | 8.4 | 7.5 | 8.5 | 8.3 | 7.78 | 8.01 | 8 | 8.38 | 7.56 | 9 | 7.46 | 8.87 | 7.9 | 7.9 | 8.3 | 7.84 | 8.34 | 8.36 | 8.06 |  |  |  | 8.2399
| 1 | 5 | 8.47 | 8.1 | 8.17 | 8.5 | 7.5 | 8.5 | 8.4 | 7.79 | 8.59 | 7.8 | 8.34 | 7.45 | 9 | 8.46 | 8.75 | 7.9 | 7.9 | 8.5 | 8.49 | 8.49 | 8.33 | 8.66 |  |  |  | 22
| 1 | 6 | 8.52 | 8.2 | 8.34 | 8.6 | 8 | 8.5 | 8.1 | 7.9 | 8.78 | 7.8 | 8.6 | 7.47 | 9 | 8.14 | 8.78 | 7.6 | 7.7 | 8.4 | 8.38 | 8.54 | 8.4 | 8.75
| Mean |  | 8.6183 | 8.1333 | 8.2217 | 8.4833 | 7.6667 | 8.4833 | 8.3333 | 7.8417 | 8.545 | 7.8833 | 8.465 | 7.475 | 9 | 8.0133 | 8.7967 | 7.75 | 7.8 | 8.3333 | 8.355 | 8.4133 | 8.4583 | 8.55
| Median |  | 8.605 | 8.15 | 8.21 | 8.5 | 7.5 | 8.5 | 8.35 | 7.845 | 8.645 | 7.9 | 8.455 | 7.515 | 9 | 8.14 | 8.815 | 7.8 | 7.8 | 8.3 | 8.45 | 8.47 | 8.455 | 8.645
| Std Dev. |  | 0.12952 | 0.08165 | 0.070828 | 0.075277 | 0.2582 | 0.040825 | 0.22509 | 0.086583 | 0.28247 | 0.13292 | 0.1084 | 0.1532 | 0 | 0.4367 | 0.1216 | 0.16432 | 0.15492 | 0.10328 | 0.25673 | 0.14652 | 0.10907 | 0.24916
| Rel.Std.Dev. |  | 1.503% | 1.004% | 0.8615% | 0.8874% | 3.368% | 0.4812% | 2.701% | 1.104% | 3.306% | 1.686% | 1.281% | 2.049% | 0% | 5.45% | 1.382% | 2.12% | 1.986% | 1.239% | 3.073% | 1.741% | 1.29% | 2.914%
| PDM3 |  | 4.593% | -1.293% | -0.2211% | 2.954% | -6.957% | 2.954% | 1.134% | -4.833% | 3.703% | -4.327% | 2.732% | -9.283% | 9.225% | -2.749% | 6.757% | -5.945% | -5.339% | 1.134% | 1.397% | 2.105% | 2.651% | 3.764%
| Z-Score (Absolute) |  | 0.91 | 0.64 | 0.36 | 0.48 | 2.12 | 0.48 | 0 | 1.57 | 0.67 | 1.43 | 0.42 | 2.73 | NA | 1.02 | 1.48 | 1.86 | 1.7 | 0 | 0.07 | 0.25 | 0.4 | 0.69
Results from laboratory 18 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 282 (Certified Value 19.2 ppm) | [26 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 18 | 18 | 24 | 21.9 | 19.3 | 19.5 | 20 | 28 | 12.1 | 18.6 | 18.6 | 18 | 17.9 | 18.9 | 20.7 | 25 | 23.8 | 18 | 17.7 | 21.4 | 19.5 | 17.7 | 20.1 | 18.1 | 20 | 1
| 1 | 2 | 18.8 | 20 | 25 | 22.6 | 18.9 | 20.4 | 20 | 28 | 12.4 | 17.9 | 18.4 | 18 | 17.8 | 20.8 | 20.8 | 20 | 20.4 | 18 | 18.3 | 21.5 | 18.4 | 17.7 | 20.2 | 17.9 | 21.2 | 29
| 1 | 3 | 18.7 | 22 | 25 | 22.2 | 18.9 | 19.7 | 20 | 28 | 12.1 | 18.3 | 18.7 | 18 | 17.9 | 18 | 21 | 22 | 20.7 | 19 | 18 | 20.1 | 18 | 17.9 | 19.2 | 17.9 | 20.6 | 16
| 1 | 4 | 18.5 | 18 | 25 | 21.9 | 18.8 | 20 | 20 | 28 | 12.1 | 18.1 | 18.1 | 18 | 18.1 | 19 | 19.7 | 20 | 20.8 | 19 | 18.1 | 20.3 | 18.8 | 16.5 | 19.4 | 17.7 | 19 | 19.1511
| 1 | 5 | 19.4 | 18 | 26 | 22.3 | 19.1 | 19.9 | 20 | 26 | 12.4 | 18.5 | 18.4 | 17 | 17.8 | 18.7 | 20.3 | 22 | 20.2 | 17 | 17.9 | 20.3 | 17.9 | 18.3 | 19.8 | 17.7 | 20.9 | 23
| 1 | 6 | 18.6 | 18 | 23 | 21.7 | 19.4 | 19.3 | 20 | 28 | 12.5 | 18.3 | 19.4 | 17 | 18.1 | 18.1 | 20.4 | 22 | 20.9 | 18 | 17.8 | 20.5 | 18 | 16.9 | 20 | 17.8 | 21
| Mean |  | 18.6667 | 19 | 24.6667 | 22.1 | 19.0667 | 19.8 | 20 | 27.6667 | 12.2667 | 18.2833 | 18.6 | 17.6667 | 17.9333 | 18.9167 | 20.4833 | 21.8333 | 21.1333 | 18.1667 | 17.9667 | 20.6833 | 18.4333 | 17.5 | 19.7833 | 17.85 | 20.45
| Median |  | 18.65 | 18 | 25 | 22.05 | 19 | 19.8 | 20 | 28 | 12.25 | 18.3 | 18.5 | 18 | 17.9 | 18.8 | 20.55 | 22 | 20.75 | 18 | 17.95 | 20.4 | 18.2 | 17.7 | 19.9 | 17.85 | 20.75
| Std Dev. |  | 0.45461 | 1.6733 | 1.0328 | 0.32863 | 0.24221 | 0.38987 | 0 | 0.8165 | 0.18619 | 0.25626 | 0.44272 | 0.5164 | 0.13663 | 1.0108 | 0.46224 | 1.8348 | 1.3322 | 0.75277 | 0.21602 | 0.608 | 0.62183 | 0.66933 | 0.40208 | 0.15166 | 0.82401
| Rel.Std.Dev. |  | 2.435% | 8.807% | 4.187% | 1.487% | 1.27% | 1.969% | 0% | 2.951% | 1.518% | 1.402% | 2.38% | 2.923% | 0.7619% | 5.343% | 2.257% | 8.404% | 6.304% | 4.144% | 1.202% | 2.94% | 3.373% | 3.825% | 2.032% | 0.8496% | 4.029%
| PDM3 |  | -2.53% | -0.789% | 28.8% | 15.4% | -0.4409% | 3.388% | 4.433% | 44.47% | -35.95% | -4.531% | -2.878% | -7.751% | -6.359% | -1.224% | 6.956% | 14.01% | 10.35% | -5.14% | -6.185% | 8.001% | -3.748% | -8.621% | 3.301% | -6.794% | 6.782%
| Z-Score (Absolute) |  | 0.18 | 0.03 | 3.61 | 1.99 | 0.07 | 0.53 | NA | 5.51 | 4.23 | 0.43 | 0.23 | 0.82 | 0.65 | 0.03 | 0.96 | 1.82 | 1.37 | 0.5 | 0.63 | 1.09 | 0.33 | 0.92 | 0.52 | 0.7 | 0.94
Results from laboratory 7 were removed due to their 10 ppm reading resolution.
| Analytical results for Dy in OREAS 282 (Certified Value 2.61 ppm) | [26 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 2.6 | 2.69 | 2.7 | 2.65 | 2.52 | 2.57 | 2.5 | 2.69 | 2.5 | 2.4 | 2.6 | 2.66 | [13 empty] | 1
| 1 | 2 | 2.65 | 2.74 | 2.6 | 2.65 | 2.53 | 2.52 | 2.7 | 2.76 | 2.5 | 2.5 | 2.6 | 2.64 | [13 empty] | 30
| 1 | 3 | 2.7 | 2.65 | 2.8 | 2.65 | 2.5 | 2.7 | 2.5 | 2.82 | 2.3 | 2.4 | 2.7 | 2.64 | [13 empty] | 16
| 1 | 4 | 2.55 | 2.69 | 2.8 | 2.65 | 2.51 | 2.63 | 2.5 | 2.84 | 2.4 | 2.5 | 2.6 | 2.55 | [13 empty] | 2.6079
| 1 | 5 | 2.6 | 2.7 | 2.6 | 2.6 | 2.53 | 2.69 | 2.5 | 2.8 | 2.3 | 2.4 | 2.7 | 2.54 | [13 empty] | 24
| 1 | 6 | 2.7 | 2.76 | 2.8 | 2.65 | 2.54 | 2.62 | 2.5 | 2.81 | 2.5 | 2.5 | 2.7 | 2.56
| Mean |  | 2.6333 | 2.705 | 2.7167 | 2.6417 | 2.5217 | 2.6217 | 2.5333 | 2.7867 | 2.4167 | 2.45 | 2.65 | 2.5983
| Median |  | 2.625 | 2.695 | 2.75 | 2.65 | 2.525 | 2.625 | 2.5 | 2.805 | 2.45 | 2.45 | 2.65 | 2.6
| Std Dev. |  | 0.060553 | 0.03937 | 0.098319 | 0.020412 | 0.01472 | 0.069113 | 0.08165 | 0.054283 | 0.098319 | 0.054772 | 0.054772 | 0.053821
| Rel.Std.Dev. |  | 2.299% | 1.455% | 3.619% | 0.7727% | 0.5837% | 2.636% | 3.223% | 1.948% | 4.068% | 2.236% | 2.067% | 2.071%
| PDM3 |  | 0.9767% | 3.725% | 4.172% | 1.296% | -3.305% | 0.5294% | -2.858% | 6.856% | -7.331% | -6.053% | 1.616% | -0.3653%
| Z-Score (Absolute) |  | 0.05 | 0.63 | 0.72 | 0.11 | 0.86 | 0.05 | 0.76 | 1.29 | 1.71 | 1.44 | 0.18 | 0.24
| Analytical results for Er in OREAS 282 (Certified Value 1.46 ppm) | [26 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 1.4 | 1.46 | 1.35 | 1.45 | 1.43 | 1.6 | 1.5 | 1.43 | 1.4 | 1.4 | 1.4 | 1.66 | [13 empty] | 1
| 1 | 2 | 1.4 | 1.47 | 1.4 | 1.45 | 1.42 | 1.53 | 1.2 | 1.46 | 1.5 | 1.4 | 1.5 | 1.69 | [13 empty] | 31
| 1 | 3 | 1.45 | 1.44 | 1.4 | 1.45 | 1.42 | 1.55 | 1.3 | 1.47 | 1.3 | 1.3 | 1.5 | 1.59 | [13 empty] | 16
| 1 | 4 | 1.4 | 1.45 | 1.4 | 1.45 | 1.42 | 1.57 | 1.4 | 1.52 | 1.5 | 1.4 | 1.4 | 1.51 | [13 empty] | 1.461
| 1 | 5 | 1.4 | 1.49 | 1.5 | 1.45 | 1.45 | 1.58 | 1.3 | 1.5 | 1.3 | 1.4 | 1.4 | 1.62 | [13 empty] | 25
| 1 | 6 | 1.5 | 1.49 | 1.5 | 1.45 | 1.41 | 1.59 | 1.5 | 1.5 | 1.4 | 1.4 | 1.5 | 1.6
| Mean |  | 1.425 | 1.4667 | 1.425 | 1.45 | 1.425 | 1.57 | 1.3667 | 1.48 | 1.4 | 1.3833 | 1.45 | 1.6117
| Median |  | 1.4 | 1.465 | 1.4 | 1.45 | 1.42 | 1.575 | 1.35 | 1.485 | 1.4 | 1.4 | 1.45 | 1.61
| Std Dev. |  | 0.041833 | 0.020656 | 0.061237 | 0 | 0.013784 | 0.026077 | 0.12111 | 0.032863 | 0.089443 | 0.040825 | 0.054772 | 0.062423
| Rel.Std.Dev. |  | 2.936% | 1.408% | 4.297% | 0% | 0.9673% | 1.661% | 8.861% | 2.22% | 6.389% | 2.951% | 3.777% | 3.873%
| PDM3 |  | -2.461% | 0.3911% | -2.461% | -0.7497% | -2.461% | 7.464% | -6.454% | 1.304% | -4.172% | -5.313% | -0.7497% | 10.32%
| Z-Score (Absolute) |  | 0.67 | 0.17 | 0.67 | 0.17 | 0.67 | 2.26 | NA | 0.44 | NA | NA | NA | 3.1
Results from laboratories 16, 19, 23 and 27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 282 (Certified Value 0.65 ppm) | [26 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 0.7 | 0.68 | 0.6 | 0.65 | 0.66 | 0.67 | 0.61 | 0.67 | 0.63 | 0.6 | 0.6 | 0.66 | [13 empty] | 1
| 1 | 2 | 0.65 | 0.68 | 0.6 | 0.65 | 0.66 | 0.68 | 0.63 | 0.67 | 0.63 | 0.6 | 0.7 | 0.68 | [13 empty] | 32
| 1 | 3 | 0.7 | 0.68 | 0.6 | 0.6 | 0.66 | 0.67 | 0.6 | 0.66 | 0.6 | 0.6 | 0.7 | 0.66 | [13 empty] | 16
| 1 | 4 | 0.65 | 0.67 | 0.6 | 0.6 | 0.65 | 0.64 | 0.62 | 0.69 | 0.58 | 0.6 | 0.7 | 0.66 | [13 empty] | 0.65426
| 1 | 5 | 0.65 | 0.68 | 0.6 | 0.65 | 0.65 | 0.71 | 0.61 | 0.69 | 0.6 | 0.6 | 0.7 | 0.64 | [13 empty] | 26
| 1 | 6 | 0.65 | 0.68 | 0.6 | 0.65 | 0.67 | 0.75 | 0.63 | 0.68 | 0.63 | 0.6 | 0.7 | 0.66
| Mean |  | 0.66667 | 0.67833 | 0.6 | 0.63333 | 0.65833 | 0.68667 | 0.61667 | 0.67667 | 0.61167 | 0.6 | 0.68333 | 0.66
| Median |  | 0.65 | 0.68 | 0.6 | 0.65 | 0.66 | 0.675 | 0.615 | 0.675 | 0.615 | 0.6 | 0.7 | 0.66
| Std Dev. |  | 0.02582 | 0.0040825 | 0 | 0.02582 | 0.0075277 | 0.038297 | 0.012111 | 0.012111 | 0.02137 | 0 | 0.040825 | 0.012649
| Rel.Std.Dev. |  | 3.873% | 0.6018% | 0% | 4.077% | 1.143% | 5.577% | 1.964% | 1.79% | 3.494% | 0% | 5.974% | 1.917%
| PDM3 |  | 1.896% | 3.68% | -8.293% | -3.198% | 0.6227% | 4.953% | -5.746% | 3.425% | -6.51% | -8.293% | 4.444% | 0.8774%
| Z-Score (Absolute) |  | 0.25 | 0.67 | NA | 0.98 | 0.06 | 0.98 | 1.59 | 0.61 | 1.78 | NA | NA | 0
Results from laboratories 7, 23 and 27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 282 (Certified Value 1.06 wt.%) | [26 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.06 | 1.14 | 0.98 | 1.09 | 1.126 | 1.09 | 1.03 | 1.05 | 1.04 | 1.15 | 1.06 | 1.05 | 1.06 | 1.03 | 1.05 | 1.01 | 1.06 | 0.96 | 1.05 | 1.01 | 1.0493 | 1.07 | 1.1 | 1.11 | 1.11 | 1
| 1 | 2 | 1.07 | 1.2 | 1.01 | 1.11 | 1.112 | 1.12 | 1.015 | 1.1 | 1.05 | 1.12 | 1.12 | 1.07 | 1.07 | 0.99 | 1.05 | 0.97 | 1.05 | 0.95 | 1.06 | 1.02 | 1.0293 | 1.1 | 1.08 | 1.1 | 1.08 | #N/A
| 1 | 3 | 1.06 | 1.22 | 1.01 | 1.09 | 1.119 | 1.13 | 1.04 | 1.1 | 1.04 | 1.14 | 1.07 | 1.08 | 1.07 | 1.05 | 1.05 | 1.04 | 1.06 | 0.97 | 1.05 | 0.99 | 1.0793 | 1.09 | 1.08 | 1.1 | 1.06 | 16
| 1 | 4 | 1.09 | 1.16 | 0.98 | 1.13 | 1.105 | 1.1 | 1.02 | 1.09 | 1.04 | 1.13 | 1.04 | 1.08 | 1.06 | 1.01 | 1.05 | 1.04 | 1.06 | 0.95 | 1.06 | 0.99 | 1.0493 | 1.07 | 1.09 | 1.1 | 1.01 | 1.0641
| 1 | 5 | 1.07 | 1.17 | 1 | 1.11 | 1.133 | 1.1 | 1.02 | 1.08 | 1.03 | 1.14 | 1 | 1.06 | 1.06 | 1 | 1.07 | 0.989 | 1.05 | 0.97 | 1.04 | 1.01 | 1.0693 | 1.07 | 1.09 | 1.09 | 1.09 | 27
| 1 | 6 | 1.08 | 1.18 | 1.1 | 1.1 | 1.119 | 1.07 | 1.02 | 1.09 | 1.03 | 1.12 | 1.07 | 1.05 | 1.08 | 1.05 | 1.07 | 0.967 | 1.06 | 0.96 | 1.04 | 1.01 | 1.1118 | 1.09 | 1.09 | 1.1 | 1.1
| Mean |  | 1.0717 | 1.1783 | 1.0133 | 1.105 | 1.119 | 1.1017 | 1.0242 | 1.085 | 1.0383 | 1.1333 | 1.06 | 1.065 | 1.0667 | 1.0217 | 1.0567 | 1.0027 | 1.0567 | 0.96 | 1.05 | 1.005 | 1.0647 | 1.0817 | 1.0883 | 1.1 | 1.075
| Median |  | 1.07 | 1.175 | 1.005 | 1.105 | 1.119 | 1.1 | 1.02 | 1.09 | 1.04 | 1.135 | 1.065 | 1.065 | 1.065 | 1.02 | 1.05 | 0.9995 | 1.06 | 0.96 | 1.05 | 1.01 | 1.0593 | 1.08 | 1.09 | 1.1 | 1.085
| Std Dev. |  | 0.01169 | 0.028577 | 0.044572 | 0.015166 | 0.0098995 | 0.02137 | 0.0091742 | 0.018708 | 0.0075277 | 0.012111 | 0.039497 | 0.013784 | 0.008165 | 0.025626 | 0.010328 | 0.032764 | 0.005164 | 0.0089443 | 0.0089443 | 0.012247 | 0.028914 | 0.013292 | 0.0075277 | 0.0063246 | 0.036194
| Rel.Std.Dev. |  | 1.091% | 2.425% | 4.399% | 1.372% | 0.8847% | 1.94% | 0.8958% | 1.724% | 0.725% | 1.069% | 3.726% | 1.294% | 0.7655% | 2.508% | 0.9774% | 3.268% | 0.4887% | 0.9317% | 0.8518% | 1.219% | 2.716% | 1.229% | 0.6917% | 0.575% | 3.367%
| PDM3 |  | 0.7068% | 10.73% | -4.775% | 3.839% | 5.155% | 3.526% | -3.757% | 1.96% | -2.426% | 6.502% | -0.3895% | 0.08036% | 0.237% | -3.992% | -0.7027% | -5.777% | -0.7027% | -9.787% | -1.329% | -5.558% | 0.05373% | 1.647% | 2.273% | 3.369% | 1.02%
| Z-Score (Absolute) |  | 0.17 | 2.87 | 1.31 | 1.01 | 1.37 | 0.93 | 1.03 | 0.51 | 0.67 | 1.73 | 0.13 | 0 | 0.04 | 1.1 | 0.21 | 1.58 | 0.21 | 2.66 | 0.38 | 1.52 | 0.01 | 0.42 | 0.59 | 0.89 | 0.25
| Analytical results for Ga in OREAS 282 (Certified Value 7.15 ppm) | [26 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 6.93 | 7.4 | 6 | 7.3 | 6.8 | 10 | 6.8 | 8 | 7.22 | 6.98 | 7 | 7.11 | 6.2 | 6.8 | 7.4 | 7.4 | 8.1 | 6.8 | 8.43 | 6.61 | 4.92 | 7.54 | 7.38 | 7.93 |  | 1
| 1 | 2 | 7.15 | 7.6 | 7 | 7.7 | 6.7 | 10 | 7 | 8 | 7.05 | 7.8 | 7.1 | 7.12 | 6.1 | 7.1 | 6.9 | 7.36 | 7.9 | 7.04 | 9.18 | 6.39 | 4.96 | 7.72 | 7.29 | 7.7 |  | 33
| 1 | 3 | 7.05 | 7.4 | 7 | 7.2 | 6.6 | 10 | 6.6 | 8 | 7.03 | 7.22 | 6.9 | 7.32 | 6.2 | 6.8 | 7.1 | 7.37 | 8.1 | 6.85 | 8.48 | 6.55 | 5 | 7.29 | 7.03 | 7.87 |  | 16
| 1 | 4 | 7.07 | 7.4 | 6 | 7.7 | 6.6 | 10 | 7 | 8 | 6.88 | 6.9 | 7.1 | 7.27 | 6.2 | 7 | 6.4 | 7.46 | 8 | 6.93 | 8.79 | 6.66 | 4.93 | 7.56 | 7.15 | 7.26 |  | 7.1461
| 1 | 5 | 7.16 | 7.4 | 7 | 7.7 | 6.8 | 10 | 6.8 | 8 | 6.87 | 7.39 | 6.9 | 7.14 | 6.1 | 7.1 | 6.8 | 7.04 | 8.1 | 7.01 | 8.75 | 6.65 | 5.01 | 7.89 | 7.26 | 7.85 |  | 28
| 1 | 6 | 7 | 7.6 | 7 | 6.9 | 6.8 | 10 | 6.8 | 8 | 7.03 | 7.33 | 7.1 | 7.38 | 6.2 | 7.2 | 7.6 | 7.25 | 8.1 | 6.75 | 8.47 | 7.12 | 4.9 | 7.5 | 7.21 | 8
| Mean |  | 7.06 | 7.4667 | 6.6667 | 7.4167 | 6.7167 | 10 | 6.8333 | 8 | 7.0133 | 7.27 | 7.0167 | 7.2233 | 6.1667 | 7 | 7.0333 | 7.3133 | 8.05 | 6.8967 | 8.6833 | 6.6633 | 4.9533 | 7.5833 | 7.22 | 7.7683
| Median |  | 7.06 | 7.4 | 7 | 7.5 | 6.75 | 10 | 6.8 | 8 | 7.03 | 7.275 | 7.05 | 7.205 | 6.2 | 7.05 | 7 | 7.365 | 8.1 | 6.89 | 8.615 | 6.63 | 4.945 | 7.55 | 7.235 | 7.86
| Std Dev. |  | 0.088091 | 0.10328 | 0.5164 | 0.33714 | 0.098319 | 0 | 0.15055 | 0 | 0.12879 | 0.32323 | 0.098319 | 0.11535 | 0.05164 | 0.16733 | 0.43205 | 0.15042 | 0.083666 | 0.11622 | 0.28759 | 0.24459 | 0.044572 | 0.20403 | 0.121 | 0.26828
| Rel.Std.Dev. |  | 1.248% | 1.383% | 7.746% | 4.546% | 1.464% | 0% | 2.203% | 0% | 1.836% | 4.446% | 1.401% | 1.597% | 0.8374% | 2.39% | 6.143% | 2.057% | 1.039% | 1.685% | 3.312% | 3.671% | 0.8998% | 2.69% | 1.676% | 3.454%
| PDM3 |  | -1.205% | 4.485% | -6.71% | 3.786% | -6.01% | 39.94% | -4.377% | 11.95% | -1.858% | 1.733% | -1.812% | 1.08% | -13.71% | -2.045% | -1.579% | 2.34% | 12.65% | -3.491% | 21.51% | -6.756% | -30.69% | 6.118% | 1.034% | 8.707%
| Z-Score (Absolute) |  | 0 | 1.08 | NA | 0.95 | 0.91 | NA | 0.6 | NA | 0.12 | 0.56 | 0.12 | 0.43 | 2.38 | 0.16 | 0.07 | 0.67 | 2.64 | 0.43 | 4.32 | 1.06 | 5.61 | 1.39 | 0.43 | 1.89
Results from laboratories 3, 7 and 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 282 (Certified Value 3.05 ppm) | [26 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 3 | 2.9 | 3.2 | 3 | 2.93 | 2.94 | 2.8 | 3.4 | 3 | 2.7 | 3 | 3.14 | [13 empty] | 1
| 1 | 2 | 3 | 2.9 | 3.3 | 3 | 2.95 | 3.06 | 3 | 3.4 | 3 | 2.7 | 3 | 3.11 | [13 empty] | 34
| 1 | 3 | 3.2 | 2.9 | 3.5 | 3.2 | 2.95 | 3.08 | 3 | 3.4 | 2.9 | 2.6 | 3.1 | 3.14 | [13 empty] | 16
| 1 | 4 | 3 | 3 | 3.3 | 3.2 | 2.96 | 3 | 2.9 | 3.4 | 2.8 | 2.7 | 3.1 | 3.03 | [13 empty] | 3.0464
| 1 | 5 | 3.2 | 3 | 3.2 | 3.2 | 3 | 3.19 | 2.9 | 3.5 | 2.8 | 2.7 | 3.2 | 3.12 | [13 empty] | 29
| 1 | 6 | 3 | 3 | 3 | 3.2 | 2.95 | 3.07 | 3 | 3.4 | 3 | 2.6 | 3.2 | 3.12
| Mean |  | 3.0667 | 2.95 | 3.25 | 3.1333 | 2.9567 | 3.0567 | 2.9333 | 3.4167 | 2.9167 | 2.6667 | 3.1 | 3.11
| Median |  | 3 | 2.95 | 3.25 | 3.2 | 2.95 | 3.065 | 2.95 | 3.4 | 2.95 | 2.7 | 3.1 | 3.12
| Std Dev. |  | 0.10328 | 0.054772 | 0.16432 | 0.10328 | 0.023381 | 0.084063 | 0.08165 | 0.040825 | 0.098319 | 0.05164 | 0.089443 | 0.040988
| Rel.Std.Dev. |  | 3.368% | 1.857% | 5.056% | 3.296% | 0.7908% | 2.75% | 2.784% | 1.195% | 3.371% | 1.936% | 2.885% | 1.318%
| PDM3 |  | 0.6656% | -3.164% | 6.684% | 2.854% | -2.945% | 0.3374% | -3.711% | 12.15% | -4.258% | -12.46% | 1.76% | 2.088%
| Z-Score (Absolute) |  | 0.03 | 0.7 | 1.17 | 0.45 | 0.65 | 0.03 | 0.8 | 2.21 | 0.9 | 2.46 | 0.24 | 0.3
| Analytical results for Ge in OREAS 282 (Indicative Value 0.19 ppm) | [26 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.20 | 2.21 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.14 | < 0.1 | < 5 | 1.3 | 0.5 | < 0.05 | 0.3 | 0.1 | < 0.1 | 0.12 | 0.6 | < 0.05 | 3.05 | < 0.05 | 0.14 | 0.28 | [9 empty] | 1
| 1 | 2 | 0.14 | < 0.1 | < 5 | 1.4 | 0.4 | < 0.05 | 0.3 | 0.1 | < 0.1 | 0.12 | 0.6 | < 0.05 | 2.57 | 0.09 | 0.12 | 0.15 | [9 empty] | 35
| 1 | 3 | 0.14 | < 0.1 | < 5 | 1.3 | 0.4 | < 0.05 | 0.3 | < 0.1 | < 0.1 | 0.13 | 0.6 | < 0.05 | 2.74 | 0.11 | 0.12 | 0.18 | [9 empty] | 16
| 1 | 4 | 0.15 | < 0.1 | < 5 | 1.3 | 0.5 | < 0.05 | 0.3 | < 0.1 | < 0.1 | 0.11 | 0.6 | < 0.05 | 3.01 | 0.07 | 0.13 | 0.22 | [9 empty] | 0.19029
| 1 | 5 | 0.15 | < 0.1 | < 5 | 1.2 | 0.4 | < 0.05 | 0.3 | 0.2 | < 0.1 | 0.11 | 0.6 | < 0.05 | 3.23 | 0.06 | 0.13 | 0.17 | [9 empty] | 9
| 1 | 6 | 0.13 | < 0.1 | < 5 | 1.3 | 0.4 | < 0.05 | 0.3 | 0.1 | < 0.1 | 0.12 | 0.6 | < 0.05 | 2.4 | 0.14 | 0.14 | 0.18
| Mean |  | 0.14167 |  |  | 1.3 | 0.43333 |  | 0.3 | 0.125 |  | 0.11833 | 0.6 |  | 2.8333 | 0.094 | 0.13 | 0.19667
| Median |  | 0.14 |  |  | 1.3 | 0.4 |  | 0.3 | 0.1 |  | 0.12 | 0.6 |  | 2.875 | 0.09 | 0.13 | 0.18
| Std Dev. |  | 0.0075277 |  |  | 0.063246 | 0.05164 |  | 0 | 0.05 |  | 0.0075277 | 0 |  | 0.31665 | 0.032094 | 0.0089443 | 0.046762
| Rel.Std.Dev. |  | 5.314% |  |  | 4.865% | 11.92% |  | 0% | 40% |  | 6.361% | 0% |  | 11.18% | 34.14% | 6.88% | 23.78%
| PDM3 |  | -25.55% |  |  | 583.2% | 127.7% |  | 57.65% | -34.31% |  | -37.81% | 215.3% |  | 1389% | -50.6% | -31.68% | 3.35%
| Z-Score (Absolute) |  | 0.04 | 0.59 | 16.21 | 7.98 | 2.04 | 0.76 | 1.13 | 0.25 | 0.59 | 0.12 | 3.18 | 0.76 | 18.5 | 0.37 | 0.04 | 0.42
| Analytical results for Hf in OREAS 282 (Certified Value 1.63 ppm) | [26 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 1.7 | 1.4 | 2.31 | 1.8 | 0.5 | 1.8 | 1.6 | 1.76 | 1.73 | 1.6 | 1.59 | 1.8 | 1.41 | 1.71 | 1.7 | 1.6 | 1.3 | 1.7 | 1.1 | 1.7 | 1.7 | 1.9 |  |  |  | 1
| 1 | 2 | 1.5 | 1.6 | 2.26 | 1.8 | < 0.5 | 1.6 | 1.6 | 1.6 | 1.67 | 1.6 | 1.5 | 1.8 | 1.45 | 1.73 | 1.6 | 1.5 | 1.39 | 1.72 | 1.1 | 1.5 | 1.7 | 1.8 |  |  |  | 19
| 1 | 3 | 1.8 | 1.6 | 2.49 | 1.8 | < 0.5 | 1.4 | 1.5 | 1.47 | 1.74 | 1.6 | 1.5 | 1.9 | 1.48 | 1.6 | 1.6 | 1.6 | 1.33 | 1.65 | 1 | 1.7 | 1.8 | 1.7 |  |  |  | 16
| 1 | 4 | 1.5 | 1.6 | 2.32 | 1.9 | 0.5 | 1.8 | 1.6 | 1.57 | 1.67 | 1.6 | 1.62 | 1.8 | 1.52 | 1.48 | 1.6 | 1.5 | 1.34 | 1.64 | 1 | 1.7 | 1.7 | 1.6 |  |  |  | 1.6265
| 1 | 5 | 1.5 | 1.6 | 2.15 | 1.9 | 1 | 1.8 | 1.6 | 1.52 | 1.82 | 1.5 | 1.63 | 1.8 | 1.46 | 1.67 | 1.5 | 1.5 | 1.5 | 1.68 | 1 | 1.8 | 1.7 | 1.8 |  |  |  | 30
| 1 | 6 | 1.6 | 1.6 | 2.4 | 1.8 | < 0.5 | 1.4 | 1.5 | 1.54 | 1.88 | 1.5 | 1.57 | 1.9 | 1.49 | 1.65 | 1.5 | 1.6 | 1.34 | 1.68 | 1 | 1.8 | 1.7 | 1.8
| Mean |  | 1.6 | 1.5667 | 2.3217 | 1.8333 | 0.66667 | 1.6333 | 1.5667 | 1.5767 | 1.7517 | 1.5667 | 1.5683 | 1.8333 | 1.4683 | 1.64 | 1.5833 | 1.55 | 1.3667 | 1.6783 | 1.0333 | 1.7 | 1.7167 | 1.7667
| Median |  | 1.55 | 1.6 | 2.315 | 1.8 | 0.5 | 1.7 | 1.6 | 1.555 | 1.735 | 1.6 | 1.58 | 1.8 | 1.47 | 1.66 | 1.6 | 1.55 | 1.34 | 1.68 | 1 | 1.7 | 1.7 | 1.8
| Std Dev. |  | 0.12649 | 0.08165 | 0.11652 | 0.05164 | 0.28868 | 0.19664 | 0.05164 | 0.10013 | 0.083766 | 0.05164 | 0.057067 | 0.05164 | 0.037639 | 0.090774 | 0.075277 | 0.054772 | 0.071461 | 0.029944 | 0.05164 | 0.10954 | 0.040825 | 0.10328
| Rel.Std.Dev. |  | 7.906% | 5.212% | 5.019% | 2.817% | 43.3% | 12.04% | 3.296% | 6.351% | 4.782% | 3.296% | 3.639% | 2.817% | 2.563% | 5.535% | 4.754% | 3.534% | 5.229% | 1.784% | 4.997% | 6.444% | 2.378% | 5.846%
| PDM3 |  | -1.629% | -3.678% | 42.74% | 12.72% | -59.01% | 0.4207% | -3.678% | -3.063% | 7.696% | -3.678% | -3.576% | 12.72% | -9.724% | 0.8305% | -2.653% | -4.703% | -15.97% | 3.187% | -36.47% | 4.519% | 5.544% | 8.618%
| Z-Score (Absolute) |  | 0.06 | 0.17 | 5.05 | 1.67 | 7.84 | 0.29 | 0.17 | 0.1 | 1.11 | 0.17 | 0.16 | 1.67 | 0.85 | 0.33 | 0.06 | 0.29 | 1.56 | 0.6 | 3.86 | 0.75 | 0.86 | 1.21
| Analytical results for Hg in OREAS 282 (Indicative Value 21.7 ppm) | [26 empty] | Indicative
Hg | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.28 | [24 empty] | ppm
|  |  | 4A*OES/MS | [24 empty] | 1
1
| 1 | 1 | 19 | [24 empty] | 1
| 1 | 2 | 20 | [24 empty] | 3
| 1 | 3 | 24 | [24 empty] | 16
| 1 | 4 | 23 | [24 empty] | 21.6667
| 1 | 5 | 24 | [24 empty] | 9
| 1 | 6 | 20
| Mean |  | 21.6667
| Median |  | 21.5
| Std Dev. |  | 2.2509
| Rel.Std.Dev. |  | 10.39%
| PDM3 |  | -1.443e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 282 (Certified Value 0.51 ppm) | [26 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 0.5 | 0.49 | 0.45 | 0.52 | 0.52 | 0.55 | 0.5 | 0.52 | 0.49 | 0.5 | 0.5 | 0.55 | [13 empty] | 1
| 1 | 2 | 0.5 | 0.49 | 0.5 | 0.52 | 0.5 | 0.5 | 0.5 | 0.54 | 0.51 | 0.5 | 0.5 | 0.53 | [13 empty] | 20
| 1 | 3 | 0.52 | 0.49 | 0.45 | 0.52 | 0.5 | 0.55 | 0.5 | 0.53 | 0.46 | 0.5 | 0.5 | 0.54 | [13 empty] | 16
| 1 | 4 | 0.5 | 0.48 | 0.45 | 0.52 | 0.5 | 0.56 | 0.5 | 0.54 | 0.48 | 0.5 | 0.5 | 0.51 | [13 empty] | 0.50778
| 1 | 5 | 0.52 | 0.49 | 0.45 | 0.52 | 0.5 | 0.57 | 0.4 | 0.54 | 0.46 | 0.5 | 0.5 | 0.54 | [13 empty] | 31
| 1 | 6 | 0.52 | 0.49 | 0.45 | 0.5 | 0.51 | 0.55 | 0.5 | 0.54 | 0.5 | 0.5 | 0.5 | 0.54
| Mean |  | 0.51 | 0.48833 | 0.45833 | 0.51667 | 0.505 | 0.54667 | 0.48333 | 0.535 | 0.48333 | 0.5 | 0.5 | 0.535
| Median |  | 0.51 | 0.49 | 0.45 | 0.52 | 0.5 | 0.55 | 0.5 | 0.54 | 0.485 | 0.5 | 0.5 | 0.54
| Std Dev. |  | 0.010954 | 0.0040825 | 0.020412 | 0.008165 | 0.0083666 | 0.024221 | 0.040825 | 0.0083666 | 0.020656 | 0 | 0 | 0.013784
| Rel.Std.Dev. |  | 2.148% | 0.836% | 4.454% | 1.58% | 1.657% | 4.431% | 8.447% | 1.564% | 4.274% | 0% | 0% | 2.576%
| PDM3 |  | 0.4376% | -3.829% | -9.737% | 1.751% | -0.547% | 7.659% | -4.814% | 5.361% | -4.814% | -1.532% | -1.532% | 5.361%
| Z-Score (Absolute) |  | 0 | 0.58 | 1.39 | 0.18 | 0.13 | 0.99 | NA | 0.67 | 0.72 | NA | NA | 0.67
Results from laboratories 16, 23 and 27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 282 (Certified Value 0.059 ppm) | [26 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.06 | 0.06 | < 2 | < 0.02 | < 0.1 | 0.05 | 0.06 | 0.05 | 0.076 | 0.05 | 0.06 | < 0.1 | 0.06 | 0.06 | 0.057 | 0.07 | 0.06 | 0.06 | 0.045 | 0.061 | 0.062 | 0.062 |  |  |  | 1
| 1 | 2 | 0.065 | 0.06 | < 2 | < 0.02 | < 0.1 | 0.05 | 0.06 | 0.06 | 0.047 | 0.05 | 0.05 | < 0.1 | 0.06 | 0.05 | 0.061 | 0.05 | 0.06 | 0.06 | 0.06 | 0.056 | 0.061 | 0.064 |  |  |  | 21
| 1 | 3 | 0.066 | 0.06 | < 2 | < 0.02 | < 0.1 | 0.05 | 0.06 | 0.06 | 0.08 | 0.05 | 0.06 | < 0.1 | 0.06 | 0.06 | 0.055 | 0.06 | 0.06 | 0.06 | 0.04 | 0.057 | 0.057 | 0.06 |  |  |  | 16
| 1 | 4 | 0.062 | 0.06 | < 2 | < 0.02 | < 0.1 | 0.05 | 0.07 | 0.06 | 0.067 | 0.05 | 0.05 | < 0.1 | 0.06 | 0.06 | 0.057 | < 0.05 | 0.06 | 0.06 | 0.044 | 0.063 | 0.06 | 0.052 |  |  |  | 0.059486
| 1 | 5 | 0.061 | 0.06 | < 2 | < 0.02 | < 0.1 | 0.05 | 0.05 | 0.06 | 0.062 | 0.05 | 0.06 | < 0.1 | 0.06 | 0.06 | 0.058 | < 0.05 | 0.06 | 0.06 | 0.061 | 0.07 | 0.061 | 0.057 |  |  |  | 32
| 1 | 6 | 0.064 | 0.06 | < 2 | < 0.02 | < 0.1 | 0.05 | 0.06 | 0.06 | 0.082 | 0.05 | 0.06 | < 0.1 | 0.06 | 0.05 | 0.057 | 0.06 | 0.06 | 0.06 | 0.039 | 0.064 | 0.06 | 0.061
| Mean |  | 0.063 | 0.06 |  |  |  | 0.05 | 0.06 | 0.058333 | 0.069 | 0.05 | 0.056667 |  | 0.06 | 0.056667 | 0.0575 | 0.06 | 0.06 | 0.06 | 0.048167 | 0.061833 | 0.060167 | 0.059333
| Median |  | 0.063 | 0.06 |  |  |  | 0.05 | 0.06 | 0.06 | 0.0715 | 0.05 | 0.06 |  | 0.06 | 0.06 | 0.057 | 0.06 | 0.06 | 0.06 | 0.0445 | 0.062 | 0.0605 | 0.0605
| Std Dev. |  | 0.0023664 | 0 |  |  |  | 7.6012e-18 | 0.0063246 | 0.0040825 | 0.013236 | 7.6012e-18 | 0.005164 |  | 0 | 0.005164 | 0.0019748 | 0.008165 | 0 | 0 | 0.0098268 | 0.0051153 | 0.0017224 | 0.004274
| Rel.Std.Dev. |  | 3.756% | 0% |  |  |  | 1.52e-14% | 10.54% | 6.999% | 19.18% | 1.52e-14% | 9.113% |  | 0% | 9.113% | 3.435% | 13.61% | 0% | 0% | 20.4% | 8.273% | 2.863% | 7.203%
| PDM3 |  | 5.908% | 0.8646% |  |  |  | -15.95% | 0.8646% | -1.937% | 15.99% | -15.95% | -4.739% |  | 0.8646% | -4.739% | -3.338% | 0.8646% | 0.8646% | 0.8646% | -19.03% | 3.947% | 1.145% | -0.2562%
| Z-Score (Absolute) |  | 1.09 | 0.3 | 245.67 | 12.75 | 2.31 | 2.31 | 0.3 | 0.13 | 2.65 | 2.31 | 0.57 | 2.31 | 0.3 | 0.57 | 0.35 | 2.74 | 0.3 | 0.3 | 2.78 | 0.78 | 0.35 | 0.13
| Analytical results for K in OREAS 282 (Certified Value 1.85 wt.%) | [26 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.82 | 1.89 | 1.99 | 1.98 | 1.9093 | 1.92 | 1.8 | 1.81 | 1.94 | 1.9453 | 1.6 | 1.767 | 1.83 | 1.7 | 1.97 | 1.85 | 1.78 | 1.79 | 1.77 | 1.94 | 1.8281 | 1.82 | 1.85 | 2.01 | 1.86 | 1
| 1 | 2 | 1.83 | 1.94 | 2.01 | 1.8 | 1.901 | 1.9 | 1.76 | 1.87 | 1.99 | 1.8761 | 1.96 | 1.805 | 1.82 | 1.63 | 2.01 | 1.77 | 1.76 | 1.72 | 1.78 | 1.95 | 1.7978 | 1.84 | 1.82 | 1.99 | 1.79 | #N/A
| 1 | 3 | 1.82 | 1.93 | 2.04 | 1.99 | 1.8761 | 1.96 | 1.81 | 1.86 | 1.97 | 1.9113 | 1.68 | 1.805 | 1.87 | 1.71 | 1.98 | 1.87 | 1.76 | 1.72 | 1.77 | 1.89 | 1.8685 | 1.82 | 1.8 | 1.99 | 1.8 | 16
| 1 | 4 | 1.88 | 1.89 | 1.97 | 1.87 | 1.8844 | 1.89 | 1.79 | 1.86 | 2.01 | 1.9042 | 1.8 | 1.818 | 1.82 | 1.68 | 1.98 | 1.89 | 1.76 | 1.77 | 1.76 | 1.88 | 1.8483 | 1.81 | 1.83 | 1.98 | 1.69 | 1.8512
| 1 | 5 | 1.83 | 1.89 | 2.05 | 1.89 | 1.9343 | 1.94 | 1.79 | 1.79 | 2.01 | 1.9247 | 1.68 | 1.795 | 1.79 | 1.65 | 2.02 | 1.79 | 1.74 | 1.78 | 1.77 | 1.93 | 1.8584 | 1.81 | 1.83 | 1.99 | 1.85 | 33
| 1 | 6 | 1.83 | 1.93 | 1.99 | 1.83 | 1.8927 | 1.85 | 1.78 | 1.85 | 1.95 | 1.9187 | 1.7 | 1.775 | 1.82 | 1.65 | 2.02 | 1.8 | 1.77 | 1.78 | 1.74 | 1.9 | 1.8786 | 1.82 | 1.84 | 1.99 | 1.86
| Mean |  | 1.835 | 1.9117 | 2.0083 | 1.8933 | 1.8996 | 1.91 | 1.7883 | 1.84 | 1.9783 | 1.9134 | 1.7367 | 1.7942 | 1.825 | 1.67 | 1.9967 | 1.8283 | 1.7617 | 1.76 | 1.765 | 1.915 | 1.8466 | 1.82 | 1.8283 | 1.9917 | 1.8083
| Median |  | 1.83 | 1.91 | 2 | 1.88 | 1.8969 | 1.91 | 1.79 | 1.855 | 1.98 | 1.915 | 1.69 | 1.8 | 1.82 | 1.665 | 1.995 | 1.825 | 1.76 | 1.775 | 1.77 | 1.915 | 1.8534 | 1.82 | 1.83 | 1.99 | 1.825
| Std Dev. |  | 0.022583 | 0.024014 | 0.031252 | 0.077632 | 0.020645 | 0.038987 | 0.017224 | 0.032249 | 0.029944 | 0.023038 | 0.12675 | 0.019539 | 0.025884 | 0.031623 | 0.022509 | 0.048339 | 0.013292 | 0.031623 | 0.013784 | 0.02881 | 0.029562 | 0.010954 | 0.017224 | 0.0098319 | 0.065549
| Rel.Std.Dev. |  | 1.231% | 1.256% | 1.556% | 4.1% | 1.087% | 2.041% | 0.9631% | 1.753% | 1.514% | 1.204% | 7.299% | 1.089% | 1.418% | 1.894% | 1.127% | 2.644% | 0.7545% | 1.797% | 0.781% | 1.504% | 1.601% | 0.6019% | 0.9421% | 0.4937% | 3.625%
| PDM3 |  | -0.8769% | 3.264% | 8.486% | 2.274% | 2.614% | 3.174% | -3.398% | -0.6068% | 6.866% | 3.357% | -6.189% | -3.083% | -1.417% | -9.79% | 7.856% | -1.237% | -4.838% | -4.928% | -4.658% | 3.445% | -0.2494% | -1.687% | -1.237% | 7.586% | -2.317%
| Z-Score (Absolute) |  | 0 | 0.74 | 1.67 | 0.56 | 0.62 | 0.72 | 0.45 | 0.05 | 1.38 | 0.76 | 0.95 | 0.39 | 0.1 | 1.59 | 1.56 | 0.06 | 0.71 | 0.72 | 0.67 | 0.77 | 0.11 | 0.14 | 0.06 | 1.51 | 0.26
| Analytical results for La in OREAS 282 (Certified Value 20.5 ppm) | [26 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 20.8 | 21.5 | 17 | 21.1 | 20 | 20.2 | 19 | 20 | 21 | 20.33 | 19.5 | 19.13 | 19.8 | 18.3 | 19.1 | 20.4 | 19.5 | 23.7 | 20.1 | 20.8 | 22.1 | 24.4 | 21.5 |  |  | 1
| 1 | 2 | 21.4 | 21.6 | 18 | 23.1 | 21 | 20.9 | 19 | 20 | 21 | 19.8 | 19.1 | 18.79 | 18.6 | 18.6 | 19.2 | 20.3 | 20.1 | 23.2 | 20.2 | 20.8 | 22.1 | 24 | 21.3 |  |  | 22
| 1 | 3 | 21.6 | 22.1 | 18 | 20.7 | 20 | 21.2 | 19.5 | 20.3 | 21 | 20.05 | 19.9 | 19.4 | 20 | 18.8 | 19.7 | 20.3 | 19.2 | 24 | 20.6 | 20.9 | 21.3 | 23.8 | 21.5 |  |  | 16
| 1 | 4 | 21.3 | 21.3 | 17 | 23.2 | 21 | 20.1 | 19 | 20.4 | 21 | 20.02 | 19.1 | 19.22 | 20 | 18.9 | 19 | 20.5 | 19.5 | 23.1 | 20.7 | 19.4 | 22.1 | 23.6 | 20.1 |  |  | 20.5254
| 1 | 5 | 21.7 | 21.4 | 18 | 21.6 | 21 | 20.8 | 19 | 19.8 | 21 | 20.03 | 19.4 | 18.8 | 19.4 | 18.7 | 19.7 | 20.1 | 20.3 | 23.5 | 20.7 | 20.6 | 23 | 23.5 | 21.4 |  |  | 34
| 1 | 6 | 21.7 | 21.8 | 19 | 19.9 | 20 | 20.1 | 19.5 | 20 | 21 | 20.13 | 20.2 | 19.51 | 19.7 | 19 | 18.8 | 19.9 | 19.5 | 22.6 | 19.8 | 20.7 | 22.3 | 23.7 | 21.6
| Mean |  | 21.4167 | 21.6167 | 17.8333 | 21.6 | 20.5 | 20.55 | 19.1667 | 20.0833 | 21 | 20.06 | 19.5333 | 19.1417 | 19.5833 | 18.7167 | 19.25 | 20.25 | 19.6833 | 23.35 | 20.35 | 20.5333 | 22.15 | 23.8333 | 21.2333
| Median |  | 21.5 | 21.55 | 18 | 21.35 | 20.5 | 20.5 | 19 | 20 | 21 | 20.04 | 19.45 | 19.175 | 19.75 | 18.75 | 19.15 | 20.3 | 19.5 | 23.35 | 20.4 | 20.75 | 22.1 | 23.75 | 21.45
| Std Dev. |  | 0.34303 | 0.29269 | 0.75277 | 1.3236 | 0.54772 | 0.47645 | 0.2582 | 0.22286 | 0 | 0.17205 | 0.44121 | 0.29969 | 0.53072 | 0.24833 | 0.37283 | 0.21679 | 0.42151 | 0.49295 | 0.37283 | 0.56451 | 0.54314 | 0.3266 | 0.56451
| Rel.Std.Dev. |  | 1.602% | 1.354% | 4.221% | 6.128% | 2.672% | 2.318% | 1.347% | 1.11% | 0% | 0.8577% | 2.259% | 1.566% | 2.71% | 1.327% | 1.937% | 1.071% | 2.141% | 2.111% | 1.832% | 2.749% | 2.452% | 1.37% | 2.659%
| PDM3 |  | 4.342% | 5.316% | -13.12% | 5.235% | -0.1239% | 0.1197% | -6.62% | -2.154% | 2.312% | -2.268% | -4.834% | -6.742% | -4.59% | -8.812% | -6.214% | -1.342% | -4.103% | 13.76% | -0.8547% | 0.03848% | 7.915% | 16.12% | 3.449%
| Z-Score (Absolute) |  | 0.81 | 0.97 | 1.92 | 0.95 | 0.11 | 0.15 | 0.9 | 0.2 | 0.5 | 0.22 | 0.62 | 0.92 | 0.59 | 1.25 | 0.84 | 0.08 | 0.51 | 2.29 | 0 | 0.14 | 1.37 | 2.66 | 0.67
| Analytical results for Li in OREAS 282 (Certified Value 21.6 ppm) | [26 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 20.3 | 21.5 | 23 | 24 | 26 | 26.8 | 21 | 30 | 25.5 | 22.8 | 21.7 | 21.3 | 22.1 | 20.1 | 23 | 18.9 | 20 | 22.5 | 20.9 | 22.1 | 21.4 | 22 | 20.9 | 21.8 | 22.7 | 1
| 1 | 2 | 21.3 | 21.5 | 23 | 23 | 27 | 28.1 | 21 | 30 | 21.5 | 22.6 | 22.4 | 21.2 | 22.1 | 19.5 | 23 | 19 | 20.4 | 22.2 | 20.3 | 22.5 | 21 | 22 | 21 | 21.5 | 22.9 | #N/A
| 1 | 3 | 21.2 | 21.5 | 22 | 21 | 27 | 28.2 | 21 | 30 | 26.5 | 22.8 | 21.1 | 21.4 | 22.3 | 20.2 | 23 | 21.5 | 19.8 | 22.8 | 20.4 | 21.6 | 21.9 | 22 | 20.4 | 20.6 | 23 | 16
| 1 | 4 | 20.9 | 21 | 22 | 27 | 27 | 27.4 | 21 | 30 | 25.4 | 22.6 | 22.4 | 21.5 | 22.5 | 20 | 23 | 19.2 | 19.6 | 22.5 | 20.2 | 21.6 | 22 | 22 | 21.2 | 21.3 | 21.6 | 21.5676
| 1 | 5 | 21.2 | 21.5 | 23 | 25 | 27 | 27.4 | 21 | 30 | 27.5 | 23.4 | 20.6 | 21 | 23.2 | 19.8 | 23 | 20.3 | 19.3 | 22.3 | 20.2 | 21.9 | 21.6 | 22 | 21.3 | 21.3 | 22.9 | 35
| 1 | 6 | 21.2 | 21.5 | 23 | 21 | 27 | 27.3 | 21 | 30 | 29.3 | 22.9 | 22.2 | 20.9 | 23 | 21.6 | 23 | 18.7 | 19.5 | 22.5 | 20 | 21.9 | 23 | 22 | 20.9 | 21.6 | 23
| Mean |  | 21.0167 | 21.4167 | 22.6667 | 23.5 | 26.8333 | 27.5333 | 21 | 30 | 25.95 | 22.85 | 21.7333 | 21.2167 | 22.5333 | 20.2 | 23 | 19.6 | 19.7667 | 22.4667 | 20.3333 | 21.9333 | 21.8167 | 22 | 20.95 | 21.35 | 22.6833
| Median |  | 21.2 | 21.5 | 23 | 23.5 | 27 | 27.4 | 21 | 30 | 26 | 22.8 | 21.95 | 21.25 | 22.4 | 20.05 | 23 | 19.1 | 19.7 | 22.5 | 20.25 | 21.9 | 21.75 | 22 | 20.95 | 21.4 | 22.9
| Std Dev. |  | 0.37639 | 0.20412 | 0.5164 | 2.3452 | 0.40825 | 0.52789 | 0 | 0 | 2.6167 | 0.29496 | 0.74744 | 0.23166 | 0.46762 | 0.72938 | 0 | 1.0881 | 0.39328 | 0.20656 | 0.30768 | 0.33862 | 0.6824 | 0 | 0.31464 | 0.41352 | 0.54191
| Rel.Std.Dev. |  | 1.791% | 0.9531% | 2.278% | 9.98% | 1.521% | 1.917% | 0% | 0% | 10.08% | 1.291% | 3.439% | 1.092% | 2.075% | 3.611% | 0% | 5.552% | 1.99% | 0.9194% | 1.513% | 1.544% | 3.128% | 0% | 1.502% | 1.937% | 2.389%
| PDM3 |  | -2.555% | -0.6999% | 5.096% | 8.96% | 24.41% | 27.66% | -2.632% | 39.1% | 20.32% | 5.946% | 0.7683% | -1.627% | 4.478% | -6.341% | 6.641% | -9.123% | -8.35% | 4.169% | -5.723% | 1.696% | 1.155% | 2.005% | -2.864% | -1.009% | 5.173%
| Z-Score (Absolute) |  | 0.67 | 0.36 | 0.62 | 1.26 | 3.86 | 4.4 | 0.68 | NA | 3.17 | 0.76 | 0.11 | 0.51 | 0.51 | 1.3 | 0.88 | 1.77 | 1.64 | 0.46 | 1.2 | 0.05 | 0.05 | 0.1 | 0.72 | 0.41 | 0.63
Results from laboratory 8 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 282 (Certified Value 0.19 ppm) | [26 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.04 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 0.2 | 0.2 | 0.2 | 0.16 | 0.19 | 0.2 | 0.2 | 0.18 | 0.18 | 0.2 | 0.2 | 0.21 | [13 empty] | 1
| 1 | 2 | 0.22 | 0.19 | 0.2 | 0.16 | 0.18 | 0.19 | 0.2 | 0.19 | 0.18 | 0.2 | 0.2 | 0.21 | [13 empty] | 23
| 1 | 3 | 0.2 | 0.2 | 0.2 | 0.16 | 0.18 | 0.21 | 0.2 | 0.19 | 0.17 | 0.2 | 0.2 | 0.2 | [13 empty] | 16
| 1 | 4 | 0.22 | 0.2 | 0.2 | 0.16 | 0.19 | 0.2 | 0.2 | 0.19 | 0.18 | 0.2 | 0.2 | 0.19 | [13 empty] | 0.18917
| 1 | 5 | 0.2 | 0.18 | 0.2 | 0.16 | 0.18 | 0.2 | 0.2 | 0.19 | 0.17 | 0.2 | 0.2 | 0.19 | [13 empty] | 36
| 1 | 6 | 0.2 | 0.19 | 0.2 | 0.16 | 0.18 | 0.22 | 0.2 | 0.19 | 0.19 | 0.2 | 0.2 | 0.2
| Mean |  | 0.20667 | 0.19333 | 0.2 | 0.16 | 0.18333 | 0.20333 | 0.2 | 0.18833 | 0.17833 | 0.2 | 0.2 | 0.2
| Median |  | 0.2 | 0.195 | 0.2 | 0.16 | 0.18 | 0.2 | 0.2 | 0.19 | 0.18 | 0.2 | 0.2 | 0.2
| Std Dev. |  | 0.010328 | 0.008165 | 3.0405e-17 | 0 | 0.005164 | 0.010328 | 3.0405e-17 | 0.0040825 | 0.0075277 | 3.0405e-17 | 3.0405e-17 | 0.0089443
| Rel.Std.Dev. |  | 4.997% | 4.223% | 1.52e-14% | 0% | 2.817% | 5.079% | 1.52e-14% | 2.168% | 4.221% | 1.52e-14% | 1.52e-14% | 4.472%
| PDM3 |  | 9.251% | 2.203% | 5.727% | -15.42% | -3.084% | 7.489% | 5.727% | -0.4405% | -5.727% | 5.727% | 5.727% | 5.727%
| Z-Score (Absolute) |  | 0.99 | 0.16 | NA | 1.92 | 0.47 | 0.78 | NA | 0.16 | 0.78 | NA | NA | 0.57
Results from laboratories 7, 16, 23 and 27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 282 (Certified Value 3.39 wt.%) | [26 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.53 | 3.31 | 4.02 | 3.45 | 3.6184 | 3.58 | 3.11 | 3.36 | 3.61 | 3.3161 | 3.4 | 3.324 | 3.29 | 3.21 | 3.55 | 3.34 | 3.59 | 3.26 | 3.37 | 3.15 | 3.185 | 3.28 | 3.62 | 3.37 | 3.6 | 1
| 1 | 2 | 3.54 | 3.37 | 3.94 | 3.47 | 3.5882 | 3.58 | 3.078 | 3.41 | 3.65 | 3.212 | 3.61 | 3.39 | 3.29 | 3.13 | 3.62 | 3.21 | 3.56 | 3.31 | 3.38 | 3.17 | 3.1066 | 3.32 | 3.55 | 3.34 | 3.48 | #N/A
| 1 | 3 | 3.53 | 3.49 | 3.92 | 3.45 | 3.6304 | 3.66 | 3.143 | 3.39 | 3.67 | 3.2575 | 3.44 | 3.409 | 3.35 | 3.23 | 3.57 | 3.33 | 3.57 | 3.28 | 3.36 | 3.09 | 3.234 | 3.28 | 3.53 | 3.34 | 3.42 | 16
| 1 | 4 | 3.63 | 3.4 | 3.91 | 3.54 | 3.6003 | 3.58 | 3.1 | 3.42 | 3.7 | 3.244 | 3.33 | 3.407 | 3.3 | 3.27 | 3.6 | 3.32 | 3.57 | 3.25 | 3.37 | 3.08 | 3.234 | 3.28 | 3.58 | 3.31 | 3.27 | 3.3938
| 1 | 5 | 3.55 | 3.35 | 3.94 | 3.39 | 3.6365 | 3.63 | 3.0965 | 3.39 | 3.7 | 3.2735 | 3.21 | 3.379 | 3.29 | 3.21 | 3.65 | 3.24 | 3.54 | 3.33 | 3.39 | 3.15 | 3.2438 | 3.29 | 3.58 | 3.34 | 3.54 | 37
| 1 | 6 | 3.59 | 3.5 | 3.9 | 3.33 | 3.5943 | 3.53 | 3.1 | 3.38 | 3.66 | 3.2591 | 3.45 | 3.331 | 3.32 | 3.2 | 3.64 | 3.23 | 3.56 | 3.26 | 3.28 | 3.11 | 3.2724 | 3.29 | 3.58 | 3.33 | 3.56
| Mean |  | 3.5617 | 3.4033 | 3.9383 | 3.4383 | 3.6113 | 3.5933 | 3.1046 | 3.3917 | 3.665 | 3.2604 | 3.4067 | 3.3733 | 3.3067 | 3.2083 | 3.605 | 3.2783 | 3.565 | 3.2817 | 3.3583 | 3.125 | 3.2126 | 3.29 | 3.5733 | 3.3383 | 3.4783
| Median |  | 3.545 | 3.385 | 3.93 | 3.45 | 3.6094 | 3.58 | 3.1 | 3.39 | 3.665 | 3.2583 | 3.42 | 3.3845 | 3.295 | 3.21 | 3.61 | 3.28 | 3.565 | 3.27 | 3.37 | 3.13 | 3.234 | 3.285 | 3.58 | 3.34 | 3.51
| Std Dev. |  | 0.040208 | 0.076855 | 0.043089 | 0.071671 | 0.019969 | 0.045461 | 0.021537 | 0.02137 | 0.033912 | 0.034339 | 0.13337 | 0.037259 | 0.024221 | 0.045789 | 0.03937 | 0.057764 | 0.016432 | 0.031885 | 0.039707 | 0.036742 | 0.059098 | 0.015492 | 0.030768 | 0.019408 | 0.12007
| Rel.Std.Dev. |  | 1.129% | 2.258% | 1.094% | 2.084% | 0.553% | 1.265% | 0.6937% | 0.6301% | 0.9253% | 1.053% | 3.915% | 1.105% | 0.7325% | 1.427% | 1.092% | 1.762% | 0.4609% | 0.9716% | 1.182% | 1.176% | 1.84% | 0.4709% | 0.861% | 0.5814% | 3.452%
| PDM3 |  | 4.945% | 0.2802% | 16.04% | 1.311% | 6.409% | 5.879% | -8.523% | -0.0636% | 7.99% | -3.932% | 0.3784% | -0.6038% | -2.568% | -5.466% | 6.222% | -3.403% | 5.044% | -3.305% | -1.046% | -7.921% | -5.339% | -3.059% | 5.289% | -1.635% | 2.49%
| Z-Score (Absolute) |  | 0.87 | 0.06 | 2.81 | 0.24 | 1.13 | 1.04 | 1.47 | 0 | 1.4 | 0.67 | 0.08 | 0.09 | 0.44 | 0.94 | 1.1 | 0.58 | 0.89 | 0.56 | 0.17 | 1.37 | 0.92 | 0.52 | 0.93 | 0.27 | 0.45
| Analytical results for Mn in OREAS 282 (Certified Value 0.025 wt.%) | [26 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.0248 | 0.0274 | 0.0214 | 0.0242 | 0.0263 | 0.0254 | 0.024 | 0.0242 | 0.0255 | 0.0262 | 0.0244 | 0.0251 | 0.0256 | 0.0246 | 0.0253 | 0.0267 | 0.0251 | 0.0221 | 0.0253 | 0.0243 | 0.0237 | 0.0258 | 0.0253 | 0.0248 | 0.0254 | 1
| 1 | 2 | 0.025 | 0.0282 | 0.022 | 0.0252 | 0.0271 | 0.026 | 0.024 | 0.025 | 0.0253 | 0.0255 | 0.0261 | 0.0251 | 0.0253 | 0.0241 | 0.0257 | 0.0257 | 0.0248 | 0.023 | 0.0257 | 0.0245 | 0.0233 | 0.026 | 0.0248 | 0.0247 | 0.0247 | 24
| 1 | 3 | 0.0249 | 0.029 | 0.0227 | 0.0243 | 0.0271 | 0.026 | 0.0245 | 0.0246 | 0.0251 | 0.0258 | 0.0248 | 0.0254 | 0.0258 | 0.0246 | 0.0253 | 0.0269 | 0.0249 | 0.023 | 0.0256 | 0.0239 | 0.0242 | 0.026 | 0.025 | 0.0247 | 0.0242 | 16
| 1 | 4 | 0.0255 | 0.0292 | 0.0215 | 0.0254 | 0.0271 | 0.0254 | 0.024 | 0.0252 | 0.0255 | 0.0257 | 0.0241 | 0.0253 | 0.0254 | 0.0247 | 0.0254 | 0.0269 | 0.0249 | 0.0227 | 0.0257 | 0.0238 | 0.0238 | 0.0261 | 0.0253 | 0.0245 | 0.0232 | 0.025112
| 1 | 5 | 0.025 | 0.0284 | 0.0216 | 0.0249 | 0.0271 | 0.0255 | 0.024 | 0.0242 | 0.0259 | 0.0259 | 0.0232 | 0.0248 | 0.0252 | 0.0244 | 0.0259 | 0.0261 | 0.0247 | 0.0233 | 0.0256 | 0.0242 | 0.0243 | 0.0261 | 0.0252 | 0.0247 | 0.0248 | 38
| 1 | 6 | 0.0254 | 0.03 | 0.0225 | 0.0231 | 0.0271 | 0.0254 | 0.024 | 0.0244 | 0.028 | 0.0258 | 0.0248 | 0.0249 | 0.0261 | 0.0247 | 0.0258 | 0.0254 | 0.025 | 0.0228 | 0.025 | 0.0241 | 0.0235 | 0.026 | 0.0251 | 0.0246 | 0.0252
| Mean |  | 0.0251 | 0.0287 | 0.02195 | 0.024517 | 0.026967 | 0.025617 | 0.024083 | 0.0246 | 0.025883 | 0.025817 | 0.024567 | 0.0251 | 0.025567 | 0.024517 | 0.025567 | 0.026283 | 0.0249 | 0.022817 | 0.025483 | 0.024133 | 0.0238 | 0.026 | 0.025117 | 0.024667 | 0.024583
| Median |  | 0.025 | 0.0287 | 0.0218 | 0.0246 | 0.0271 | 0.02545 | 0.024 | 0.0245 | 0.0255 | 0.0258 | 0.0246 | 0.0251 | 0.0255 | 0.0246 | 0.02555 | 0.0264 | 0.0249 | 0.0229 | 0.0256 | 0.02415 | 0.02375 | 0.026 | 0.02515 | 0.0247 | 0.02475
| Std Dev. |  | 0.00028284 | 0.00090111 | 0.00054681 | 0.00084242 | 0.0003266 | 0.00029944 | 0.00020412 | 0.00041952 | 0.0010704 | 0.00023166 | 0.00095638 | 0.00022804 | 0.00033862 | 0.00023166 | 0.00026583 | 0.00064627 | 0.00014142 | 0.00040702 | 0.00027869 | 0.0002582 | 0.00038987 | 0.00010954 | 0.00019408 | 0.00010328 | 0.00079603
| Rel.Std.Dev. |  | 1.127% | 3.14% | 2.491% | 3.436% | 1.211% | 1.169% | 0.8476% | 1.705% | 4.135% | 0.8973% | 3.893% | 0.9085% | 1.324% | 0.9449% | 1.04% | 2.459% | 0.568% | 1.784% | 1.094% | 1.07% | 1.638% | 0.4213% | 0.7727% | 0.4187% | 3.238%
| PDM3 |  | -0.0459% | 14.29% | -12.59% | -2.369% | 7.388% | 2.012% | -4.095% | -2.037% | 3.074% | 2.808% | -2.17% | -0.0459% | 1.812% | -2.369% | 1.812% | 4.666% | -0.8423% | -9.139% | 1.481% | -3.895% | -5.223% | 3.538% | 0.02047% | -1.772% | -2.103%
| Z-Score (Absolute) |  | 0 | 4.16 | 3.64 | 0.67 | 2.17 | 0.6 | 1.18 | 0.58 | 0.91 | 0.83 | 0.62 | 0 | 0.54 | 0.67 | 0.54 | 1.37 | 0.23 | 2.64 | 0.44 | 1.12 | 1.5 | 1.04 | 0.02 | 0.5 | 0.6
| Analytical results for Mo in OREAS 282 (Certified Value 13.5 ppm) | [26 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 13 | 13.5 | 12 | 13.58 | 14.4 | 12 | 12.5 | 13.4 | 13.8 | 13.39 | 13 | 13.7 | 13.4 | 13.76 | 14.3 | 14.1 | 13.9 | 12.7 | 15.1 | 12.96 | 19.25 | 14.2 | 13.84 | 14.45 |  | 1
| 1 | 2 | 13.6 | 13.5 | 12 | 13.92 | 14.5 | 12 | 12.5 | 13.4 | 13.2 | 13.39 | 13.1 | 13.6 | 12.9 | 13.98 | 13.6 | 14.15 | 13.6 | 12.91 | 15.1 | 12.94 | 20.16 | 13.85 | 13.58 | 14.05 |  | 25
| 1 | 3 | 13.45 | 13.5 | 12 | 13.67 | 13.9 | 13 | 13 | 13.2 | 13.6 | 13.47 | 13.7 | 13.7 | 13.5 | 13.69 | 13.7 | 14.15 | 13.8 | 12.78 | 14.9 | 12.96 | 20.23 | 13.3 | 13.8 | 14.65 |  | 16
| 1 | 4 | 13.45 | 13.5 | 12 | 13.38 | 14.2 | 12 | 12.5 | 13.5 | 13.4 | 13.12 | 13.6 | 13.7 | 12.9 | 13.67 | 14 | 14.1 | 13.5 | 12.7 | 14.9 | 13.07 | 18.98 | 13.9 | 13.41 | 13.6 |  | 13.5416
| 1 | 5 | 13.45 | 13.5 | 12 | 13.61 | 14.6 | 13 | 12.5 | 13.5 | 13.6 | 13.14 | 12.8 | 13.8 | 12.9 | 13.94 | 13.6 | 13.75 | 13.6 | 12.76 | 15 | 13.07 | 20.59 | 14.2 | 13.33 | 14.7 |  | 39
| 1 | 6 | 13.25 | 13.5 | 12 | 12.88 | 14.8 | 13 | 13 | 13.3 | 13.3 | 13.62 | 12.5 | 14 | 13.4 | 13.87 | 13.9 | 14 | 13.9 | 12.57 | 14.8 | 13.42 | 19.64 | 13.8 | 13.34 | 14.65
| Mean |  | 13.3667 | 13.5 | 12 | 13.5067 | 14.4 | 12.5 | 12.6667 | 13.3833 | 13.4833 | 13.355 | 13.1167 | 13.75 | 13.1667 | 13.8183 | 13.85 | 14.0417 | 13.7167 | 12.7367 | 14.9667 | 13.07 | 19.8083 | 13.875 | 13.55 | 14.35
| Median |  | 13.45 | 13.5 | 12 | 13.595 | 14.45 | 12.5 | 12.5 | 13.4 | 13.5 | 13.39 | 13.05 | 13.7 | 13.15 | 13.815 | 13.8 | 14.1 | 13.7 | 12.73 | 14.95 | 13.015 | 19.9 | 13.875 | 13.495 | 14.55
| Std Dev. |  | 0.21134 | 0 | 0 | 0.3528 | 0.31623 | 0.54772 | 0.2582 | 0.1169 | 0.22286 | 0.19357 | 0.46224 | 0.13784 | 0.29439 | 0.13075 | 0.27386 | 0.15303 | 0.17224 | 0.11219 | 0.12111 | 0.18089 | 0.62269 | 0.33129 | 0.22786 | 0.43932
| Rel.Std.Dev. |  | 1.581% | 0% | 0% | 2.612% | 2.196% | 4.382% | 2.038% | 0.8735% | 1.653% | 1.449% | 3.524% | 1.002% | 2.236% | 0.9462% | 1.977% | 1.09% | 1.256% | 0.8808% | 0.8092% | 1.384% | 3.144% | 2.388% | 1.682% | 3.061%
| PDM3 |  | -1.292% | -0.3072% | -11.38% | -0.258% | 6.339% | -7.692% | -6.461% | -1.169% | -0.4303% | -1.378% | -3.138% | 1.539% | -2.769% | 2.044% | 2.277% | 3.693% | 1.293% | -5.944% | 10.52% | -3.483% | 46.28% | 2.462% | 0.06203% | 5.97%
| Z-Score (Absolute) |  | 0.33 | 0.06 | NA | 0.04 | 1.76 | NA | 1.74 | 0.29 | 0.09 | 0.35 | 0.83 | 0.45 | 0.73 | 0.59 | 0.65 | 1.04 | 0.38 | 1.6 | 2.9 | 0.92 | 12.67 | 0.7 | 0.04 | 1.66
Results from laboratories 3 and 7 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 282 (Certified Value 0.065 wt.%) | [26 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.06 | 0.05 | 0.31 | 0.35 | 0.067 | 0.08 | 0.06 | 0.08 | 0.065 | 0.07 | 0.062 | 0.06 | 0.06 | 0.07 | 0.071 | 0.06 | 0.059 | 0.062 | 0.07 | 0.0635 | 0.06 | 0.07 | 0.08 | 0.07 |  | 1
| 1 | 2 | 0.06 | 0.04 | 0.29 | 0.37 | 0.067 | 0.07 | 0.06 | 0.08 | 0.0627 | 0.07 | 0.062 | 0.06 | 0.06 | 0.07 | 0.0705 | 0.06 | 0.056 | 0.062 | 0.07 | 0.0619 | 0.07 | 0.07 | 0.07 | 0.07 |  | 8
| 1 | 3 | 0.06 | 0.05 | 0.3 | 0.3 | 0.067 | 0.07 | 0.06 | 0.08 | 0.0633 | 0.07 | 0.063 | 0.06 | 0.06 | 0.07 | 0.07 | 0.06 | 0.058 | 0.061 | 0.07 | 0.064 | 0.07 | 0.07 | 0.07 | 0.06 |  | 16
| 1 | 4 | 0.06 | 0.05 | 0.28 | 0.38 | 0.067 | 0.08 | 0.06 | 0.09 | 0.0637 | 0.07 | 0.064 | 0.06 | 0.06 | 0.07 | 0.075 | 0.06 | 0.058 | 0.061 | 0.07 | 0.0639 | 0.06 | 0.07 | 0.07 | 0.06 |  | 0.065404
| 1 | 5 | 0.06 | 0.05 | 0.31 | 0.35 | 0.067 | 0.08 | 0.06 | 0.09 | 0.0641 | 0.07 | 0.064 | 0.06 | 0.06 | 0.07 | 0.071 | 0.06 | 0.057 | 0.061 | 0.07 | 0.0638 | 0.07 | 0.07 | 0.07 | 0.07 |  | 40
| 1 | 6 | 0.06 | 0.05 | 0.27 | 0.29 | 0.067 | 0.08 | 0.06 | 0.08 | 0.0637 | 0.07 | 0.064 | 0.06 | 0.06 | 0.07 | 0.067 | 0.06 | 0.058 | 0.06 | 0.07 | 0.0646 | 0.07 | 0.07 | 0.07 | 0.07
| Mean |  | 0.06 | 0.048333 | 0.29333 | 0.34 | 0.067 | 0.076667 | 0.06 | 0.083333 | 0.06375 | 0.07 | 0.063167 | 0.06 | 0.06 | 0.07 | 0.07075 | 0.06 | 0.057667 | 0.061167 | 0.07 | 0.063617 | 0.066667 | 0.07 | 0.071667 | 0.066667
| Median |  | 0.06 | 0.05 | 0.295 | 0.35 | 0.067 | 0.08 | 0.06 | 0.08 | 0.0637 | 0.07 | 0.0635 | 0.06 | 0.06 | 0.07 | 0.07075 | 0.06 | 0.058 | 0.061 | 0.07 | 0.06385 | 0.07 | 0.07 | 0.07 | 0.07
| Std Dev. |  | 0 | 0.0040825 | 0.01633 | 0.036878 | 0 | 0.005164 | 0 | 0.005164 | 0.00077395 | 0 | 0.00098319 | 0 | 0 | 0 | 0.0025642 | 0 | 0.0010328 | 0.00075277 | 0 | 0.00091524 | 0.005164 | 0 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 0% | 8.447% | 5.567% | 10.85% | 0% | 6.736% | 0% | 6.197% | 1.214% | 0% | 1.557% | 0% | 0% | 0% | 3.624% | 0% | 1.791% | 1.231% | 0% | 1.439% | 7.746% | 0% | 5.696% | 7.746%
| PDM3 |  | -8.263% | -26.1% | 348.5% | 419.8% | 2.44% | 17.22% | -8.263% | 27.41% | -2.529% | 7.027% | -3.421% | -8.263% | -8.263% | 7.027% | 8.173% | -8.263% | -11.83% | -6.479% | 7.027% | -2.733% | 1.93% | 7.027% | 9.575% | 1.93%
| Z-Score (Absolute) |  | 0.86 | 2.35 | 29.11 | 35.11 | 0.01 | 1.28 | 0.86 | 2.14 | 0.37 | 0.43 | 0.45 | 0.86 | 0.86 | 0.43 | 0.52 | 0.86 | 1.16 | 0.71 | 0.43 | 0.39 | 0 | 0.43 | 0.64 | 0
| Analytical results for Nb in OREAS 282 (Certified Value 5.73 ppm) | [26 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5.6 | 5.5 | < 5 | 6.2 | 5.7 | < 1 | 6 | 5.3 | 5.91 | 5.52 | 5.4 | 5.86 | 4.1 | 5.9 | 5.9 | 5.8 | 6 | 5.39 | 6.3 | 5.33 | 3.8 | 5.9 | 5.8 | 5.9 |  | 1
| 1 | 2 | 5.7 | 5.5 | < 5 | 6.3 | 5.7 | < 1 | 5.5 | 5.1 | 5.81 | 5.26 | 5.4 | 5.8 | 3.5 | 6.2 | 5.4 | 5.8 | 5.8 | 5.56 | 6.6 | 5.28 | 3.6 | 6 | 5.7 | 5.9 |  | 26
| 1 | 3 | 5.6 | 5.5 | < 5 | 6.1 | 5.6 | < 1 | 6 | 4.8 | 5.88 | 5.41 | 5.4 | 5.9 | 3.3 | 6.1 | 5.6 | 5.8 | 5.9 | 5.45 | 6.4 | 5.31 | 3.5 | 5.7 | 5.7 | 6.1 |  | 16
| 1 | 4 | 5.7 | 5.5 | < 5 | 6.3 | 5.6 | < 1 | 5.5 | 5.2 | 5.89 | 5.18 | 5.4 | 5.86 | 4.2 | 6 | 5.8 | 5.8 | 5.7 | 5.5 | 6.4 | 5.42 | 3.8 | 5.7 | 5.8 | 5.6 |  | 5.7269
| 1 | 5 | 5.7 | 5.5 | < 5 | 6.4 | 5.6 | < 1 | 6 | 4.7 | 5.91 | 5.38 | 5.4 | 5.78 | 3.4 | 6.1 | 5.7 | 5.7 | 5.7 | 5.51 | 6.3 | 5.55 | 3.9 | 6 | 5.7 | 6.1 |  | 41
| 1 | 6 | 5.6 | 5.5 | < 5 | 5.9 | 5.8 | < 1 | 5.5 | 5.5 | 5.86 | 5.75 | 5.3 | 5.86 | 3.2 | 5.9 | 5.9 | 5.8 | 5.9 | 5.35 | 6.1 | 5.84 | 3.5 | 5.7 | 5.7 | 6.1
| Mean |  | 5.65 | 5.5 |  | 6.2 | 5.6667 |  | 5.75 | 5.1 | 5.8767 | 5.4167 | 5.3833 | 5.8433 | 3.6167 | 6.0333 | 5.7167 | 5.7833 | 5.8333 | 5.46 | 6.35 | 5.455 | 3.6833 | 5.8333 | 5.7333 | 5.95
| Median |  | 5.65 | 5.5 |  | 6.25 | 5.65 |  | 5.75 | 5.15 | 5.885 | 5.395 | 5.4 | 5.86 | 3.45 | 6.05 | 5.75 | 5.8 | 5.85 | 5.475 | 6.35 | 5.375 | 3.7 | 5.8 | 5.7 | 6
| Std Dev. |  | 0.054772 | 0 |  | 0.17889 | 0.08165 |  | 0.27386 | 0.30332 | 0.037771 | 0.20186 | 0.040825 | 0.044572 | 0.42622 | 0.12111 | 0.19408 | 0.040825 | 0.12111 | 0.078994 | 0.16432 | 0.21249 | 0.17224 | 0.15055 | 0.05164 | 0.19748
| Rel.Std.Dev. |  | 0.9694% | 0% |  | 2.885% | 1.441% |  | 4.763% | 5.947% | 0.6427% | 3.727% | 0.7584% | 0.7628% | 11.78% | 2.007% | 3.395% | 0.7059% | 2.076% | 1.447% | 2.588% | 3.895% | 4.676% | 2.581% | 0.9007% | 3.319%
| PDM3 |  | -1.343% | -3.962% |  | 8.261% | -1.052% |  | 0.4034% | -10.95% | 2.615% | -5.417% | -5.999% | 2.033% | -36.85% | 5.351% | -0.1787% | 0.9854% | 1.858% | -4.66% | 10.88% | -4.748% | -35.68% | 1.858% | 0.1123% | 3.896%
| Z-Score (Absolute) |  | 0.12 | 0.55 | 9.19 | 1.46 | 0.07 | 14.95 | 0.17 | 1.7 | 0.53 | 0.79 | 0.89 | 0.44 | 5.98 | 0.98 | 0.07 | 0.26 | 0.41 | 0.67 | 1.9 | 0.68 | 5.78 | 0.41 | 0.12 | 0.74
| Analytical results for Nd in OREAS 282 (Certified Value 17.5 ppm) | [26 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 1
2
| 1 | 1 | 17.3 | 17.3 | 18 | 18.2 | 16.68 | 17.65 | 17.1 | 17.6 | 16.3 | 16.3 | 17.5 | 18 | [13 empty] | 1
| 1 | 2 | 17.8 | 17.4 | 18.5 | 18.4 | 16.72 | 16.7 | 17 | 17.6 | 15.3 | 16.9 | 18 | 17.9 | [13 empty] | 27
| 1 | 3 | 17.7 | 17.1 | 19 | 18.2 | 16.69 | 17.79 | 17.5 | 17.7 | 16.5 | 16.5 | 18.2 | 17.5 | [13 empty] | 16
| 1 | 4 | 17.7 | 17.2 | 18.5 | 18.5 | 16.69 | 17.31 | 17.1 | 18 | 14.6 | 16.9 | 18.1 | 16.4 | [13 empty] | 17.4517
| 1 | 5 | 17.6 | 17 | 18.5 | 18.2 | 16.81 | 17.1 | 17.1 | 17.7 | 16.4 | 16.9 | 18.6 | 17.5 | [13 empty] | 42
| 1 | 6 | 17.7 | 17.5 | 19 | 17.9 | 16.79 | 17.67 | 17.3 | 18.1 | 16.1 | 16.5 | 18.6 | 17.6
| Mean |  | 17.6333 | 17.25 | 18.5833 | 18.2333 | 16.73 | 17.37 | 17.1833 | 17.7833 | 15.8667 | 16.6667 | 18.1667 | 17.4833
| Median |  | 17.7 | 17.25 | 18.5 | 18.2 | 16.705 | 17.48 | 17.1 | 17.7 | 16.2 | 16.7 | 18.15 | 17.55
| Std Dev. |  | 0.17512 | 0.18708 | 0.37639 | 0.20656 | 0.056214 | 0.41718 | 0.18348 | 0.2137 | 0.75542 | 0.26583 | 0.41312 | 0.57067
| Rel.Std.Dev. |  | 0.9931% | 1.085% | 2.025% | 1.133% | 0.336% | 2.402% | 1.068% | 1.202% | 4.761% | 1.595% | 2.274% | 3.264%
| PDM3 |  | 1.041% | -1.156% | 6.485% | 4.479% | -4.135% | -0.468% | -1.538% | 1.9% | -9.082% | -4.498% | 4.097% | 0.1815%
| Z-Score (Absolute) |  | 0.26 | 0.23 | 1.48 | 1.03 | 0.89 | 0.07 | 0.31 | 0.46 | 2 | 0.97 | 0.95 | 0.07
| Analytical results for Ni in OREAS 282 (Certified Value 21.8 ppm) | [26 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 20.9 | 26 | 18 | 22 | 28.2 | 21.5 | 25 | < 5 | 26.1 | 22.2 | 22.9 | 21 | 22.2 | 23.2 | 21 | 29 | 22.9 | 30 | 21.8 | 20.2 | 21.4 | 21 | 23.3 | 24.3 | 24.2 | 1
| 1 | 2 | 21.5 | 20 | 19 | 22 | 28.7 | 22 | 20 | < 5 | 26.1 | 21.7 | 23.6 | 21 | 22 | 21.8 | 21 | 27 | 23 | 29 | 22.3 | 20.6 | 21.4 | 21 | 22.9 | 24.2 | 24.2 | 28
| 1 | 3 | 21.6 | 28 | 19 | 20 | 28 | 21.1 | 20 | < 5 | 25.9 | 21.9 | 23.4 | 21 | 22.2 | 22.1 | 21 | 29 | 22.9 | 29 | 22 | 19.9 | 21.6 | 23 | 22.1 | 24 | 24.2 | 16
| 1 | 4 | 20.8 | 22 | 18 | 22 | 27.9 | 21.3 | 20 | 5 | 25.6 | 21.5 | 22.3 | 21 | 21.9 | 23.4 | 21 | 26 | 22.4 | 29 | 22.1 | 20.2 | 21.4 | 20 | 23 | 23.8 | 21.9 | 21.8135
| 1 | 5 | 21.3 | 24 | 19 | 21 | 28.2 | 21.7 | 20 | < 5 | 25.7 | 22.3 | 22.5 | 21 | 21.9 | 21.9 | 21 | 27 | 22.5 | 29 | 21.8 | 20 | 21.1 | 23 | 23.4 | 23.8 | 24.1 | 43
| 1 | 6 | 21.3 | 24 | 19 | 21 | 28.2 | 21.5 | 20 | < 5 | 25.4 | 21.8 | 23.7 | 21 | 22.5 | 22.4 | 21 | 28 | 22.9 | 30 | 21.4 | 20.1 | 20.5 | 21 | 22.8 | 23.9 | 24
| Mean |  | 21.2333 | 24 | 18.6667 | 21.3333 | 28.2 | 21.5167 | 20.8333 | 5 | 25.8 | 21.9 | 23.0667 | 21 | 22.1167 | 22.4667 | 21 | 27.6667 | 22.7667 | 29.3333 | 21.9 | 20.1667 | 21.2333 | 21.5 | 22.9167 | 24 | 23.7667
| Median |  | 21.3 | 24 | 19 | 21.5 | 28.2 | 21.5 | 20 | 5 | 25.8 | 21.85 | 23.15 | 21 | 22.1 | 22.25 | 21 | 27.5 | 22.9 | 29 | 21.9 | 20.15 | 21.4 | 21 | 22.95 | 23.95 | 24.15
| Std Dev. |  | 0.32042 | 2.8284 | 0.5164 | 0.8165 | 0.27568 | 0.31252 | 2.0412 |  | 0.28284 | 0.30332 | 0.58878 | 0 | 0.23166 | 0.6802 | 0 | 1.2111 | 0.25033 | 0.5164 | 0.30984 | 0.24221 | 0.39328 | 1.2247 | 0.46224 | 0.20976 | 0.91797
| Rel.Std.Dev. |  | 1.509% | 11.79% | 2.766% | 3.827% | 0.9776% | 1.452% | 9.798% |  | 1.096% | 1.385% | 2.553% | 0% | 1.047% | 3.028% | 0% | 4.377% | 1.1% | 1.76% | 1.415% | 1.201% | 1.852% | 5.696% | 2.017% | 0.874% | 3.862%
| PDM3 |  | -2.66% | 10.02% | -14.43% | -2.201% | 29.28% | -1.361% | -4.493% | -77.08% | 18.28% | 0.3965% | 5.745% | -3.729% | 1.39% | 2.994% | -3.729% | 26.83% | 4.37% | 34.47% | 0.3965% | -7.55% | -2.66% | -1.437% | 5.057% | 10.02% | 8.954%
| Z-Score (Absolute) |  | 0.44 | 1.39 | 2.14 | 0.38 | 4.18 | 0.25 | 0.71 | 12.59 | 2.59 | 0 | 0.77 | 0.6 | 0.14 | 0.38 | 0.6 | 3.82 | 0.57 | 4.93 | 0 | 1.15 | 0.44 | 0.27 | 0.67 | 1.39 | 1.24
| Analytical results for P in OREAS 282 (Certified Value 0.032 wt.%) | [26 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.034 | 0.035 | 0.0253 | 0.0306 | 0.0305 | 0.036 | 0.029 | 0.04 | 0.038 | 0.0329 | 0.0303 | 0.032 | 0.0324 | 0.03 | 0.03 | 0.0345 | 0.033 | 0.031 | 0.03 | 0.0319 | 0.0305 | 0.0311 | 0.033 | 0.0324 | 0.034 | 1
| 1 | 2 | 0.033 | 0.035 | 0.0264 | 0.0318 | 0.0301 | 0.0362 | 0.028 | 0.04 | 0.037 | 0.0319 | 0.0299 | 0.032 | 0.0337 | 0.03 | 0.03 | 0.0375 | 0.032 | 0.031 | 0.029 | 0.0315 | 0.0305 | 0.0315 | 0.031 | 0.032 | 0.033 | #N/A
| 1 | 3 | 0.033 | 0.035 | 0.0263 | 0.0313 | 0.0292 | 0.0363 | 0.029 | 0.04 | 0.036 | 0.0325 | 0.0301 | 0.032 | 0.0339 | 0.03 | 0.03 | 0.0325 | 0.031 | 0.031 | 0.03 | 0.031 | 0.0309 | 0.0305 | 0.032 | 0.0325 | 0.031 | 16
| 1 | 4 | 0.035 | 0.035 | 0.0258 | 0.032 | 0.0297 | 0.0376 | 0.029 | 0.04 | 0.036 | 0.0326 | 0.0302 | 0.031 | 0.0297 | 0.03 | 0.03 | 0.036 | 0.032 | 0.03 | 0.03 | 0.031 | 0.031 | 0.031 | 0.033 | 0.0316 | 0.031 | 0.031805
| 1 | 5 | 0.034 | 0.035 | 0.0265 | 0.0312 | 0.0297 | 0.0364 | 0.0285 | 0.04 | 0.038 | 0.0327 | 0.03 | 0.032 | 0.0319 | 0.031 | 0.03 | 0.0365 | 0.031 | 0.031 | 0.029 | 0.0311 | 0.031 | 0.031 | 0.032 | 0.0322 | 0.034 | 44
| 1 | 6 | 0.034 | 0.035 | 0.0266 | 0.0316 | 0.0319 | 0.0354 | 0.029 | 0.04 | 0.038 | 0.0323 | 0.0303 | 0.031 | 0.0307 | 0.031 | 0.03 | 0.04 | 0.032 | 0.03 | 0.029 | 0.0306 | 0.0308 | 0.0311 | 0.032 | 0.0323 | 0.033
| Mean |  | 0.033833 | 0.035 | 0.02615 | 0.031417 | 0.030183 | 0.036317 | 0.02875 | 0.04 | 0.037167 | 0.032483 | 0.030133 | 0.031667 | 0.03205 | 0.030333 | 0.03 | 0.036167 | 0.031833 | 0.030667 | 0.0295 | 0.031183 | 0.030783 | 0.031033 | 0.032167 | 0.032167 | 0.032667
| Median |  | 0.034 | 0.035 | 0.02635 | 0.03145 | 0.0299 | 0.03625 | 0.029 | 0.04 | 0.0375 | 0.03255 | 0.03015 | 0.032 | 0.03215 | 0.03 | 0.03 | 0.03625 | 0.032 | 0.031 | 0.0295 | 0.03105 | 0.03085 | 0.03105 | 0.032 | 0.03225 | 0.033
| Std Dev. |  | 0.00075277 | 0 | 0.000501 | 0.00049967 | 0.00094745 | 0.00072226 | 0.00041833 | 0 | 0.00098319 | 0.00034881 | 0.0001633 | 0.0005164 | 0.0016514 | 0.0005164 | 0 | 0.0025626 | 0.00075277 | 0.0005164 | 0.00054772 | 0.0004535 | 0.00023166 | 0.00032042 | 0.00075277 | 0.0003266 | 0.0013663
| Rel.Std.Dev. |  | 2.225% | 0% | 1.916% | 1.59% | 3.139% | 1.989% | 1.455% | 0% | 2.645% | 1.074% | 0.5419% | 1.631% | 5.152% | 1.702% | 0% | 7.085% | 2.365% | 1.684% | 1.857% | 1.454% | 0.7526% | 1.032% | 2.34% | 1.015% | 4.182%
| PDM3 |  | 6.376% | 10.04% | -17.78% | -1.222% | -5.1% | 14.18% | -9.606% | 25.77% | 16.86% | 2.132% | -5.257% | -0.4358% | 0.7695% | -4.628% | -5.676% | 13.71% | 0.08822% | -3.58% | -7.248% | -1.955% | -3.213% | -2.427% | 1.136% | 1.136% | 2.708%
| Z-Score (Absolute) |  | 1.1 | 1.69 | 2.79 | 0.13 | 0.75 | 2.35 | 1.48 | 4.21 | 2.78 | 0.41 | 0.78 | 0 | 0.19 | 0.67 | 0.84 | 2.28 | 0.08 | 0.51 | 1.1 | 0.24 | 0.45 | 0.32 | 0.25 | 0.25 | 0.51
| Analytical results for Pb in OREAS 282 (Certified Value 18.8 ppm) | [26 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 18.9 | 18 | 18 | 18.3 | 18.2 | 21.1 | 17 | 16 | 20.2 | 19.3 | 17.3 | 18 | 19 | 19.9 | 17 | 21 | 18.5 | 19 | 17.76 | 16.4 | 20.39 | 20.7 | 19.7 | 20 | 20.6 | 1
| 1 | 2 | 19.3 | 18 | 18 | 19.2 | 18.1 | 24.7 | 16 | 16 | 19.5 | 18.6 | 17.2 | 18 | 18.1 | 19.2 | 17.2 | 18 | 17.8 | 19 | 18.29 | 16.9 | 19.7 | 20 | 19.3 | 19.5 | 20.8 | 29
| 1 | 3 | 19 | 18 | 17 | 18 | 18 | 23.9 | 17 | 16 | 20 | 19.1 | 17.5 | 18 | 18.6 | 19.8 | 17.1 | 20 | 18.4 | 20 | 17.9 | 16.3 | 19.01 | 20.7 | 19 | 19.2 | 21 | 16
| 1 | 4 | 19.3 | 18 | 18 | 18.6 | 18.1 | 23.4 | 17 | 17 | 19.8 | 18.9 | 17 | 18 | 18.8 | 19.6 | 17 | 19 | 18.5 | 19 | 18.4 | 16.2 | 19.11 | 19.7 | 19.5 | 19.4 | 19.6 | 18.7597
| 1 | 5 | 19.6 | 18 | 18 | 18.5 | 18.1 | 23.1 | 17 | 16 | 19.8 | 18.9 | 17.7 | 18 | 18.6 | 19.1 | 17.2 | 20 | 18.4 | 19 | 18.22 | 17.4 | 19.6 | 19.4 | 20 | 19.5 | 21.1 | 45
| 1 | 6 | 19.3 | 18 | 18 | 17.8 | 18.3 | 22.3 | 17 | 17 | 19.9 | 19 | 17.9 | 18 | 20 | 19.8 | 17.4 | 19 | 18.2 | 20 | 18.03 | 16.3 | 18.6 | 20.7 | 19.5 | 19.7 | 21
| Mean |  | 19.2333 | 18 | 17.8333 | 18.4 | 18.1333 | 23.0833 | 16.8333 | 16.3333 | 19.8667 | 18.9667 | 17.4333 | 18 | 18.85 | 19.5667 | 17.15 | 19.5 | 18.3 | 19.3333 | 18.1 | 16.5833 | 19.4017 | 20.2 | 19.5 | 19.55 | 20.6833
| Median |  | 19.3 | 18 | 18 | 18.4 | 18.1 | 23.25 | 17 | 16 | 19.85 | 18.95 | 17.4 | 18 | 18.7 | 19.7 | 17.15 | 19.5 | 18.4 | 19 | 18.125 | 16.35 | 19.355 | 20.35 | 19.5 | 19.5 | 20.9
| Std Dev. |  | 0.25033 | 0 | 0.40825 | 0.49396 | 0.10328 | 1.2592 | 0.40825 | 0.5164 | 0.23381 | 0.23381 | 0.33267 | 0 | 0.63797 | 0.33862 | 0.15166 | 1.0488 | 0.26833 | 0.5164 | 0.24536 | 0.47081 | 0.63028 | 0.57966 | 0.34059 | 0.27386 | 0.56006
| Rel.Std.Dev. |  | 1.302% | 0% | 2.289% | 2.685% | 0.5696% | 5.455% | 2.425% | 3.162% | 1.177% | 1.233% | 1.908% | 0% | 3.384% | 1.731% | 0.8843% | 5.379% | 1.466% | 2.671% | 1.356% | 2.839% | 3.249% | 2.87% | 1.747% | 1.401% | 2.708%
| PDM3 |  | 2.525% | -4.05% | -4.938% | -1.918% | -3.339% | 23.05% | -10.27% | -12.93% | 5.901% | 1.103% | -7.07% | -4.05% | 0.4812% | 4.301% | -8.581% | 3.946% | -2.451% | 3.058% | -3.517% | -11.6% | 3.422% | 7.678% | 3.946% | 4.213% | 10.25%
| Z-Score (Absolute) |  | 0.22 | NA | 0.92 | 0.46 | 0.67 | 3.33 | NA | NA | 0.73 | 0 | 1.24 | NA | 0.09 | 0.49 | 1.47 | NA | 0.54 | NA | 0.7 | 1.93 | 0.35 | 1 | 0.43 | 0.47 | 1.39
Results from laboratories 2, 7, 8, 13, 19 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 282 (Certified Value 4.6 ppm) | [26 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 4.68 | 4.5 | 4.8 | 4.85 | 4.53 | 4.78 | 4.5 | 4.67 | 4.25 | 4.3 | 4.6 | 4.54 | [13 empty] | 1
| 1 | 2 | 4.7 | 4.6 | 4.7 | 4.85 | 4.44 | 4.62 | 4.4 | 4.71 | 4.39 | 4.5 | 4.7 | 4.55 | [13 empty] | 30
| 1 | 3 | 4.8 | 4.5 | 4.7 | 4.7 | 4.51 | 4.79 | 4.5 | 4.67 | 4.39 | 4.4 | 4.6 | 4.59 | [13 empty] | 16
| 1 | 4 | 4.68 | 4.6 | 4.7 | 4.95 | 4.5 | 4.61 | 4.4 | 4.84 | 4.22 | 4.4 | 4.6 | 4.21 | [13 empty] | 4.6041
| 1 | 5 | 4.7 | 4.6 | 4.8 | 4.85 | 4.56 | 4.92 | 4 | 4.76 | 4.29 | 4.5 | 4.8 | 4.52 | [13 empty] | 46
| 1 | 6 | 4.76 | 4.6 | 4.7 | 4.75 | 4.52 | 4.92 | 4.4 | 4.77 | 4.24 | 4.4 | 4.8 | 4.54
| Mean |  | 4.72 | 4.5667 | 4.7333 | 4.825 | 4.51 | 4.7733 | 4.3667 | 4.7367 | 4.2967 | 4.4167 | 4.6833 | 4.4917
| Median |  | 4.7 | 4.6 | 4.7 | 4.85 | 4.515 | 4.785 | 4.4 | 4.735 | 4.27 | 4.4 | 4.65 | 4.54
| Std Dev. |  | 0.04899 | 0.05164 | 0.05164 | 0.088034 | 0.04 | 0.13677 | 0.18619 | 0.066232 | 0.075807 | 0.075277 | 0.098319 | 0.13992
| Rel.Std.Dev. |  | 1.038% | 1.131% | 1.091% | 1.825% | 0.8869% | 2.865% | 4.264% | 1.398% | 1.764% | 1.704% | 2.099% | 3.115%
| PDM3 |  | 2.516% | -0.8139% | 2.806% | 4.797% | -2.045% | 3.675% | -5.158% | 2.878% | -6.678% | -4.072% | 1.72% | -2.443%
| Z-Score (Absolute) |  | 0.52 | 0.32 | 0.59 | 1.09 | 0.62 | 0.81 | 1.4 | 0.61 | 1.78 | 1.13 | 0.32 | 0.72
| Analytical results for Rb in OREAS 282 (Certified Value 126 ppm) | [26 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 126.5 | 124 | 126.9 | 126 | 122 | 130 | 129.4 | 122.4 | 113.37 | 115.1 | 133.72 | 115 | 133 | 128 | 118.5 | 124 | 120.4 | 145 | 116.6 | 121.8 | 131 | 131.6 | 130 |  |  | 1
| 1 | 2 | 128.5 | 129 | 129.9 | 125 | 121.5 | 125 | 129.1 | 122 | 130.35 | 115.6 | 133.44 | 109 | 133 | 122 | 119 | 126 | 126.7 | 138 | 113.7 | 125.3 | 131 | 131.3 | 129.5 |  |  | 31
| 1 | 3 | 127 | 131 | 127.9 | 124 | 122.5 | 125 | 127.4 | 122.1 | 117.42 | 114 | 135.22 | 116 | 133 | 118 | 118 | 125 | 127.7 | 139 | 118.6 | 125.7 | 125.5 | 128.1 | 127.5 |  |  | 16
| 1 | 4 | 130 | 126 | 131 | 124 | 121 | 132 | 130.6 | 122.07 | 131.29 | 115.9 | 133.29 | 114 | 132 | 120 | 121 | 124 | 127 | 141 | 118.6 | 110.5 | 127 | 129.6 | 121 |  |  | 126.2468
| 1 | 5 | 128 | 128 | 130.7 | 124 | 122 | 146 | 130.1 | 122.35 | 129.75 | 113.9 | 130.07 | 112 | 134 | 118 | 118 | 122 | 127.1 | 143 | 119.6 | 127.3 | 129 | 131.5 | 128 |  |  | 47
| 1 | 6 | 129 | 130 | 126.2 | 128 | 121 | 131 | 127.3 | 122.77 | 120.75 | 115.1 | 131.07 | 115 | 133 | 131 | 117 | 124 | 123.9 | 141 | 125.4 | 124.4 | 128 | 131.4 | 129
| Mean |  | 128.1667 | 128 | 128.7667 | 125.1667 | 121.6667 | 131.5 | 128.9833 | 122.2817 | 123.8217 | 114.9333 | 132.8017 | 113.5 | 133 | 122.8333 | 118.5833 | 124.1667 | 125.4667 | 141.1667 | 118.75 | 122.5 | 128.5833 | 130.5833 | 127.5
| Median |  | 128.25 | 128.5 | 128.9 | 124.5 | 121.75 | 130.5 | 129.25 | 122.225 | 125.25 | 115.1 | 133.365 | 114.5 | 133 | 121 | 118.25 | 124 | 126.85 | 141 | 118.6 | 124.85 | 128.5 | 131.35 | 128.5
| Std Dev. |  | 1.291 | 2.6077 | 2.0412 | 1.6021 | 0.60553 | 7.7136 | 1.3703 | 0.28785 | 7.6576 | 0.82138 | 1.8867 | 2.5884 | 0.63246 | 5.4559 | 1.3571 | 1.3292 | 2.8161 | 2.5626 | 3.876 | 6.1517 | 2.2004 | 1.4275 | 3.3166
| Rel.Std.Dev. |  | 1.007% | 2.037% | 1.585% | 1.28% | 0.4977% | 5.866% | 1.062% | 0.2354% | 6.184% | 0.7147% | 1.421% | 2.281% | 0.4755% | 4.442% | 1.144% | 1.07% | 2.245% | 1.815% | 3.264% | 5.022% | 1.711% | 1.093% | 2.601%
| PDM3 |  | 1.521% | 1.389% | 1.996% | -0.8555% | -3.628% | 4.161% | 2.168% | -3.141% | -1.921% | -8.961% | 5.192% | -10.1% | 5.349% | -2.704% | -6.07% | -1.648% | -0.6179% | 11.82% | -5.938% | -2.968% | 1.851% | 3.435% | 0.9927%
| Z-Score (Absolute) |  | 0.55 | 0.52 | 0.67 | 0.06 | 0.78 | 1.23 | 0.72 | 0.65 | 0.34 | 2.15 | 1.5 | 2.45 | 1.54 | 0.54 | 1.41 | 0.27 | 0 | 3.21 | 1.37 | 0.61 | 0.64 | 1.05 | 0.42
| Analytical results for Re in OREAS 282 (Certified Value 0.022 ppm) | [26 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | 0.023 | < 0.1 | < 0.05 | < 0.1 | 0.03 | 0.021 | 0.021 | < 0.05 | 0.023 | 0.01 | 0.022 | < 0.05 | 0.02 | 0.02 | 0.015 | 0.023 | 0.025 | 0.028 | [7 empty] | 1
| 1 | 2 | 0.024 | < 0.1 | < 0.05 | < 0.1 | 0.026 | 0.019 | 0.027 | < 0.05 | 0.02 | 0.01 | 0.02 | < 0.05 | 0.023 | 0.02 | 0.012 | 0.023 | 0.023 | 0.023 | [7 empty] | 32
| 1 | 3 | 0.025 | < 0.1 | < 0.05 | < 0.1 | 0.033 | 0.018 | 0.019 | < 0.05 | 0.022 | 0.01 | 0.02 | < 0.05 | 0.021 | 0.02 | 0.014 | 0.021 | 0.024 | 0.021 | [7 empty] | 16
| 1 | 4 | 0.024 | < 0.1 | < 0.05 | < 0.1 | 0.025 | 0.022 | 0.024 | < 0.05 | 0.022 | 0.01 | 0.02 | < 0.05 | 0.023 | 0.021 | 0.019 | 0.019 | 0.022 | 0.023 | [7 empty] | 0.022087
| 1 | 5 | 0.022 | < 0.1 | < 0.05 | < 0.1 | 0.031 | 0.019 | 0.024 | < 0.05 | 0.021 | 0.01 | 0.023 | < 0.05 | 0.019 | 0.023 | 0.015 | 0.024 | 0.026 | 0.028 | [7 empty] | 48
| 1 | 6 | 0.023 | < 0.1 | < 0.05 | < 0.1 | 0.028 | 0.022 | 0.016 | < 0.05 | 0.022 | 0.01 | 0.023 | < 0.05 | 0.02 | 0.02 | 0.014 | 0.027 | 0.025 | 0.022
| Mean |  | 0.0235 |  |  |  | 0.028833 | 0.020167 | 0.021833 |  | 0.021667 | 0.01 | 0.021333 |  | 0.021 | 0.020667 | 0.014833 | 0.022833 | 0.024167 | 0.024167
| Median |  | 0.0235 |  |  |  | 0.029 | 0.02 | 0.0225 |  | 0.022 | 0.01 | 0.021 |  | 0.0205 | 0.02 | 0.0145 | 0.023 | 0.0245 | 0.023
| Std Dev. |  | 0.0010488 |  |  |  | 0.0030605 | 0.0017224 | 0.0039707 |  | 0.0010328 | 0 | 0.0015055 |  | 0.0016733 | 0.0012111 | 0.0023166 | 0.0027142 | 0.001472 | 0.0030605
| Rel.Std.Dev. |  | 4.463% |  |  |  | 10.61% | 8.541% | 18.19% |  | 4.767% | 0% | 7.057% |  | 7.968% | 5.86% | 15.62% | 11.89% | 6.091% | 12.66%
| PDM3 |  | 6.399% |  |  |  | 30.55% | -8.693% | -1.147% |  | -1.902% | -54.72% | -3.411% |  | -4.92% | -6.429% | -32.84% | 3.381% | 9.417% | 9.417%
| Z-Score (Absolute) |  | 0.12 | 9.87 | 0.67 | 9.87 | 2.08 | 1.1 | 0.49 | 0.67 | 0.55 | 4.84 | 0.67 | 0.67 | 0.8 | 0.92 | 3.07 | 0.12 | 0.37 | 0.37
| Analytical results for S in OREAS 282 (Certified Value 0.717 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.71 | 0.715 | 0.71 | 0.76 | 1.08 | 0.67 | 0.73 | 0.78 | 0.7653 | 0.7 | 0.673 | 0.68 | 0.74 | 0.75 | 0.838 | 0.7 | 0.668 | 0.63 | 0.651 | 0.68 | 0.73 | 0.73 | 0.74 |  |  | 1
| 1 | 2 | 0.72 | 0.71 | 0.72 | 0.79 | 1.1 | 0.67 | 0.755 | 0.79 | 0.7577 | 0.76 | 0.682 | 0.67 | 0.75 | 0.75 | 0.868 | 0.7 | 0.666 | 0.64 | 0.644 | 0.69 | 0.71 | 0.72 | 0.71 |  |  | 17
| 1 | 3 | 0.72 | 0.71 | 0.73 | 0.78 | 1.1 | 0.68 | 0.755 | 0.78 | 0.7592 | 0.69 | 0.643 | 0.66 | 0.75 | 0.77 | 0.774 | 0.7 | 0.663 | 0.64 | 0.669 | 0.69 | 0.71 | 0.72 | 0.7 |  |  | 16
| 1 | 4 | 0.73 | 0.715 | 0.71 | 0.76 | 1.07 | 0.67 | 0.755 | 0.81 | 0.7544 | 0.74 | 0.671 | 0.66 | 0.75 | 0.76 | 0.824 | 0.7 | 0.666 | 0.65 | 0.657 | 0.67 | 0.72 | 0.71 | 0.67 |  |  | 0.71732
| 1 | 5 | 0.72 | 0.705 | 0.73 | 0.78 | 1.08 | 0.675 | 0.72 | 0.8 | 0.7632 | 0.71 | 0.64 | 0.65 | 0.76 | 0.77 | 0.848 | 0.7 | 0.667 | 0.64 | 0.67 | 0.68 | 0.71 | 0.72 | 0.72 |  |  | 49
| 1 | 6 | 0.72 | 0.7 | 0.73 | 0.77 | 1.11 | 0.67 | 0.73 | 0.79 | 0.7575 | 0.7 | 0.627 | 0.68 | 0.75 | 0.75 | 0.877 | 0.7 | 0.669 | 0.62 | 0.6916 | 0.68 | 0.72 | 0.72 | 0.73
| Mean |  | 0.72 | 0.70917 | 0.72167 | 0.77333 | 1.09 | 0.6725 | 0.74083 | 0.79167 | 0.75955 | 0.71667 | 0.656 | 0.66667 | 0.75 | 0.75833 | 0.83817 | 0.7 | 0.6665 | 0.63667 | 0.66377 | 0.68167 | 0.71667 | 0.72 | 0.71167
| Median |  | 0.72 | 0.71 | 0.725 | 0.775 | 1.09 | 0.67 | 0.7425 | 0.79 | 0.75845 | 0.705 | 0.657 | 0.665 | 0.75 | 0.755 | 0.843 | 0.7 | 0.6665 | 0.64 | 0.663 | 0.68 | 0.715 | 0.72 | 0.715
| Std Dev. |  | 0.0063246 | 0.0058452 | 0.0098319 | 0.012111 | 0.015492 | 0.0041833 | 0.015943 | 0.01169 | 0.0040163 | 0.027325 | 0.022163 | 0.012111 | 0.0063246 | 0.0098319 | 0.036913 | 1.2162e-16 | 0.0020736 | 0.010328 | 0.016973 | 0.0075277 | 0.008165 | 0.0063246 | 0.024833
| Rel.Std.Dev. |  | 0.8784% | 0.8242% | 1.362% | 1.566% | 1.421% | 0.6221% | 2.152% | 1.477% | 0.5288% | 3.813% | 3.379% | 1.817% | 0.8433% | 1.297% | 4.404% | 1.737e-14% | 0.3111% | 1.622% | 2.557% | 1.104% | 1.139% | 0.8784% | 3.489%
| PDM3 |  | 0.3736% | -1.137% | 0.606% | 7.809% | 51.95% | -6.248% | 3.278% | 10.36% | 5.887% | -0.09106% | -8.548% | -7.061% | 4.556% | 5.718% | 16.85% | -2.415% | -7.085% | -11.24% | -7.466% | -4.97% | -0.09106% | 0.3736% | -0.7881%
| Z-Score (Absolute) |  | 0.03 | 0.15 | 0.06 | 0.91 | 6.17 | 0.76 | 0.37 | 1.22 | 0.68 | 0.03 | 1.04 | 0.86 | 0.53 | 0.66 | 1.99 | NA | 0.86 | 1.36 | 0.91 | 0.61 | 0.03 | 0.03 | 0.11
Results from laboratory 20 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Sb in OREAS 282 (Certified Value 15.2 ppm) | [26 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 14.9 | 14.6 | 14 | 14.3 | 17 | 16.9 | 11.7 | 15.2 | 14.59 | 14.7 | 16.77 | 12.6 | 15.15 | 14.8 | 15.05 | 14.9 | 14.18 | 17.4 | 14.63 | 16.79 | 15.1 | 15.2 | 15.4 |  |  | 1
| 1 | 2 | 15.55 | 14.8 | 14 | 13.6 | 17 | 15.2 | 11.4 | 14.87 | 14.81 | 14.7 | 17.97 | 12.4 | 15.48 | 14.9 | 15.05 | 14.9 | 14.34 | 16.6 | 14.74 | 16.42 | 14.95 | 15 | 15.7 |  |  | 33
| 1 | 3 | 18.35 | 14.8 | 15 | 14.5 | 17 | 15.7 | 11.5 | 14.93 | 14.49 | 14.8 | 16.93 | 12.8 | 15.45 | 15.3 | 15.05 | 14.4 | 14.31 | 17.1 | 14.74 | 16.89 | 14.35 | 15 | 15.8 |  |  | 16
| 1 | 4 | 15.65 | 14.4 | 14 | 14.5 | 17 | 16.3 | 11.7 | 14.89 | 14.33 | 14.8 | 16.4 | 12.7 | 14.93 | 14.8 | 15.15 | 14.8 | 14.1 | 16.5 | 15.07 | 15.95 | 14.75 | 15 | 14.5 |  |  | 15.2068
| 1 | 5 | 15.45 | 14.6 | 14 | 14.5 | 17 | 17.5 | 12 | 14.97 | 14.95 | 14.4 | 16.33 | 12.6 | 15.63 | 15.4 | 14.8 | 14.3 | 14.11 | 16.7 | 15.51 | 16.71 | 15 | 14.7 | 15.65 |  |  | 50
| 1 | 6 | 15.25 | 14.8 | 14 | 14.1 | 16.5 | 15 | 12.2 | 14.8 | 15.06 | 14.5 | 16.73 | 12.8 | 15.38 | 14.7 | 14.95 | 15 | 13.93 | 15.4 | 15.35 | 16.43 | 14.9 | 14.8 | 16.15
| Mean |  | 15.8583 | 14.6667 | 14.1667 | 14.25 | 16.9167 | 16.1 | 11.75 | 14.9433 | 14.705 | 14.65 | 16.855 | 12.65 | 15.3367 | 14.9833 | 15.0083 | 14.7167 | 14.1617 | 16.6167 | 15.0067 | 16.5317 | 14.8417 | 14.95 | 15.5333
| Median |  | 15.5 | 14.7 | 14 | 14.4 | 17 | 16 | 11.7 | 14.91 | 14.7 | 14.7 | 16.75 | 12.65 | 15.415 | 14.85 | 15.05 | 14.85 | 14.145 | 16.65 | 14.905 | 16.57 | 14.925 | 15 | 15.675
| Std Dev. |  | 1.2492 | 0.1633 | 0.40825 | 0.35637 | 0.20412 | 0.98184 | 0.30166 | 0.13823 | 0.28183 | 0.16432 | 0.59234 | 0.15166 | 0.25359 | 0.29269 | 0.12007 | 0.29269 | 0.15118 | 0.68532 | 0.36313 | 0.34295 | 0.26724 | 0.17607 | 0.56184
| Rel.Std.Dev. |  | 7.877% | 1.113% | 2.882% | 2.501% | 1.207% | 6.098% | 2.567% | 0.925% | 1.917% | 1.122% | 3.514% | 1.199% | 1.653% | 1.953% | 0.8% | 1.989% | 1.068% | 4.124% | 2.42% | 2.075% | 1.801% | 1.178% | 3.617%
| PDM3 |  | 4.285% | -3.552% | -6.84% | -6.292% | 11.24% | 5.874% | -22.73% | -1.732% | -3.3% | -3.661% | 10.84% | -16.81% | 0.8542% | -1.469% | -1.305% | -3.223% | -6.873% | 9.271% | -1.316% | 8.713% | -2.401% | -1.689% | 2.147%
| Z-Score (Absolute) |  | 0.94 | NA | 0.84 | 0.75 | 2.05 | 1.19 | 3.38 | 0.02 | 0.27 | 0.33 | 1.98 | 2.43 | 0.39 | 0.02 | 0.04 | 0.26 | 0.85 | 1.73 | 0.04 | 1.64 | 0.13 | 0.02 | 0.6
Results from laboratory 2 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 282 (Certified Value 3.94 ppm) | [26 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.6 | 4 | < 5 | 3.8 | 4.1 | 3.8 | 4 | 3 | 4.7 | 4 | 3.3 | 4 | 4.3 | 4 | 4.2 | 4 | 3.9 | 3 | 3.7 | 3.7 | 3.8 | 4.4 | 4.1 | 4 | 4 | 1
| 1 | 2 | 3.8 | 5 | < 5 | 3.9 | 4.1 | 3.8 | 4 | 3 | 4.6 | 3.9 | 3.2 | 4 | 4.1 | 4 | 4.3 | 3.7 | 4 | 3 | 3.7 | 3.7 | 3.7 | 4.4 | 4 | 4 | 3.9 | 34
| 1 | 3 | 3.7 | 5 | < 5 | 3.8 | 4 | 3.9 | 4 | 3 | 4.7 | 3.7 | 4.4 | 4 | 4 | 4 | 4.3 | 4.4 | 4 | 3 | 3.7 | 3.6 | 3.8 | 4.6 | 3.9 | 3.8 | 4 | 16
| 1 | 4 | 3.7 | 4 | < 5 | 4 | 4 | 3.8 | 4 | 3 | 4.6 | 3.8 | 3.8 | 4 | 4 | 4 | 4.3 | 3.8 | 4 | 3 | 3.7 | 3.6 | 3.8 | 4.2 | 4 | 4 | 3.7 | 3.9361
| 1 | 5 | 3.9 | 4 | < 5 | 4.1 | 4.2 | 3.9 | 4 | 3 | 4.7 | 4 | 3.7 | 4 | 3.9 | 4 | 4.3 | 4.2 | 4 | 3 | 3.7 | 3.6 | 3.7 | 4.6 | 4.1 | 4 | 4 | 51
| 1 | 6 | 3.7 | 5 | < 5 | 3.9 | 4.1 | 3.7 | 3 | 3 | 4.6 | 3.9 | 3.8 | 4 | 3.8 | 4 | 4.3 | 4 | 3.9 | 3 | 3.6 | 3.6 | 3.9 | 4.3 | 4.1 | 4 | 4
| Mean |  | 3.7333 | 4.5 |  | 3.9167 | 4.0833 | 3.8167 | 3.8333 | 3 | 4.65 | 3.8833 | 3.7 | 4 | 4.0167 | 4 | 4.2833 | 4.0167 | 3.9667 | 3 | 3.6833 | 3.6333 | 3.7833 | 4.4167 | 4.0333 | 3.9667 | 3.9333
| Median |  | 3.7 | 4.5 |  | 3.9 | 4.1 | 3.8 | 4 | 3 | 4.65 | 3.9 | 3.75 | 4 | 4 | 4 | 4.3 | 4 | 4 | 3 | 3.7 | 3.6 | 3.8 | 4.4 | 4.05 | 4 | 4
| Std Dev. |  | 0.10328 | 0.54772 |  | 0.1169 | 0.075277 | 0.075277 | 0.40825 | 0 | 0.054772 | 0.1169 | 0.42895 | 0 | 0.17224 | 0 | 0.040825 | 0.25626 | 0.05164 | 0 | 0.040825 | 0.05164 | 0.075277 | 0.16021 | 0.08165 | 0.08165 | 0.12111
| Rel.Std.Dev. |  | 2.766% | 12.17% |  | 2.985% | 1.844% | 1.972% | 10.65% | 0% | 1.178% | 3.01% | 11.59% | 0% | 4.288% | 0% | 0.9531% | 6.38% | 1.302% | 0% | 1.108% | 1.421% | 1.99% | 3.627% | 2.024% | 2.058% | 3.079%
| PDM3 |  | -5.151% | 14.33% |  | -0.4932% | 3.741% | -3.034% | -2.61% | -23.78% | 18.14% | -1.34% | -5.998% | 1.624% | 2.047% | 1.624% | 8.822% | 2.047% | 0.7771% | -23.78% | -6.421% | -7.692% | -3.881% | 12.21% | 2.471% | 0.7771% | -0.06974%
| Z-Score (Absolute) |  | 0.9 | NA | 6.44 | 0.07 | 0.67 | 0.52 | NA | NA | 3.22 | 0.22 | 1.05 | NA | 0.37 | NA | 1.57 | 0.37 | 0.15 | NA | 1.12 | 1.35 | 0.67 | 2.17 | 0.45 | 0.15 | 0
Results from laboratories 2, 7, 8, 13, 16 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 282 (Certified Value 2.35 ppm) | [26 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 3 | < 5 | < 5 | < 2 | < 2 | 10 | < 5 | 1.9 | < 0.5 | 2 | 2.3 | 2 | 3 | 3 | 2 | 4 | 2.3 | 3.1 | 2.4 | < 10 | 2 | 2 | 2 |  |  | 1
| 1 | 2 | 3 | < 5 | < 5 | < 2 | 4.5 | 15 | < 5 | 2.2 | 1.2 | 2 | 2.5 | 1.5 | 3 | 3 | 3 | 3 | 2.3 | 2.9 | 2.1 | < 10 | 1 | 2 | 2 |  |  | 35
| 1 | 3 | 3 | < 5 | < 5 | < 2 | < 2 | 10 | < 5 | 2.1 | 3.2 | 2 | 2.1 | 2.1 | 3 | 2 | 2 | 4 | 2.3 | 3.1 | 2.7 | < 10 | 1 | 2 | 2 |  |  | 16
| 1 | 4 | 3 | < 5 | < 5 | < 2 | < 2 | 10 | < 5 | 2.1 | < 0.5 | 2 | 2.2 | 2 | 3 | 2 | 2 | 4 | 2.4 | 3.1 | 2.1 | < 10 | 2 | 2 | 1 |  |  | 2.3538
| 1 | 5 | 3 | < 5 | < 5 | < 2 | 5.9 | 10 | < 5 | 1.8 | < 0.5 | 2 | 2.2 | 1.6 | 3 | 2 | 2 | 3 | 2.3 | 2.8 | 2.9 | < 10 | 2 | 2 | 2 |  |  | 52
| 1 | 6 | 3 | < 5 | < 5 | < 2 | 3.2 | 10 | < 5 | 2.2 | 2.7 | 2 | 2.3 | 1.9 | 3 | 2 | 2 | 4 | 2.4 | 3.1 | 3.1 | < 10 | 1 | 2 | 2
| Mean |  | 3 |  |  |  | 4.5333 | 10.8333 |  | 2.05 | 2.3667 | 2 | 2.2667 | 1.85 | 3 | 2.3333 | 2.1667 | 3.6667 | 2.3333 | 3.0167 | 2.55 |  | 1.5 | 2 | 1.8333
| Median |  | 3 |  |  |  | 4.5 | 10 |  | 2.1 | 2.7 | 2 | 2.25 | 1.95 | 3 | 2 | 2 | 4 | 2.3 | 3.1 | 2.55 |  | 1.5 | 2 | 2
| Std Dev. |  | 0 |  |  |  | 1.3503 | 2.0412 |  | 0.16432 | 1.0408 | 0 | 0.13663 | 0.2429 | 0 | 0.5164 | 0.40825 | 0.5164 | 0.05164 | 0.13292 | 0.41833 |  | 0.54772 | 0 | 0.40825
| Rel.Std.Dev. |  | 0% |  |  |  | 29.79% | 18.84% |  | 8.015% | 43.98% | 0% | 6.028% | 13.13% | 0% | 22.13% | 18.84% | 14.08% | 2.213% | 4.406% | 16.41% |  | 36.51% | 0% | 22.27%
| PDM3 |  | 27.45% |  |  |  | 92.6% | 360.2% |  | -12.91% | 0.5462% | -15.03% | -3.702% | -21.4% | 27.45% | -0.8699% | -7.951% | 55.78% | -0.8699% | 28.16% | 8.335% |  | -36.27% | -15.03% | -22.11%
| Z-Score (Absolute) |  | NA | 0.22 | 0.22 | 3.82 | 0.94 | NA | 0.22 | 0.99 | 2.99 | NA | 0.4 | 1.53 | NA | NA | NA | NA | 0.22 | 1.62 | 0.36 | NA | NA | NA | NA
Results from laboratories 1, 7, 13, 18, 19, 20, 21, 28, 29, 30 and 31 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 282 (Certified Value 3.33 ppm) | [26 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [13 empty] | 2
3
| 1 | 1 | 3.4 | 3.3 | 3.3 | 3.45 | 3.37 | 3.29 | 2.9 | 3.3 | 3.2 | 3.1 | 3.2 | 3.42 | [13 empty] | 1
| 1 | 2 | 3.6 | 3.3 | 3.4 | 3.45 | 3.24 | 3.23 | 2.4 | 3.3 | 3.2 | 3.2 | 3.3 | 3.42 | [13 empty] | 19
| 1 | 3 | 3.55 | 3.4 | 3.4 | 3.55 | 3.39 | 3.37 | 3.3 | 3.4 | 3.1 | 3.1 | 3.3 | 3.35 | [13 empty] | 16
| 1 | 4 | 3.55 | 3.3 | 3.4 | 3.55 | 3.18 | 3.21 | 3.2 | 3.4 | 3 | 3.2 | 3.1 | 3.22 | [13 empty] | 3.3267
| 1 | 5 | 3.5 | 3.4 | 3.3 | 3.45 | 3.31 | 3.49 | 2.4 | 3.4 | 3.2 | 3.2 | 3.3 | 3.39 | [13 empty] | 53
| 1 | 6 | 3.65 | 3.4 | 3.4 | 3.55 | 3.27 | 3.29 | 3.3 | 3.3 | 3.1 | 3.1 | 3.4 | 3.3
| Mean |  | 3.5417 | 3.35 | 3.3667 | 3.5 | 3.2933 | 3.3133 | 2.9167 | 3.35 | 3.1333 | 3.15 | 3.2667 | 3.35
| Median |  | 3.55 | 3.35 | 3.4 | 3.5 | 3.29 | 3.29 | 3.05 | 3.35 | 3.15 | 3.15 | 3.3 | 3.37
| Std Dev. |  | 0.08612 | 0.054772 | 0.05164 | 0.054772 | 0.079666 | 0.10309 | 0.42622 | 0.054772 | 0.08165 | 0.054772 | 0.10328 | 0.078486
| Rel.Std.Dev. |  | 2.432% | 1.635% | 1.534% | 1.565% | 2.419% | 3.111% | 14.61% | 1.635% | 2.606% | 1.739% | 3.162% | 2.343%
| PDM3 |  | 6.463% | 0.7014% | 1.202% | 5.21% | -1.002% | -0.4008% | -12.32% | 0.7014% | -5.812% | -5.311% | -1.804% | 0.7014%
| Z-Score (Absolute) |  | 2.74 | 0.24 | 0.46 | 2.2 | 0.5 | 0.24 | 5.42 | 0.24 | 2.59 | 2.37 | 0.85 | 0.24
| Analytical results for Sn in OREAS 282 (Certified Value 1.46 ppm) | [26 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.5 | 1 | < 10 | 1.4 | 1.35 | 1.5 | 1.5 | 1.6 | 1.5 | 1.4 | 1.5 | 1.5 | 2 | 1.3 | 1.5 | 1.4 | 1.8 | 1.3 | 1.3 | 1.7 | 1.4 | 1.6 | 1.5 |  |  | 1
| 1 | 2 | 1.5 | 1 | < 10 | 1.3 | 1.35 | 1 | 1.6 | 1.3 | 1.5 | 1.5 | 1.4 | 1.5 | 1 | 1.4 | 1.6 | 1.4 | 1.9 | 1.4 | 1.3 | 1.7 | 1.5 | 1.5 | 1.6 |  |  | 20
| 1 | 3 | 1.6 | 1 | < 10 | 1.4 | 1.36 | 1.5 | 1.8 | 1.3 | 1.5 | 1.5 | 1.3 | 1.5 | 1 | 1.3 | 1.6 | 1.4 | 1.9 | 1.4 | 1.4 | 1.6 | 1.4 | 1.5 | 1.5 |  |  | 16
| 1 | 4 | 1.5 | 1 | < 10 | 1.3 | 1.35 | 1.5 | 1.6 | 1.3 | 1.5 | 1.4 | 1.5 | 1.4 | 2 | 1.3 | 1.6 | 1.5 | 1.9 | 1.4 | 1.3 | 1.6 | 1.5 | 1.6 | 1.4 |  |  | 1.4607
| 1 | 5 | 1.5 | 1 | < 10 | 1.4 | 1.34 | 1.5 | 1.6 | 1.4 | 1.5 | 1.5 | 1.4 | 1.5 | 1 | 1.3 | 1.6 | 1.4 | 1.8 | 1.4 | 1.4 | 1.8 | 1.5 | 1.5 | 1.5 |  |  | 54
| 1 | 6 | 1.5 | 1 | < 10 | 1.4 | 1.37 | 1.5 | 1.8 | 1.3 | 1.4 | 1.4 | 1.4 | 1.5 | 1 | 1.3 | 1.5 | 1.5 | 1.8 | 1.3 | 1.4 | 1.9 | 1.5 | 1.5 | 1.5
| Mean |  | 1.5167 | 1 |  | 1.3667 | 1.3533 | 1.4167 | 1.65 | 1.3667 | 1.4833 | 1.45 | 1.4167 | 1.4833 | 1.3333 | 1.3167 | 1.5667 | 1.4333 | 1.85 | 1.3667 | 1.35 | 1.7167 | 1.4667 | 1.5333 | 1.5
| Median |  | 1.5 | 1 |  | 1.4 | 1.35 | 1.5 | 1.6 | 1.3 | 1.5 | 1.45 | 1.4 | 1.5 | 1 | 1.3 | 1.6 | 1.4 | 1.85 | 1.4 | 1.35 | 1.7 | 1.5 | 1.5 | 1.5
| Std Dev. |  | 0.040825 | 0 |  | 0.05164 | 0.010328 | 0.20412 | 0.12247 | 0.12111 | 0.040825 | 0.054772 | 0.075277 | 0.040825 | 0.5164 | 0.040825 | 0.05164 | 0.05164 | 0.054772 | 0.05164 | 0.054772 | 0.1169 | 0.05164 | 0.05164 | 0.063246
| Rel.Std.Dev. |  | 2.692% | 0% |  | 3.779% | 0.7631% | 14.41% | 7.423% | 8.861% | 2.752% | 3.777% | 5.314% | 2.752% | 38.73% | 3.101% | 3.296% | 3.603% | 2.961% | 3.779% | 4.057% | 6.81% | 3.521% | 3.368% | 4.216%
| PDM3 |  | 3.831% | -31.54% |  | -6.438% | -7.35% | -3.015% | 12.96% | -6.438% | 1.549% | -0.7326% | -3.015% | 1.549% | -8.72% | -9.861% | 7.254% | -1.874% | 26.65% | -6.438% | -7.579% | 17.52% | 0.4084% | 4.972% | 2.69%
| Z-Score (Absolute) |  | 0.34 | NA | 23.83 | 0.67 | 0.76 | 0.34 | 1.24 | 0.67 | 0.11 | 0.11 | 0.34 | 0.11 | NA | 1.01 | 0.67 | 0.22 | 2.58 | 0.67 | 0.79 | 1.69 | 0 | 0.45 | 0.22
Results from laboratories 2 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 282 (Certified Value 277 ppm) | [26 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 311 | 281 | 220 | 270.4 | 278 | 266 | 270 | 281 | 268 | 280.5 | 274 | 261.7 | 264.6 | 280 | 289 | 292 | 299 | 290 | 267 | 267 | 265 | 270.8 | 310 | 287.2 | 304 | 1
| 1 | 2 | 312 | 285 | 220 | 274.8 | 276 | 278 | 265 | 286 | 268 | 269.52 | 272 | 258.1 | 262.01 | 253 | 296 | 286 | 302 | 284 | 273 | 268 | 258 | 277.5 | 304 | 283.7 | 293 | 21
| 1 | 3 | 310 | 286 | 216 | 271.9 | 274 | 277 | 270 | 270 | 266 | 274.97 | 273.6 | 258.9 | 264.33 | 268 | 290 | 280 | 303 | 277 | 271 | 261 | 267 | 279.7 | 303 | 283.8 | 288 | 16
| 1 | 4 | 319 | 282 | 213 | 274.8 | 272 | 272 | 270 | 287 | 266 | 279.3 | 268 | 260.8 | 263.12 | 266 | 290 | 283 | 305 | 288 | 271 | 262 | 270 | 268.5 | 308 | 282.7 | 274 | 277.4968
| 1 | 5 | 313 | 298 | 220 | 274 | 275 | 272 | 270 | 279 | 263 | 275.11 | 274.8 | 255.8 | 259.42 | 265 | 297 | 275 | 299 | 280 | 271 | 265 | 268 | 282.7 | 308 | 283.1 | 299 | 55
| 1 | 6 | 317 | 295 | 230 | 272.8 | 274 | 261 | 270 | 281 | 266 | 276.64 | 276 | 255.5 | 265.64 | 269 | 295 | 294 | 302 | 286 | 268 | 261 | 260 | 276.8 | 306 | 284 | 300
| Mean |  | 313.6667 | 287.8333 | 219.8333 | 273.1167 | 274.8333 | 271 | 269.1667 | 280.6667 | 266.1667 | 276.0067 | 273.0667 | 258.4667 | 263.1867 | 266.8333 | 292.8333 | 285 | 301.6667 | 284.1667 | 270.1667 | 264 | 264.6667 | 276 | 306.5 | 284.0833 | 293
| Median |  | 312.5 | 285.5 | 220 | 273.4 | 274.5 | 272 | 270 | 281 | 266 | 275.875 | 273.8 | 258.5 | 263.725 | 267 | 292.5 | 284.5 | 302 | 285 | 271 | 263.5 | 266 | 277.15 | 307 | 283.75 | 296
| Std Dev. |  | 3.559 | 7.0261 | 5.7417 | 1.7555 | 2.0412 | 6.5115 | 2.0412 | 6.0882 | 1.8348 | 3.8831 | 2.8133 | 2.5351 | 2.2316 | 8.6583 | 3.5449 | 7.2111 | 2.3381 | 4.916 | 2.2286 | 3.0984 | 4.7188 | 5.3807 | 2.6646 | 1.6018 | 10.8812
| Rel.Std.Dev. |  | 1.135% | 2.441% | 2.612% | 0.6428% | 0.7427% | 2.403% | 0.7584% | 2.169% | 0.6894% | 1.407% | 1.03% | 0.9808% | 0.8479% | 3.245% | 1.211% | 2.53% | 0.7751% | 1.73% | 0.8249% | 1.174% | 1.783% | 1.95% | 0.8694% | 0.5638% | 3.714%
| PDM3 |  | 13.03% | 3.725% | -20.78% | -1.578% | -0.9598% | -2.341% | -3.002% | 1.142% | -4.083% | -0.537% | -1.596% | -6.858% | -5.157% | -3.843% | 5.527% | 2.704% | 8.71% | 2.404% | -2.642% | -4.864% | -4.624% | -0.5394% | 10.45% | 2.374% | 5.587%
| Z-Score (Absolute) |  | 2.81 | 0.94 | 3.97 | 0.12 | 0 | 0.28 | 0.41 | 0.42 | 0.63 | 0.08 | 0.13 | 1.18 | 0.84 | 0.58 | 1.3 | 0.73 | 1.94 | 0.67 | 0.34 | 0.78 | 0.73 | 0.08 | 2.29 | 0.67 | 1.31
| Analytical results for Ta in OREAS 282 (Certified Value 0.4 ppm) | [26 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.38 | 0.6 | < 5 | 0.66 | 0.34 | 0.5 | 0.5 | 0.4 | 0.42 | 0.39 | 0.39 | 0.38 | < 0.1 | 0.35 | 0.41 | 0.43 | 0.45 | 0.4 | 0.56 | 0.4 | 0.69 | 0.43 | 0.42 | 0.45 |  | 1
| 1 | 2 | 0.37 | 0.6 | < 5 | 0.71 | 0.34 | < 0.5 | 0.5 | 0.4 | 0.41 | 0.38 | 0.39 | 0.36 | < 0.1 | 0.39 | 0.41 | 0.43 | 0.43 | 0.4 | 0.56 | 0.4 | 0.66 | 0.4 | 0.42 | 0.45 |  | 22
| 1 | 3 | 0.39 | 0.6 | < 5 | 0.78 | 0.34 | 0.5 | 0.5 | 0.4 | 0.41 | 0.37 | 0.39 | 0.37 | < 0.1 | 0.39 | 0.39 | 0.43 | 0.42 | 0.4 | 0.56 | 0.4 | 0.82 | 0.4 | 0.42 | 0.43 |  | 16
| 1 | 4 | 0.37 | 0.6 | < 5 | 0.71 | 0.34 | < 0.5 | 0.4 | 0.4 | 0.41 | 0.38 | 0.39 | 0.39 | < 0.1 | 0.36 | 0.42 | 0.43 | 0.44 | 0.4 | 0.56 | 0.4 | 0.75 | 0.38 | 0.43 | 0.41 |  | 0.39859
| 1 | 5 | 0.38 | 0.6 | < 5 | 0.52 | 0.34 | 0.5 | 0.5 | 0.4 | 0.4 | 0.37 | 0.39 | 0.37 | < 0.1 | 0.4 | 0.38 | 0.43 | 0.47 | 0.4 | 0.54 | 0.4 | 0.75 | 0.4 | 0.41 | 0.43 |  | 56
| 1 | 6 | 0.38 | 0.6 | < 5 | 0.52 | 0.36 | 0.5 | 0.5 | 0.4 | 0.41 | 0.39 | 0.39 | 0.4 | < 0.1 | 0.35 | 0.38 | 0.44 | 0.43 | 0.4 | 0.52 | 0.5 | 0.61 | 0.42 | 0.41 | 0.44
| Mean |  | 0.37833 | 0.6 |  | 0.65 | 0.34333 | 0.5 | 0.48333 | 0.4 | 0.41 | 0.38 | 0.39 | 0.37833 |  | 0.37333 | 0.39833 | 0.43167 | 0.44 | 0.4 | 0.55 | 0.41667 | 0.71333 | 0.405 | 0.41833 | 0.435
| Median |  | 0.38 | 0.6 |  | 0.685 | 0.34 | 0.5 | 0.5 | 0.4 | 0.41 | 0.38 | 0.39 | 0.375 |  | 0.375 | 0.4 | 0.43 | 0.435 | 0.4 | 0.56 | 0.4 | 0.72 | 0.4 | 0.42 | 0.435
| Std Dev. |  | 0.0075277 | 0 |  | 0.1077 | 0.008165 | 0 | 0.040825 | 6.0809e-17 | 0.0063246 | 0.0089443 | 6.0809e-17 | 0.01472 |  | 0.022509 | 0.017224 | 0.0040825 | 0.017889 | 6.0809e-17 | 0.016733 | 0.040825 | 0.075011 | 0.017607 | 0.0075277 | 0.015166
| Rel.Std.Dev. |  | 1.99% | 0% |  | 16.57% | 2.378% | 0% | 8.447% | 1.52e-14% | 1.543% | 2.354% | 1.559e-14% | 3.891% |  | 6.029% | 4.324% | 0.9457% | 4.066% | 1.52e-14% | 3.042% | 9.798% | 10.52% | 4.347% | 1.799% | 3.486%
| PDM3 |  | -5.082% | 50.53% |  | 63.07% | -13.86% | 25.44% | 21.26% | 0.3538% | 2.863% | -4.664% | -2.155% | -5.082% |  | -6.336% | -0.06433% | 8.298% | 10.39% | 0.3538% | 37.99% | 4.535% | 78.96% | 1.608% | 4.953% | 9.135%
| Z-Score (Absolute) |  | 0.67 | NA | 48.38 | 5.61 | 1.48 | NA | NA | NA | 0.06 | 0.64 | 0.4 | 0.67 | 8.27 | 0.79 | 0.21 | 0.56 | 0.75 | NA | 3.29 | NA | 7.07 | 0.06 | 0.25 | 0.64
Results from laboratories 2, 7, 8, 9, 23 and 27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 282 (Certified Value 0.44 ppm) | [26 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 0.44 | 0.48 | 0.41 | 0.45 | 0.44 | 0.42 | 0.47 | 0.4 | 0.44 | 0.44 | 0.4 | 0.4 | 0.49 | [12 empty] | 1
| 1 | 2 | 0.44 | 0.51 | 0.4 | 0.4 | 0.46 | 0.4 | 0.46 | 0.4 | 0.46 | 0.46 | 0.4 | 0.4 | 0.48 | [12 empty] | 23
| 1 | 3 | 0.42 | 0.47 | 0.4 | 0.4 | 0.46 | 0.43 | 0.45 | 0.4 | 0.45 | 0.43 | 0.4 | 0.5 | 0.48 | [12 empty] | 16
| 1 | 4 | 0.42 | 0.48 | 0.4 | 0.45 | 0.46 | 0.42 | 0.45 | 0.4 | 0.47 | 0.43 | 0.4 | 0.5 | 0.46 | [12 empty] | 0.4445
| 1 | 5 | 0.42 | 0.49 | 0.4 | 0.4 | 0.44 | 0.41 | 0.49 | 0.4 | 0.47 | 0.43 | 0.4 | 0.5 | 0.49 | [12 empty] | 57
| 1 | 6 | 0.44 | 0.47 | 0.41 | 0.4 | 0.44 | 0.41 | 0.47 | 0.4 | 0.47 | 0.45 | 0.4 | 0.5 | 0.49
| Mean |  | 0.43 | 0.48333 | 0.40333 | 0.41667 | 0.45 | 0.415 | 0.465 | 0.4 | 0.46 | 0.44 | 0.4 | 0.46667 | 0.48167
| Median |  | 0.43 | 0.48 | 0.4 | 0.4 | 0.45 | 0.415 | 0.465 | 0.4 | 0.465 | 0.435 | 0.4 | 0.5 | 0.485
| Std Dev. |  | 0.010954 | 0.015055 | 0.005164 | 0.02582 | 0.010954 | 0.010488 | 0.015166 | 6.0809e-17 | 0.012649 | 0.012649 | 6.0809e-17 | 0.05164 | 0.01169
| Rel.Std.Dev. |  | 2.548% | 3.115% | 1.28% | 6.197% | 2.434% | 2.527% | 3.261% | 1.52e-14% | 2.75% | 2.875% | 1.52e-14% | 11.07% | 2.427%
| PDM3 |  | -3.262% | 8.736% | -9.261% | -6.262% | 1.237% | -6.637% | 4.612% | -10.01% | 3.487% | -1.012% | -10.01% | 4.987% | 8.361%
| Z-Score (Absolute) |  | 0.42 | 1.07 | 1.16 | 0.79 | 0.14 | 0.84 | 0.56 | NA | 0.42 | 0.14 | NA | NA | 1.02
Results from laboratories 16, 23 and 27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 282 (Certified Value 0.23 ppm) | [26 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 0.23 | 0.2 | < 5 | 0.24 | < 1 | 0.2 | 0.3 | 0.3 | < 0.01 | 0.2 | 0.3 | 0.3 | 0.25 | 0.2 | 0.24 | 0.2 | 0.22 | 0.21 | 0.21 | 0.32 | 0.2 | 0.27 |  |  |  | 1
| 1 | 2 | 0.25 | 0.2 | < 5 | 0.23 | < 1 | < 0.2 | 0.3 | 0.3 | 0.25 | 0.2 | 0.3 | 0.2 | 0.23 | 0.2 | 0.22 | 0.3 | 0.23 | 0.22 | 0.18 | 0.3 | 0.33 | 0.21 |  |  |  | 24
| 1 | 3 | 0.24 | 0.2 | < 5 | 0.26 | < 1 | 0.2 | 0.5 | 0.2 | 0.33 | 0.3 | 0.3 | 0.2 | 0.23 | 0.2 | 0.23 | 0.2 | 0.24 | 0.26 | 0.18 | 0.24 | 0.25 | 0.26 |  |  |  | 16
| 1 | 4 | 0.25 | 0.2 | < 5 | 0.25 | < 1 | 0.2 | 0.3 | 0.2 | 0.39 | 0.2 | 0.3 | 0.3 | 0.26 | 0.2 | 0.22 | 0.3 | 0.21 | 0.21 | 0.17 | 0.23 | 0.2 | 0.22 |  |  |  | 0.2347
| 1 | 5 | 0.26 | 0.2 | < 5 | 0.23 | < 1 | < 0.2 | 0.5 | 0.2 | 0.42 | 0.2 | 0.3 | 0.3 | 0.25 | 0.1 | 0.25 | 0.3 | 0.24 | 0.25 | 0.2 | 0.29 | 0.18 | 0.27 |  |  |  | 58
| 1 | 6 | 0.25 | 0.2 | < 5 | 0.24 | < 1 | < 0.2 | 0.5 | 0.3 | 0.24 | 0.2 | 0.3 | 0.2 | 0.22 | 0.2 | 0.23 | 0.3 | 0.2 | 0.21 | 0.21 | 0.28 | 0.36 | 0.19
| Mean |  | 0.24667 | 0.2 |  | 0.24167 |  | 0.2 | 0.4 | 0.25 | 0.326 | 0.21667 | 0.3 | 0.25 | 0.24 | 0.18333 | 0.23167 | 0.26667 | 0.22333 | 0.22667 | 0.19167 | 0.27667 | 0.25333 | 0.23667
| Median |  | 0.25 | 0.2 |  | 0.24 |  | 0.2 | 0.4 | 0.25 | 0.33 | 0.2 | 0.3 | 0.25 | 0.24 | 0.2 | 0.23 | 0.3 | 0.225 | 0.215 | 0.19 | 0.285 | 0.225 | 0.24
| Std Dev. |  | 0.010328 | 3.0405e-17 |  | 0.01169 |  | 3.3993e-17 | 0.10954 | 0.054772 | 0.080808 | 0.040825 | 0 | 0.054772 | 0.015492 | 0.040825 | 0.01169 | 0.05164 | 0.01633 | 0.022509 | 0.017224 | 0.035024 | 0.075277 | 0.034448
| Rel.Std.Dev. |  | 4.187% | 1.52e-14% |  | 4.837% |  | 1.7e-14% | 27.39% | 21.91% | 24.79% | 18.84% | 0% | 21.91% | 6.455% | 22.27% | 5.046% | 19.36% | 7.312% | 9.931% | 8.986% | 12.66% | 29.71% | 14.56%
| PDM3 |  | 5.099% | -14.78% |  | 2.968% |  | -14.78% | 70.43% | 6.519% | 38.9% | -7.684% | 27.82% | 6.519% | 2.258% | -21.89% | -1.292% | 13.62% | -4.843% | -3.423% | -18.34% | 17.88% | 7.939% | 0.8379%
| Z-Score (Absolute) |  | 0.22 | NA | 101.52 | 0 | 11.61 | NA | NA | NA | 1.39 | NA | NA | NA | 0.07 | NA | 0.45 | NA | 0.82 | 0.67 | 2.25 | 1.57 | 0.52 | 0.22
Results from laboratories 2, 8, 9, 10, 13, 15, 16, 19 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 282 (Certified Value 5.36 ppm) | [26 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.16 | 5.1 | 6.1 | 5.7 | 5.3 | 4.8 | 5.4 | 5.56 | 5.3 | 5.42 | 5.02 | 5.4 | 5.7 | 5.23 | 5.34 | 5.64 | 4.9 | 5.1 | 5.1 | 5.58 | 5.78 | 6.2 | 5.33 |  |  | 1
| 1 | 2 | 5.29 | 5.4 | 6.5 | 5.72 | 5.3 | 4.8 | 5.4 | 5.4 | 5.2 | 5.39 | 5.06 | 5.2 | 5.7 | 5.15 | 5.41 | 5.57 | 5.1 | 5.1 | 5 | 5.45 | 5.62 | 6.2 | 5.35 |  |  | 25
| 1 | 3 | 5.2 | 5.3 | 6 | 5.63 | 5.3 | 4.4 | 5.4 | 5.5 | 5.4 | 5.32 | 5.17 | 5.3 | 5.7 | 5.5 | 5.47 | 5.48 | 5 | 5.1 | 5.3 | 5.37 | 5.51 | 6.1 | 5.55 |  |  | 16
| 1 | 4 | 5.29 | 5.2 | 6.5 | 5.69 | 4.9 | 4.7 | 5.4 | 5.49 | 5.2 | 5.41 | 5.15 | 5.2 | 5.7 | 5.05 | 5.45 | 5.62 | 5.1 | 5 | 5.2 | 5.49 | 5.66 | 6.1 | 5.12 |  |  | 5.3621
| 1 | 5 | 5.32 | 5.2 | 6.4 | 5.71 | 5.2 | 4.8 | 5.4 | 5.5 | 5.4 | 5.33 | 5.05 | 5.2 | 5.8 | 5.48 | 5.46 | 5.44 | 5.1 | 5 | 5.2 | 5.33 | 5.85 | 6.1 | 5.55 |  |  | 59
| 1 | 6 | 5.31 | 5.4 | 6.3 | 5.78 | 5.1 | 4.2 | 5.5 | 5.51 | 5.4 | 5.31 | 5.17 | 5.4 | 5.8 | 5.3 | 5.43 | 5.68 | 5 | 4.7 | 5.4 | 5.47 | 5.71 | 6.1 | 5.51
| Mean |  | 5.2617 | 5.2667 | 6.3 | 5.705 | 5.1833 | 4.6167 | 5.4167 | 5.4933 | 5.3167 | 5.3633 | 5.1033 | 5.2833 | 5.7333 | 5.285 | 5.4267 | 5.5717 | 5.0333 | 5 | 5.2 | 5.4483 | 5.6883 | 6.1333 | 5.4017
| Median |  | 5.29 | 5.25 | 6.35 | 5.705 | 5.25 | 4.75 | 5.4 | 5.5 | 5.35 | 5.36 | 5.105 | 5.25 | 5.7 | 5.265 | 5.44 | 5.595 | 5.05 | 5.05 | 5.2 | 5.46 | 5.685 | 6.1 | 5.43
| Std Dev. |  | 0.065549 | 0.12111 | 0.20976 | 0.048477 | 0.16021 | 0.25626 | 0.040825 | 0.052026 | 0.098319 | 0.048854 | 0.067429 | 0.098319 | 0.05164 | 0.17942 | 0.04761 | 0.094322 | 0.08165 | 0.15492 | 0.14142 | 0.089088 | 0.12024 | 0.05164 | 0.16905
| Rel.Std.Dev. |  | 1.246% | 2.299% | 3.33% | 0.8497% | 3.091% | 5.551% | 0.7537% | 0.9471% | 1.849% | 0.9109% | 1.321% | 1.861% | 0.9007% | 3.395% | 0.8773% | 1.693% | 1.622% | 3.098% | 2.72% | 1.635% | 2.114% | 0.842% | 3.13%
| PDM3 |  | -1.873% | -1.779% | 17.49% | 6.395% | -3.334% | -13.9% | 1.018% | 2.448% | -0.847% | 0.02331% | -4.826% | -1.469% | 6.924% | -1.438% | 1.204% | 3.909% | -6.131% | -6.753% | -3.023% | 1.609% | 6.084% | 14.38% | 0.7382%
| Z-Score (Absolute) |  | 0.42 | 0.4 | 3.87 | 1.41 | 0.74 | 3.08 | 0.22 | 0.54 | 0.19 | 0 | 1.07 | 0.33 | 1.53 | 0.32 | 0.26 | 0.86 | 1.36 | 1.5 | 0.67 | 0.35 | 1.34 | 3.18 | 0.16
| Analytical results for Ti in OREAS 282 (Certified Value 0.161 wt.%) | [26 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.159 | 0.18 | 0.11 | 0.16 | 0.1756 | 0.15 | 0.15 | 0.165 | 0.154 | 0.1699 | 0.16 | 0.1549 | 0.164 | 0.161 | 0.17 | 0.154 | 0.16 | 0.145 | 0.159 | 0.16 | 0.1576 | 0.17 | 0.162 | 0.16 | 0.164 | 1
| 1 | 2 | 0.16 | 0.18 | 0.11 | 0.17 | 0.1756 | 0.15 | 0.148 | 0.17 | 0.154 | 0.1642 | 0.17 | 0.1578 | 0.1652 | 0.165 | 0.17 | 0.157 | 0.158 | 0.149 | 0.159 | 0.16 | 0.1545 | 0.17 | 0.159 | 0.16 | 0.16 | 26
| 1 | 3 | 0.159 | 0.19 | 0.11 | 0.16 | 0.1774 | 0.15 | 0.152 | 0.17 | 0.154 | 0.167 | 0.16 | 0.1582 | 0.1674 | 0.162 | 0.17 | 0.153 | 0.159 | 0.149 | 0.157 | 0.16 | 0.1616 | 0.16 | 0.159 | 0.16 | 0.157 | 16
| 1 | 4 | 0.164 | 0.18 | 0.11 | 0.17 | 0.178 | 0.15 | 0.15 | 0.17 | 0.153 | 0.1663 | 0.15 | 0.1588 | 0.1641 | 0.158 | 0.17 | 0.153 | 0.16 | 0.146 | 0.16 | 0.16 | 0.1596 | 0.17 | 0.159 | 0.16 | 0.149 | 0.16051
| 1 | 5 | 0.16 | 0.19 | 0.11 | 0.17 | 0.1774 | 0.15 | 0.15 | 0.165 | 0.152 | 0.1677 | 0.15 | 0.1573 | 0.1639 | 0.164 | 0.17 | 0.154 | 0.157 | 0.151 | 0.158 | 0.16 | 0.1576 | 0.17 | 0.161 | 0.16 | 0.162 | 60
| 1 | 6 | 0.161 | 0.19 | 0.12 | 0.16 | 0.1762 | 0.16 | 0.15 | 0.165 | 0.153 | 0.1669 | 0.16 | 0.1548 | 0.1682 | 0.163 | 0.17 | 0.159 | 0.161 | 0.147 | 0.156 | 0.16 | 0.1616 | 0.16 | 0.162 | 0.16 | 0.163
| Mean |  | 0.1605 | 0.185 | 0.11167 | 0.165 | 0.1767 | 0.15167 | 0.15 | 0.1675 | 0.15333 | 0.167 | 0.15833 | 0.15697 | 0.16547 | 0.16217 | 0.17 | 0.155 | 0.15917 | 0.14783 | 0.15817 | 0.16 | 0.15875 | 0.16667 | 0.16033 | 0.16 | 0.15917
| Median |  | 0.16 | 0.185 | 0.11 | 0.165 | 0.1768 | 0.15 | 0.15 | 0.1675 | 0.1535 | 0.16695 | 0.16 | 0.15755 | 0.16465 | 0.1625 | 0.17 | 0.154 | 0.1595 | 0.148 | 0.1585 | 0.16 | 0.1586 | 0.17 | 0.16 | 0.16 | 0.161
| Std Dev. |  | 0.0018708 | 0.0054772 | 0.0040825 | 0.0054772 | 0.0010334 | 0.0040825 | 0.0012649 | 0.0027386 | 0.0008165 | 0.0018569 | 0.0075277 | 0.0017119 | 0.0018843 | 0.0024833 | 0 | 0.0024495 | 0.001472 | 0.0022286 | 0.001472 | 0 | 0.002745 | 0.005164 | 0.0015055 | 0 | 0.0055648
| Rel.Std.Dev. |  | 1.166% | 2.961% | 3.656% | 3.32% | 0.5849% | 2.692% | 0.8433% | 1.635% | 0.5325% | 1.112% | 4.754% | 1.091% | 1.139% | 1.531% | 0% | 1.58% | 0.9248% | 1.508% | 0.9306% | 0% | 1.729% | 3.098% | 0.939% | 0% | 3.496%
| PDM3 |  | -0.008072% | 15.26% | -30.43% | 2.795% | 10.08% | -5.511% | -6.55% | 4.353% | -4.473% | 4.041% | -1.358% | -2.209% | 3.086% | 1.03% | 5.91% | -3.435% | -0.8387% | -7.899% | -1.462% | -0.3196% | -1.098% | 3.834% | -0.1119% | -0.3196% | -0.8387%
| Z-Score (Absolute) |  | 0.07 | 3.37 | 6.52 | 0.67 | 2.26 | 1.12 | 1.35 | 1.01 | 0.9 | 0.94 | 0.22 | 0.41 | 0.74 | 0.29 | 1.35 | 0.67 | 0.11 | 1.64 | 0.25 | 0 | 0.17 | 0.9 | 0.04 | 0 | 0.11
| Analytical results for Tl in OREAS 282 (Certified Value 28.8 ppm) | [26 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 28.9 | 28.8 | 29.33 | 27.35 | 22.9 | 29.9 | 28.36 | 29.85 | 28.4 | 29.11 | 30.3 | 26.23 | 27.2 | 28.4 | 27.7 | 29.04 | 28.63 | 28.22 | 30.4 | 29.15 | 30.1 |  |  |  |  | 1
| 1 | 2 | 29.8 | 29.5 | 29.86 | 27.05 | 23.3 | 29.6 | 27.43 | 28.94 | 28.56 | 29.04 | 29.4 | 26.8 | 27.4 | 28.4 | 28.2 | 30.06 | 28.63 | 26.93 | 29.6 | 28.24 | 29.8 |  |  |  |  | 27
| 1 | 3 | 28.9 | 29.6 | 28.73 | 27.5 | 22.9 | 29.1 | 27.83 | 29.69 | 27.83 | 29.4 | 30.5 | 26.75 | 28.6 | 29.1 | 28.1 | 29.14 | 29.05 | 27.32 | 29.2 | 27.83 | 30.4 |  |  |  |  | 16
| 1 | 4 | 30.3 | 29 | 29.82 | 27.65 | 23.6 | 29.7 | 27.86 | 28.6 | 28.47 | 29.51 | 29.7 | 27.05 | 27.8 | 28.8 | 28.6 | 29.98 | 29.05 | 28.85 | 29.7 | 28.21 | 28.4 |  |  |  |  | 28.7725
| 1 | 5 | 29.8 | 28.8 | 29.33 | 26.9 | 23.3 | 29.7 | 27.83 | 29.36 | 28.02 | 29 | 29.4 | 26.63 | 28.8 | 29.4 | 27.9 | 30.05 | 29.25 | 26.29 | 30.5 | 27.96 | 31.4 |  |  |  |  | 61
| 1 | 6 | 28.7 | 29.7 | 29.02 | 27.35 | 22.7 | 29.3 | 28.05 | 30.06 | 27.81 | 29.55 | 30.1 | 27.02 | 27.4 | 28.4 | 28.5 | 29.46 | 28.91 | 27.29 | 30.2 | 28.63 | 30.7
| Mean |  | 29.4 | 29.2333 | 29.3483 | 27.3 | 23.1167 | 29.55 | 27.8933 | 29.4167 | 28.1817 | 29.2683 | 29.9 | 26.7467 | 27.8667 | 28.75 | 28.1667 | 29.6217 | 28.92 | 27.4833 | 29.9333 | 28.3367 | 30.1333
| Median |  | 29.35 | 29.25 | 29.33 | 27.35 | 23.1 | 29.65 | 27.845 | 29.525 | 28.21 | 29.255 | 29.9 | 26.775 | 27.6 | 28.6 | 28.15 | 29.72 | 28.98 | 27.305 | 29.95 | 28.225 | 30.25
| Std Dev. |  | 0.65115 | 0.41312 | 0.44161 | 0.27928 | 0.33714 | 0.29496 | 0.30533 | 0.56102 | 0.33523 | 0.24669 | 0.46904 | 0.30004 | 0.67725 | 0.42778 | 0.34448 | 0.46914 | 0.2494 | 0.91703 | 0.51251 | 0.48405 | 1.0113
| Rel.Std.Dev. |  | 2.215% | 1.413% | 1.505% | 1.023% | 1.458% | 0.9982% | 1.095% | 1.907% | 1.19% | 0.8429% | 1.569% | 1.122% | 2.43% | 1.488% | 1.223% | 1.584% | 0.8624% | 3.337% | 1.712% | 1.708% | 3.356%
| PDM3 |  | 2.181% | 1.602% | 2.001% | -5.118% | -19.66% | 2.702% | -3.056% | 2.239% | -2.053% | 1.723% | 3.919% | -7.041% | -3.148% | -0.0782% | -2.106% | 2.951% | 0.5126% | -4.481% | 4.035% | -1.515% | 4.73%
| Z-Score (Absolute) |  | 0.44 | 0.29 | 0.39 | 1.48 | 5.3 | 0.58 | 0.94 | 0.45 | 0.67 | 0.32 | 0.9 | 1.98 | 0.96 | 0.16 | 0.69 | 0.64 | 0 | 1.31 | 0.93 | 0.53 | 1.11
| Analytical results for Tm in OREAS 282 (Certified Value 0.2 ppm) | [26 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [14 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [14 empty] | 2
3
| 1 | 1 | 0.2 | 0.15 | 0.2 | 0.2 | 0.21 | 0.2 | 0.18 | 0.2 | 0.2 | 0.2 | 0.19 | [14 empty] | 1
| 1 | 2 | 0.2 | 0.15 | 0.2 | 0.19 | 0.2 | 0.2 | 0.19 | 0.2 | 0.2 | 0.2 | 0.2 | [14 empty] | 28
| 1 | 3 | 0.2 | 0.15 | 0.2 | 0.19 | 0.21 | 0.2 | 0.19 | 0.18 | 0.2 | 0.2 | 0.2 | [14 empty] | 16
| 1 | 4 | 0.2 | 0.15 | 0.2 | 0.2 | 0.19 | 0.2 | 0.19 | 0.19 | 0.2 | 0.2 | 0.19 | [14 empty] | 0.19773
| 1 | 5 | 0.2 | 0.15 | 0.2 | 0.2 | 0.22 | 0.2 | 0.19 | 0.19 | 0.2 | 0.2 | 0.19 | [14 empty] | 62
| 1 | 6 | 0.2 | 0.15 | 0.2 | 0.2 | 0.21 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.19
| Mean |  | 0.2 | 0.15 | 0.2 | 0.19667 | 0.20667 | 0.2 | 0.19 | 0.19333 | 0.2 | 0.2 | 0.19333
| Median |  | 0.2 | 0.15 | 0.2 | 0.2 | 0.21 | 0.2 | 0.19 | 0.195 | 0.2 | 0.2 | 0.19
| Std Dev. |  | 3.0405e-17 | 0 | 3.0405e-17 | 0.005164 | 0.010328 | 3.0405e-17 | 0.0063246 | 0.008165 | 3.0405e-17 | 3.0405e-17 | 0.005164
| Rel.Std.Dev. |  | 1.52e-14% | 0% | 1.52e-14% | 2.626% | 4.997% | 1.52e-14% | 3.329% | 4.223% | 1.52e-14% | 1.52e-14% | 2.671%
| PDM3 |  | 1.146% | -24.14% | 1.146% | -0.5394% | 4.518% | 1.146% | -3.911% | -2.225% | 1.146% | 1.146% | -2.225%
| Z-Score (Absolute) |  | 0 | 10.11 | 0 | 0.67 | 1.35 | 0 | 2.02 | 1.35 | 0 | 0 | 1.35
| Analytical results for U in OREAS 282 (Certified Value 7.31 ppm) | [26 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.9 | 7.4 | 7.7 | 7.51 | 6.6 | 7.3 | 7.9 | 7.47 | 7.56 | 7.59 | 7.51 | 7.4 | 7.49 | 6.85 | 7.3 | 7.19 | 6.9 | 6.49 | 7 | 7.7 | 7.1 | 7.8 | 7.1 |  |  | 1
| 1 | 2 | 7.1 | 7.6 | 8 | 7.66 | 6.5 | 7.4 | 7.9 | 7.22 | 7.38 | 7.53 | 7.66 | 7.1 | 7.72 | 6.81 | 7.2 | 7.04 | 7.2 | 6.57 | 7 | 7.5 | 6.9 | 7.5 | 7.1 |  |  | 29
| 1 | 3 | 7.2 | 7.6 | 7.5 | 7.5 | 6.7 | 7.4 | 7.9 | 7.36 | 7.47 | 7.52 | 7.82 | 7.3 | 7.47 | 7.2 | 7.3 | 7.09 | 7.1 | 6.53 | 7.1 | 7.4 | 6.8 | 7.4 | 7.3 |  |  | 16
| 1 | 4 | 7.1 | 7.5 | 8 | 7.63 | 6.7 | 7.3 | 8 | 7.32 | 7.16 | 7.58 | 7.61 | 7.4 | 7.71 | 6.8 | 7.3 | 7.26 | 7.3 | 6.48 | 7.1 | 7.6 | 7 | 7.5 | 6.8 |  |  | 7.3108
| 1 | 5 | 7.2 | 7.5 | 7.8 | 7.65 | 6.5 | 7.3 | 7.8 | 7.36 | 7.58 | 7.49 | 7.63 | 7.3 | 7.63 | 7.15 | 7.3 | 7.03 | 7.3 | 6.48 | 7.2 | 7.3 | 7.1 | 7.6 | 7.4 |  |  | 63
| 1 | 6 | 7.1 | 7.7 | 7.5 | 7.68 | 6.6 | 7.3 | 7.8 | 7.46 | 7.69 | 7.49 | 7.68 | 7.4 | 7.69 | 6.81 | 7.2 | 7.22 | 7.2 | 6.51 | 7.2 | 7.5 | 7.2 | 7.6 | 7.3
| Mean |  | 7.1 | 7.55 | 7.75 | 7.605 | 6.6 | 7.3333 | 7.8833 | 7.365 | 7.4733 | 7.5333 | 7.6517 | 7.3167 | 7.6183 | 6.9367 | 7.2667 | 7.1383 | 7.1667 | 6.51 | 7.1 | 7.5 | 7.0167 | 7.5667 | 7.1667
| Median |  | 7.1 | 7.55 | 7.75 | 7.64 | 6.6 | 7.3 | 7.9 | 7.36 | 7.515 | 7.525 | 7.645 | 7.35 | 7.66 | 6.83 | 7.3 | 7.14 | 7.2 | 6.5 | 7.1 | 7.5 | 7.05 | 7.55 | 7.2
| Std Dev. |  | 0.10954 | 0.10488 | 0.22583 | 0.079183 | 0.089443 | 0.05164 | 0.075277 | 0.092898 | 0.18587 | 0.043205 | 0.10147 | 0.1169 | 0.11179 | 0.18608 | 0.05164 | 0.09786 | 0.15055 | 0.035214 | 0.089443 | 0.14142 | 0.1472 | 0.13663 | 0.21602
| Rel.Std.Dev. |  | 1.543% | 1.389% | 2.914% | 1.041% | 1.355% | 0.7042% | 0.9549% | 1.261% | 2.487% | 0.5735% | 1.326% | 1.598% | 1.467% | 2.683% | 0.7106% | 1.371% | 2.101% | 0.5409% | 1.26% | 1.886% | 2.098% | 1.806% | 3.014%
| PDM3 |  | -2.883% | 3.272% | 6.008% | 4.024% | -9.723% | 0.3083% | 7.831% | 0.7414% | 2.223% | 3.044% | 4.663% | 0.08029% | 4.207% | -5.118% | -0.6036% | -2.359% | -1.971% | -10.95% | -2.883% | 2.588% | -4.023% | 3.5% | -1.971%
| Z-Score (Absolute) |  | 0.67 | 0.63 | 1.2 | 0.79 | 2.12 | 0 | 1.59 | 0.09 | 0.4 | 0.58 | 0.92 | 0.05 | 0.82 | 1.15 | 0.19 | 0.56 | 0.48 | 2.38 | 0.67 | 0.48 | 0.92 | 0.67 | 0.48
| Analytical results for V in OREAS 282 (Certified Value 105 ppm) | [26 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 108 | 110 | 99 | 110 | 119 | 114 | 100 | 95 | 115 | 106 | 100 | 105 | 105 | 103 | 103 | 100 | 105 | 96 | 108 | 105 | 103 | 111 | 110 | 105 | 108 | 1
| 1 | 2 | 109 | 105 | 102 | 116 | 122 | 117 | 100 | 100 | 114 | 107 | 104 | 105 | 105 | 99 | 104 | 92 | 104 | 95 | 111 | 106 | 101 | 112 | 107 | 104 | 104 | 30
| 1 | 3 | 109 | 110 | 101 | 109 | 122 | 117 | 100 | 100 | 114 | 106 | 99 | 106 | 107 | 103 | 106 | 100 | 104 | 95 | 109 | 103 | 104 | 111 | 105 | 105 | 103 | 16
| 1 | 4 | 111 | 110 | 99 | 115 | 119 | 112 | 100 | 95 | 115 | 109 | 100 | 106 | 105 | 101 | 106 | 99 | 104 | 96 | 110 | 104 | 104 | 111 | 108 | 104 | 100 | 105.3
| 1 | 5 | 109 | 100 | 101 | 114 | 119 | 115 | 100 | 100 | 116 | 110 | 98 | 103 | 105 | 100 | 105 | 93 | 103 | 96 | 109 | 104 | 107 | 111 | 108 | 103 | 106 | 64
| 1 | 6 | 111 | 105 | 110 | 104 | 122 | 112 | 100 | 100 | 113 | 108 | 103 | 104 | 106 | 103 | 106 | 97 | 104 | 96 | 107 | 103 | 107 | 111 | 108 | 103 | 107
| Mean |  | 109.5 | 106.6667 | 102 | 111.3333 | 120.5 | 114.5 | 100 | 98.3333 | 114.5 | 107.6667 | 100.6667 | 104.8333 | 105.5 | 101.5 | 105 | 96.8333 | 104 | 95.6667 | 109 | 104.1667 | 104.3333 | 111.1667 | 107.6667 | 104 | 104.6667
| Median |  | 109 | 107.5 | 101 | 112 | 120.5 | 114.5 | 100 | 100 | 114.5 | 107.5 | 100 | 105 | 105 | 102 | 105.5 | 98 | 104 | 96 | 109 | 104 | 104 | 111 | 108 | 104 | 105
| Std Dev. |  | 1.2247 | 4.0825 | 4.0988 | 4.5461 | 1.6432 | 2.2583 | 0 | 2.582 | 1.0488 | 1.633 | 2.3381 | 1.169 | 0.83666 | 1.7607 | 1.2649 | 3.5449 | 0.63246 | 0.5164 | 1.4142 | 1.169 | 2.3381 | 0.40825 | 1.633 | 0.89443 | 2.9439
| Rel.Std.Dev. |  | 1.118% | 3.827% | 4.018% | 4.083% | 1.364% | 1.972% | 0% | 2.626% | 0.916% | 1.517% | 2.323% | 1.115% | 0.793% | 1.735% | 1.205% | 3.661% | 0.6081% | 0.5398% | 1.297% | 1.122% | 2.241% | 0.3672% | 1.517% | 0.86% | 2.813%
| PDM3 |  | 3.989% | 1.298% | -3.134% | 5.73% | 14.43% | 8.737% | -5.033% | -6.616% | 8.737% | 2.248% | -4.4% | -0.4432% | 0.1899% | -3.609% | -0.2849% | -8.041% | -1.235% | -9.148% | 3.514% | -1.076% | -0.918% | 5.571% | 2.248% | -1.235% | -0.6015%
| Z-Score (Absolute) |  | 0.98 | 0.37 | 0.62 | 1.37 | 3.32 | 2.04 | NA | 1.4 | 2.04 | 0.59 | 0.9 | 0.02 | 0.12 | 0.73 | 0.02 | 1.72 | 0.2 | 1.97 | 0.87 | 0.16 | 0.12 | 1.33 | 0.59 | 0.2 | 0.05
Results from laboratory 7 were removed due to their 10 ppm reading resolution.
| Analytical results for W in OREAS 282 (Certified Value 25.9 ppm) | [26 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 25.9 | 25 | 49 | 25.4 | 29.6 | 23 | 28 | 24.5 | 27 | 24.95 | 24.5 | 28.6 | 27.9 | 25.8 | 27.3 | 26.1 | 24.1 | 24.2 | 23.3 | 25.8 | 26.7 | 26.7 | 28.6 |  |  | 1
| 1 | 2 | 25.2 | 26 | 46 | 26.7 | 29.3 | 22 | 28 | 24.2 | 26.5 | 24.7 | 24.4 | 28.4 | 26.6 | 26 | 26.4 | 26 | 24.4 | 24.7 | 24.5 | 25.4 | 25.4 | 26.5 | 27.1 |  |  | 31
| 1 | 3 | 25.2 | 26 | 44 | 25 | 28.8 | 24 | 28 | 24.6 | 27.1 | 25.44 | 23.8 | 28.7 | 28.3 | 25.7 | 27.3 | 25.2 | 23.8 | 24.2 | 23.9 | 25.6 | 25.3 | 26.4 | 27 |  |  | 16
| 1 | 4 | 26.5 | 25.5 | 46 | 26.6 | 29.5 | 23 | 27 | 24.5 | 27.1 | 24.89 | 24.3 | 28.7 | 28 | 26.5 | 25.9 | 25.9 | 23.4 | 25 | 24.5 | 25.2 | 25.4 | 26.1 | 25.7 |  |  | 25.9244
| 1 | 5 | 25.9 | 25 | 46 | 25.5 | 29.3 | 23 | 27 | 24.6 | 27.3 | 25.32 | 23.9 | 28.1 | 27.3 | 26.3 | 26.3 | 25.5 | 24.1 | 25 | 24.7 | 25 | 26.6 | 26.1 | 27.3 |  |  | 65
| 1 | 6 | 25.4 | 26.5 | 47 | 24.7 | 30 | 22 | 27.5 | 24.5 | 26.8 | 25.74 | 23.8 | 29 | 26.7 | 26.1 | 27 | 26.1 | 23.7 | 24.6 | 24.8 | 25.5 | 26.7 | 26.4 | 28.3
| Mean |  | 25.6833 | 25.6667 | 46.3333 | 25.65 | 29.4167 | 22.8333 | 27.5833 | 24.4833 | 26.9667 | 25.1733 | 24.1167 | 28.5833 | 27.4667 | 26.0667 | 26.7 | 25.8 | 23.9167 | 24.6167 | 24.2833 | 25.4167 | 26.0167 | 26.3667 | 27.3333
| Median |  | 25.65 | 25.75 | 46 | 25.45 | 29.4 | 23 | 27.75 | 24.5 | 27.05 | 25.135 | 24.1 | 28.65 | 27.6 | 26.05 | 26.7 | 25.95 | 23.95 | 24.65 | 24.5 | 25.45 | 26 | 26.4 | 27.2
| Std Dev. |  | 0.51153 | 0.60553 | 1.633 | 0.82644 | 0.39707 | 0.75277 | 0.4916 | 0.1472 | 0.28048 | 0.3919 | 0.31885 | 0.30605 | 0.71181 | 0.30111 | 0.5831 | 0.36878 | 0.35449 | 0.36009 | 0.57417 | 0.28577 | 0.71391 | 0.23381 | 1.0367
| Rel.Std.Dev. |  | 1.992% | 2.359% | 3.524% | 3.222% | 1.35% | 3.297% | 1.782% | 0.6012% | 1.04% | 1.557% | 1.322% | 1.071% | 2.592% | 1.155% | 2.184% | 1.429% | 1.482% | 1.463% | 2.364% | 1.124% | 2.744% | 0.8868% | 3.793%
| PDM3 |  | -0.9299% | -0.9941% | 78.72% | -1.058% | 13.47% | -11.92% | 6.399% | -5.559% | 4.02% | -2.897% | -6.973% | 10.26% | 5.949% | 0.5488% | 2.992% | -0.4798% | -7.745% | -5.044% | -6.33% | -1.958% | 0.3559% | 1.706% | 5.435%
| Z-Score (Absolute) |  | 0.07 | 0.08 | 11.7 | 0.09 | 2.06 | 1.69 | 1.02 | 0.75 | 0.66 | 0.36 | 0.96 | 1.59 | 0.95 | 0.15 | 0.51 | 0 | 1.07 | 0.67 | 0.86 | 0.22 | 0.12 | 0.32 | 0.87
| Analytical results for Y in OREAS 282 (Certified Value 16.6 ppm) | [26 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 16.8 | 16 | 13 | 16.1 | 15.6 | 15.4 | 17.4 | 16.7 | 17.5 | 16.32 | 15.84 | 16.1 | 15.9 | 15.2 | 17.3 | 17.1 | 16.8 | 16.3 | 16 | 18.1 | 15.8 | 17 | 16.9 | 18.1 | 17.2 | 1
| 1 | 2 | 17.8 | 16.2 | 13 | 17.2 | 15.7 | 15.9 | 17.6 | 17.1 | 17.4 | 15.66 | 15.7 | 16.1 | 15.89 | 14.7 | 17.6 | 17.5 | 17 | 16.5 | 16.4 | 17.5 | 15.7 | 17 | 16.8 | 17.8 | 16.9 | 32
| 1 | 3 | 16.9 | 16.4 | 13 | 16 | 15.5 | 16.1 | 18 | 16.3 | 17.7 | 16.17 | 16.05 | 16 | 16.08 | 15.5 | 17.9 | 17.9 | 16.6 | 16.6 | 16.2 | 17.2 | 16.1 | 17 | 16.2 | 17.8 | 17.1 | 16
| 1 | 4 | 17.2 | 16 | 12 | 17.4 | 15.4 | 15.4 | 17.3 | 17.2 | 17.5 | 15.96 | 15.86 | 16.1 | 16.07 | 15.5 | 17.3 | 17.1 | 16.9 | 16.3 | 16.3 | 17.6 | 16 | 17 | 16.5 | 17.6 | 15.9 | 16.5977
| 1 | 5 | 17.9 | 16 | 13 | 17 | 15.8 | 15.7 | 17.8 | 16.8 | 17.4 | 15.98 | 16.13 | 15.6 | 15.75 | 14.6 | 17.7 | 16.3 | 16.5 | 16.6 | 16.3 | 18.4 | 16.1 | 17 | 16.9 | 17.4 | 17 | 66
| 1 | 6 | 17.1 | 16.4 | 13 | 16.6 | 16 | 15.2 | 17.5 | 16.9 | 17.4 | 15.95 | 16.14 | 15.6 | 16.15 | 15.5 | 17.3 | 17.2 | 16.6 | 16.4 | 15.9 | 17.9 | 16.5 | 17 | 16.7 | 17.7 | 17.2
| Mean |  | 17.2833 | 16.1667 | 12.8333 | 16.7167 | 15.6667 | 15.6167 | 17.6 | 16.8333 | 17.4833 | 16.0067 | 15.9533 | 15.9167 | 15.9733 | 15.1667 | 17.5167 | 17.1833 | 16.7333 | 16.45 | 16.1833 | 17.7833 | 16.0333 | 17 | 16.6667 | 17.7333 | 16.8833
| Median |  | 17.15 | 16.1 | 13 | 16.8 | 15.65 | 15.55 | 17.55 | 16.85 | 17.45 | 15.97 | 15.955 | 16.05 | 15.985 | 15.35 | 17.45 | 17.15 | 16.7 | 16.45 | 16.25 | 17.75 | 16.05 | 17 | 16.75 | 17.75 | 17.05
| Std Dev. |  | 0.46224 | 0.19664 | 0.40825 | 0.58109 | 0.21602 | 0.34303 | 0.26077 | 0.32042 | 0.1169 | 0.22411 | 0.17952 | 0.24833 | 0.15108 | 0.41793 | 0.25626 | 0.53072 | 0.19664 | 0.13784 | 0.19408 | 0.43551 | 0.28048 | 0 | 0.27325 | 0.23381 | 0.49565
| Rel.Std.Dev. |  | 2.674% | 1.216% | 3.181% | 3.476% | 1.379% | 2.197% | 1.482% | 1.903% | 0.6687% | 1.4% | 1.125% | 1.56% | 0.9459% | 2.756% | 1.463% | 3.089% | 1.175% | 0.8379% | 1.199% | 2.449% | 1.749% | 0% | 1.64% | 1.318% | 2.936%
| PDM3 |  | 4.131% | -2.597% | -22.68% | 0.7169% | -5.609% | -5.911% | 6.039% | 1.42% | 5.336% | -3.561% | -3.882% | -4.103% | -3.762% | -8.622% | 5.537% | 3.529% | 0.8173% | -0.8898% | -2.496% | 7.143% | -3.4% | 2.424% | 0.4156% | 6.842% | 1.721%
| Z-Score (Absolute) |  | 0.63 | 0.51 | NA | 0.05 | 1.02 | 1.07 | 0.95 | 0.17 | 0.83 | 0.67 | 0.73 | 0.77 | 0.71 | 1.53 | 0.87 | 0.53 | 0.07 | 0.22 | 0.49 | 1.14 | 0.65 | NA | 0 | 1.09 | 0.22
Results from laboratories 3 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 282 (Certified Value 1.31 ppm) | [26 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.23 | 2.27 | 2.28 | [12 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 1.35 | 1.3 | 1.29 | 1.4 | 1.25 | 1.27 | 1.35 | 1.3 | 1.3 | 1.3 | 1.2 | 1.3 | 1.45 | [12 empty] | 1
| 1 | 2 | 1.4 | 1.3 | 1.3 | 1.4 | 1.25 | 1.3 | 1.33 | 1.3 | 1.3 | 1.4 | 1.3 | 1.3 | 1.4 | [12 empty] | 33
| 1 | 3 | 1.45 | 1.3 | 1.29 | 1.45 | 1.25 | 1.3 | 1.42 | 1.3 | 1.3 | 1.2 | 1.2 | 1.3 | 1.44 | [12 empty] | 16
| 1 | 4 | 1.3 | 1.3 | 1.34 | 1.35 | 1.2 | 1.31 | 1.45 | 1.3 | 1.4 | 1.2 | 1.2 | 1.3 | 1.46 | [12 empty] | 1.3057
| 1 | 5 | 1.4 | 1.3 | 1.31 | 1.3 | 1.2 | 1.29 | 1.36 | 1.2 | 1.3 | 1.2 | 1.2 | 1.3 | 1.42 | [12 empty] | 67
| 1 | 6 | 1.4 | 1.3 | 1.32 | 1.4 | 1.25 | 1.29 | 1.39 | 1.3 | 1.3 | 1.3 | 1.2 | 1.3 | 1.48
| Mean |  | 1.3833 | 1.3 | 1.3083 | 1.3833 | 1.2333 | 1.2933 | 1.3833 | 1.2833 | 1.3167 | 1.2667 | 1.2167 | 1.3 | 1.4417
| Median |  | 1.4 | 1.3 | 1.305 | 1.4 | 1.25 | 1.295 | 1.375 | 1.3 | 1.3 | 1.25 | 1.2 | 1.3 | 1.445
| Std Dev. |  | 0.05164 | 0 | 0.019408 | 0.05164 | 0.02582 | 0.013663 | 0.045461 | 0.040825 | 0.040825 | 0.08165 | 0.040825 | 0 | 0.028577
| Rel.Std.Dev. |  | 3.733% | 0% | 1.483% | 3.733% | 2.094% | 1.056% | 3.286% | 3.181% | 3.101% | 6.446% | 3.355% | 0% | 1.982%
| PDM3 |  | 5.946% | -0.4361% | 0.2021% | 5.946% | -5.542% | -0.9467% | 5.946% | -1.713% | 0.8403% | -2.989% | -6.818% | -0.4361% | 10.41%
| Z-Score (Absolute) |  | 1.69 | 0 | 0.17 | 1.69 | 1.35 | 0.13 | 1.69 | 0.34 | 0.34 | 0.67 | 1.69 | 0 | 2.87
| Analytical results for Zn in OREAS 282 (Certified Value 121 ppm) | [26 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 130 | 122 | 111 | 126 | 120 | 117 | 110 | 122 | 141 | 124 | 127 | 121 | 121 | 112 | 113 | 132 | 132 | 108 | 110.7 | 115 | 119 | 119 | 136 | 117 | 134 | 1
| 1 | 2 | 131 | 124 | 108 | 130 | 116 | 119 | 110 | 122 | 125 | 122 | 134 | 121 | 120 | 112 | 115 | 129 | 130 | 103 | 113 | 115 | 118 | 121 | 133 | 115 | 131 | 34
| 1 | 3 | 130 | 124 | 105 | 126 | 115 | 121 | 115 | 122 | 129 | 124 | 127 | 119 | 121 | 114 | 115 | 134 | 131 | 100 | 111.5 | 111 | 122 | 121 | 133 | 117 | 128 | 16
| 1 | 4 | 133 | 120 | 104 | 130 | 115 | 115 | 110 | 122 | 131 | 123 | 122 | 121 | 119 | 118 | 111 | 131 | 131 | 108 | 112.9 | 113 | 119 | 120 | 134 | 115 | 122 | 120.6753
| 1 | 5 | 131 | 120 | 107 | 129 | 120 | 120 | 115 | 122 | 125 | 123 | 119 | 120 | 119 | 108 | 113 | 128 | 130 | 99 | 111.5 | 113 | 117 | 121 | 135 | 116 | 132 | 68
| 1 | 6 | 133 | 126 | 115 | 122 | 119 | 115 | 115 | 124 | 123 | 124 | 126 | 119 | 121 | 115 | 113 | 138 | 130 | 108 | 109.1 | 114 | 116 | 119 | 133 | 116 | 133
| Mean |  | 131.3333 | 122.6667 | 108.3333 | 127.1667 | 117.5 | 117.8333 | 112.5 | 122.3333 | 129 | 123.3333 | 125.8333 | 120.1667 | 120.1667 | 113.1667 | 113.3333 | 132 | 130.6667 | 104.3333 | 111.45 | 113.5 | 118.5 | 120.1667 | 134 | 116 | 130
| Median |  | 131 | 123 | 107.5 | 127.5 | 117.5 | 118 | 112.5 | 122 | 127 | 123.5 | 126.5 | 120.5 | 120.5 | 113 | 113 | 131.5 | 130.5 | 105.5 | 111.5 | 113.5 | 118.5 | 120.5 | 133.5 | 116 | 131.5
| Std Dev. |  | 1.3663 | 2.4221 | 4.0825 | 3.1252 | 2.429 | 2.5626 | 2.7386 | 0.8165 | 6.5727 | 0.8165 | 5.1153 | 0.98319 | 0.98319 | 3.3714 | 1.5055 | 3.6332 | 0.8165 | 4.2269 | 1.4557 | 1.5166 | 2.0736 | 0.98319 | 1.2649 | 0.89443 | 4.4272
| Rel.Std.Dev. |  | 1.04% | 1.975% | 3.768% | 2.458% | 2.067% | 2.175% | 2.434% | 0.6674% | 5.095% | 0.662% | 4.065% | 0.8182% | 0.8182% | 2.979% | 1.328% | 2.752% | 0.6249% | 4.051% | 1.306% | 1.336% | 1.75% | 0.8182% | 0.944% | 0.7711% | 3.406%
| PDM3 |  | 8.832% | 1.65% | -10.23% | 5.379% | -2.631% | -2.355% | -6.775% | 1.374% | 6.898% | 2.203% | 4.274% | -0.4215% | -0.4215% | -6.222% | -6.084% | 9.384% | 8.28% | -13.54% | -7.645% | -5.946% | -1.803% | -0.4215% | 11.04% | -3.874% | 7.727%
| Z-Score (Absolute) |  | 1.1 | 0.25 | 1.17 | 0.69 | 0.26 | 0.23 | 0.76 | 0.21 | 0.87 | 0.31 | 0.56 | 0 | 0 | 0.69 | 0.67 | 1.17 | 1.04 | 1.56 | 0.86 | 0.66 | 0.16 | 0 | 1.37 | 0.41 | 0.97
| Analytical results for Zr in OREAS 282 (Certified Value 58 ppm) | [26 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 61.5 | 55 | 48 | 78.8 | 71 | 55 | 52 | 67.2 | 52.4 | 76.9 | 50.8 | 52.3 | 65 | 53.1 | 67.6 | 53.5 | 53.2 | 47.5 | 60 | 50.6 | 59 | 60.8 | 58.3 | 57.5 |  | 1
| 1 | 2 | 60.2 | 60 | 46 | 72.6 | 67 | 60 | 49 | 68.3 | 48.9 | 64.2 | 51.4 | 52.9 | 61 | 57.6 | 61.6 | 54 | 55 | 48.5 | 59.5 | 49.9 | 60 | 77.8 | 58.4 | 56.3 |  | 35
| 1 | 3 | 57.6 | 57 | 45 | 75.2 | 71 | 60 | 49 | 66.2 | 45.6 | 62.3 | 52.3 | 53.3 | 66 | 60.7 | 66.3 | 54.2 | 53.6 | 47.5 | 59.4 | 50 | 60 | 58.5 | 59.7 | 61.6 |  | 16
| 1 | 4 | 63.7 | 55 | 45 | 73.8 | 64 | 60 | 51 | 67.1 | 48.9 | 66.2 | 52.5 | 48.8 | 63 | 58.4 | 53 | 55.6 | 55.4 | 47.7 | 58.5 | 53 | 50 | 63.5 | 56.2 | 57.3 |  | 57.8114
| 1 | 5 | 63.4 | 58 | 45 | 70.5 | 65 | 60 | 51 | 66.3 | 48.6 | 64.3 | 52.7 | 51.6 | 59 | 61 | 67.2 | 52.5 | 54 | 48.6 | 59 | 51.4 | 60 | 81.7 | 58 | 57.2 |  | 69
| 1 | 6 | 67.6 | 56 | 46 | 69.2 | 66 | 55 | 45 | 66.9 | 48.3 | 68.3 | 50.6 | 57.8 | 69 | 59.9 | 56.6 | 52.7 | 52.9 | 46.4 | 57.1 | 54.4 | 58 | 56.8 | 60.8 | 56.2
| Mean |  | 62.3333 | 56.8333 | 45.8333 | 73.35 | 67.3333 | 58.3333 | 49.5 | 67 | 48.7833 | 67.0333 | 51.7167 | 52.7833 | 63.8333 | 58.45 | 62.05 | 53.75 | 54.0167 | 47.7 | 58.9167 | 51.55 | 57.8333 | 66.5167 | 58.5667 | 57.6833
| Median |  | 62.45 | 56.5 | 45.5 | 73.2 | 66.5 | 60 | 50 | 67 | 48.75 | 65.25 | 51.85 | 52.6 | 64 | 59.15 | 63.95 | 53.75 | 53.8 | 47.6 | 59.2 | 51 | 59.5 | 62.15 | 58.35 | 57.25
| Std Dev. |  | 3.4174 | 1.9408 | 1.169 | 3.4419 | 3.0111 | 2.582 | 2.51 | 0.75895 | 2.1683 | 5.2451 | 0.90646 | 2.9308 | 3.6009 | 2.9331 | 6.1187 | 1.1327 | 0.99683 | 0.8025 | 1.0226 | 1.8042 | 3.92 | 10.567 | 1.568 | 1.9934
| Rel.Std.Dev. |  | 5.482% | 3.415% | 2.551% | 4.692% | 4.472% | 4.426% | 5.071% | 1.133% | 4.445% | 7.825% | 1.753% | 5.553% | 5.641% | 5.018% | 9.861% | 2.107% | 1.845% | 1.682% | 1.736% | 3.5% | 6.778% | 15.89% | 2.677% | 3.456%
| PDM3 |  | 7.822% | -1.692% | -20.72% | 26.88% | 16.47% | 0.9028% | -14.38% | 15.89% | -15.62% | 15.95% | -10.54% | -8.697% | 10.42% | 1.105% | 7.332% | -7.025% | -6.564% | -17.49% | 1.912% | -10.83% | 0.03796% | 15.06% | 1.306% | -0.2215%
| Z-Score (Absolute) |  | 0.52 | 0.15 | 1.5 | 1.86 | 1.13 | 0.03 | 1.05 | 1.09 | 1.14 | 1.09 | 0.78 | 0.65 | 0.7 | 0.04 | 0.48 | 0.53 | 0.5 | 1.27 | 0.1 | 0.8 | 0.03 | 1.03 | 0.06 | 0.05
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C184: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C240: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C259: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C278: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C296: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C315: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C334: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C352: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C371: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C389: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C407: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C425: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C443: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C462: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C480: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C498: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C516: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C535: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C554: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C572: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C590: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C609: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C627: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C645: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C663: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C681: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C718: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C736: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C754: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C772: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C810: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C829: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C848: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C866: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C885: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C903: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C922: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C941: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C960: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C978: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C996: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1014: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1032: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1050: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1069: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1087: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1106: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1124: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1142: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 282 (Indicative Value 10.87 ppm) | [12 empty] | Indicative
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.00 | 2.03 | 2.07 | 2.08 | 2.10 | 2.12 | 2.19 | 2.21 | 2.23 | 2.26 | 2.27 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | 2
|  |  | 0.085g | 30g | 25g | 40g | 10g | 25g | 25g | 30g | 30g | 25g | 15g | 2
| 1 | 1 | 13.7119 |  | 9.18 | N.A. | 10.64 | > 4 |  | NA | 11.0397 | 13 | N.A. | 1
| 1 | 2 | 14.1603 |  | 9.25 | N.A. | 10.34 | > 4 |  | NA | 11.1663 | 12.8 | N.A. | 5
| 1 | 3 | 14.2625 |  | 8.965 | N.A. | 10.49 | > 4 |  | NA | 10.8946 | 12.9 | N.A. | 16
| 1 | 4 | 14.2924 |  | 9.12 | N.A. | 10.54 | > 4 |  | NA | 10.7309 | 13.1 | N.A. | 10.8661
| 1 | 5 | 13.9612 |  | 9.12 | N.A. | 10.21 | > 4 |  | NA | 10.8173 | 13 | N.A. | 11
| 1 | 6 | 13.8885 |  | 9.21 | N.A. | 10.54 | > 4 |  | NA | 10.8321 | 12.9 | N.A.
|  | 7 | 14.0546
|  | 8 | 14.021
|  | 9 | 14.1988
|  | 10 | 13.9041
|  | 11 | 13.6999
|  | 12 | 13.9543
|  | 13 | 14.0465
|  | 14 | 14.002
|  | 15 | 13.7712
|  | 16 | 13.7865
|  | 17 | 14.2296
|  | 18 | 14.0894
|  | 19 | 14.0851
|  | 20 | 14.0355
| Mean |  | 14.0078 |  | 9.1408 |  | 10.46 |  |  |  | 10.9135 | 12.95
| Median |  | 14.0283 |  | 9.15 |  | 10.515 |  |  |  | 10.8634 | 12.95
| Std Dev. |  | 0.17613 |  | 0.10002 |  | 0.15684 |  |  |  | 0.16088 | 0.10488
| Rel.Std.Dev. |  | 1.257% |  | 1.094% |  | 1.499% |  |  |  | 1.474% | 0.8099%
| PDM3 |  | 28.91% |  | -15.88% |  | -3.737% |  |  |  | 0.4363% | 19.18%
| Z-Score (Absolute) |  | NA | NA | 1.18 | NA | 0.17 | NA | NA | NA | 0.17 | 1.72 | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 282 (Certified Value 0.577 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 0.6 | 0.5 | 0.5 | 0.5 | 0.6 | 0.45 | 0.6 | 0.59 | 0.6 | 0.57 | 0.487 | 0.56 | 0.61 | 0.58 | 0.6 | 0.57 | 0.55 | 0.606 | < 2 | 0.563 | 0.8 | 0.65 | 0.58 | 0.57 |  | 1
| 1 | 2 | 0.6 | 0.5 | 0.6 | 0.5 | 0.6 | 0.5 | 0.5 | 0.62 | 0.58 | 0.56 | 0.494 | 0.6 | 0.586 | 0.57 | 0.5 | 0.59 | 0.57 | 0.606 | < 2 | 0.573 | 0.9 | 0.62 | 0.58 | 0.52 |  | 19
| 1 | 3 | 0.57 | 0.45 | 0.6 | 0.5 | 0.5 | 0.5 | 0.5 | 0.57 | 0.63 | 0.56 | 0.49 | 0.56 | 0.609 | 0.58 | 0.6 | 0.58 | 0.56 | 0.592 | < 2 | 0.563 | 0.7 | 0.64 | 0.57 | 0.58 |  | 16
| 1 | 4 | 0.61 | 0.5 | 0.6 | 0.4 | 0.6 | 0.5 | 0.5 | 0.63 | 0.6 | 0.57 | 0.475 | 0.57 | 0.573 | 0.58 | 0.5 | 0.58 | 0.57 | 0.593 | < 2 | 0.552 | 0.8 | 0.62 | 0.56 | 0.54 |  | 0.57659
| 1 | 5 | 0.6 | 0.7 | 0.6 | 0.5 | 0.6 | 0.5 | 0.5 | 0.6 | 0.64 | 0.57 | 0.482 | 0.57 | 0.584 | 0.57 | 0.6 | 0.59 | 0.59 | 0.593 | < 2 | 0.534 | 0.7 | 0.6 | 0.56 | 0.55 |  | 71
| 1 | 6 | 0.57 | 0.5 | 0.6 | 0.6 | 0.6 | 0.5 | 0.5 | 0.54 | 0.64 | 0.56 | 0.476 | 0.56 | 0.621 | 0.58 | 0.7 | 0.58 | 0.59 | 0.58 | < 2 | 0.521 | 0.9 | 0.6 | 0.57 | 0.58
| Mean |  | 0.59167 | 0.525 | 0.58333 | 0.5 | 0.58333 | 0.49167 | 0.51667 | 0.59167 | 0.615 | 0.565 | 0.484 | 0.57 | 0.59717 | 0.57667 | 0.58333 | 0.58167 | 0.57167 | 0.595 |  | 0.551 | 0.8 | 0.62167 | 0.57 | 0.55667
| Median |  | 0.6 | 0.5 | 0.6 | 0.5 | 0.6 | 0.5 | 0.5 | 0.595 | 0.615 | 0.565 | 0.4845 | 0.565 | 0.5975 | 0.58 | 0.6 | 0.58 | 0.57 | 0.593 |  | 0.5575 | 0.8 | 0.62 | 0.57 | 0.56
| Std Dev. |  | 0.017224 | 0.088034 | 0.040825 | 0.063246 | 0.040825 | 0.020412 | 0.040825 | 0.033116 | 0.0251 | 0.0054772 | 0.0076681 | 0.015492 | 0.018734 | 0.005164 | 0.075277 | 0.0075277 | 0.016021 | 0.0098387 |  | 0.019809 | 0.089443 | 0.020412 | 0.0089443 | 0.024221
| Rel.Std.Dev. |  | 2.911% | 16.77% | 6.999% | 12.65% | 6.999% | 4.152% | 7.902% | 5.597% | 4.081% | 0.9694% | 1.584% | 2.718% | 3.137% | 0.8955% | 12.9% | 1.294% | 2.802% | 1.654% |  | 3.595% | 11.18% | 3.283% | 1.569% | 4.351%
| PDM3 |  | 2.615% | -8.947% | 1.17% | -13.28% | 1.17% | -14.73% | -10.39% | 2.615% | 6.662% | -2.01% | -16.06% | -1.143% | 3.569% | 0.01349% | 1.17% | 0.8807% | -0.8537% | 3.193% |  | -4.438% | 38.75% | 7.818% | -1.143% | -3.455%
| Z-Score (Absolute) |  | 0.62 | 1.73 | NA | NA | NA | 2.9 | NA | 0.62 | 1.44 | 0.32 | 3.17 | 0.15 | 0.81 | 0.09 | NA | 0.26 | 0.09 | 0.73 | 14.98 | 0.82 | NA | 1.67 | 0.15 | 0.62
Results from laboratories 3, 4, 7, 9, 19 and 28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 282 (Certified Value 0.815 wt.%) | [26 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.85 | 0.79 | 0.84 | 0.77 | 0.81 | 0.6555 | 0.76 | 0.84 | 0.781 | 1.035 | 0.8466 | 0.9675 | 1.28 | 0.85 | 0.751 | 0.77 | 0.989 | 0.84 | 0.73 | 0.85 | 0.91 | 0.81 | 0.81 |  |  | 1
| 1 | 2 | 0.84 | 0.74 | 0.84 | 0.78 | 0.81 | 0.683 | 0.78 | 0.86 | 0.779 | 1.053 | 0.8358 | 1.0177 | 1.28 | 0.86 | 0.738 | 0.75 | 0.992 | 0.86 | 0.74 | 0.85 | 0.86 | 0.81 | 0.8 |  |  | #N/A
| 1 | 3 | 0.84 | 0.75 | 0.83 | 0.76 | 0.83 | 0.652 | 0.78 | 0.83 | 0.751 | 1.058 | 0.8571 | 0.98 | 1.37 | 0.84 | 0.74 | 0.75 | 0.997 | 0.84 | 0.71 | 0.84 | 0.85 | 0.83 | 0.8 |  |  | 16
| 1 | 4 | 0.85 | 0.79 | 0.81 | 0.78 | 0.81 | 0.6705 | 0.77 | 0.83 | 0.772 | 1.028 | 0.8497 | 0.9458 | 1.22 | 0.83 | 0.733 | 0.75 | 0.991 | 0.84 | 0.74 | 0.8 | 0.86 | 0.81 | 0.81 |  |  | 0.81479
| 1 | 5 | 0.87 | 0.78 | 0.81 | 0.77 | 0.82 | 0.676 | 0.75 | 0.86 | 0.786 | 1.047 | 0.8461 | 0.9443 | 1.27 | 0.85 | 0.751 | 0.77 | 0.986 | 0.86 | 0.74 | 0.81 | 0.83 | 0.8 | 0.8 |  |  | 72
| 1 | 6 | 0.87 | 0.74 | 0.83 | 0.77 | 0.83 | 0.664 | 0.76 | 0.83 | 0.763 | 1.025 | 0.8495 | 0.9747 | 1.28 | 0.83 | 0.744 | 0.74 | 0.992 | 0.86 | 0.79 | 0.84 | 0.85 | 0.79 | 0.81
| Mean |  | 0.85333 | 0.765 | 0.82667 | 0.77167 | 0.81833 | 0.66683 | 0.76667 | 0.84167 | 0.772 | 1.041 | 0.84747 | 0.97167 | 1.2833 | 0.84333 | 0.74283 | 0.755 | 0.99117 | 0.85 | 0.74167 | 0.83167 | 0.86 | 0.80833 | 0.805
| Median |  | 0.85 | 0.765 | 0.83 | 0.77 | 0.815 | 0.66725 | 0.765 | 0.835 | 0.7755 | 1.041 | 0.84805 | 0.9711 | 1.28 | 0.845 | 0.742 | 0.75 | 0.9915 | 0.85 | 0.74 | 0.84 | 0.855 | 0.81 | 0.805
| Std Dev. |  | 0.013663 | 0.02429 | 0.013663 | 0.0075277 | 0.0098319 | 0.011961 | 0.012111 | 0.01472 | 0.013023 | 0.013638 | 0.0069376 | 0.026956 | 0.048442 | 0.012111 | 0.0072503 | 0.012247 | 0.003656 | 0.010954 | 0.026394 | 0.02137 | 0.026833 | 0.013292 | 0.0054772
| Rel.Std.Dev. |  | 1.601% | 3.175% | 1.653% | 0.9755% | 1.201% | 1.794% | 1.58% | 1.749% | 1.687% | 1.31% | 0.8186% | 2.774% | 3.775% | 1.436% | 0.976% | 1.622% | 0.3689% | 1.289% | 3.559% | 2.57% | 3.12% | 1.644% | 0.6804%
| PDM3 |  | 4.73% | -6.111% | 1.457% | -5.293% | 0.4346% | -18.16% | -5.906% | 3.298% | -5.252% | 27.76% | 4.01% | 19.25% | 57.5% | 3.503% | -8.832% | -7.338% | 21.65% | 4.321% | -8.975% | 2.071% | 5.548% | -0.7927% | -1.202%
| Z-Score (Absolute) |  | 0.33 | 0.76 | 0 | 0.68 | 0.1 | 1.97 | 0.74 | 0.19 | 0.67 | 2.64 | 0.26 | 1.79 | 5.63 | 0.21 | 1.03 | 0.88 | 2.03 | 0.29 | 1.05 | 0.06 | 0.41 | 0.23 | 0.27
| Analytical results for As in OREAS 282 (Certified Value 2320 ppm) | [26 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | 0
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0
| 1 | 1 | 2320 | 2450 | 2285 | 2459 | 1790 | 2231 | 2240 | > 1000 | 2386.86 | 2690.4 | 2355 | 2345 | 2259 | 2890 | 2156 | 2450 | 2330 | 2240 | 2548.4 | 2160 | 2028.6 | 2555 | 2590 | 2326 | 2260 | 1
| 1 | 2 | 2290 | 2430 | 2313 | 2391 | 1760 | 2247 | 2230 | > 1000 | 2376.47 | 2641.5 | 2254 | 2357 | 2346 | 2830 | 2150 | 2440 | 2350 | 2200 | 2531.7 | 2240 | 2048.2 | 2519 | 2470 | 2279.2 | 2270 | 20
| 1 | 3 | 2290 | 2460 | 2246 | 2435 | 1860 | 2201 | 2250 | > 1000 | 2317.62 | 2687.3 | 2344 | 2358 | 2301 | 2910 | 2198 | 2420 | 2320 | 2210 | 2526.6 | 2120 | 2028.6 | 2557 | 2470 | 2289.1 | 2270 | 16
| 1 | 4 | 2290 | 2460 | 2284 | 2469 | 1940 | 2228 | 2360 | > 1000 | 2327.02 | 2779.2 | 2230 | 2352 | 2355 | 2730 | 2185 | 2300 | 2350 | 2180 | 2498.8 | 2210 | 2009 | 2578 | 2510 | 2272 | 2310 | 2319.5166
| 1 | 5 | 2340 | 2490 | 2223 | 2451 | 1880 | 2226 | 2350 | > 1000 | 2384.52 | 2701.5 | 2254 | 2336 | 2296 | 2810 | 2168 | 2310 | 2370 | 2170 | 2523 | 2220 | 2009 | 2492 | 2440 | 2279.9 | 2250 | 73
| 1 | 6 | 2290 | 2460 | 2282 | 2443 | 1800 | 2230 | 2240 | > 1000 | 2154.72 | 2730.8 | 2201 | 2296 | 2304 | 2880 | 2201 | 2470 | 2310 | 2180 | 2511.2 | 2150 | 2040 | 2571 | 2530 | 2271.8 | 2260
| Mean |  | 2303.3333 | 2458.3333 | 2272.1667 | 2441.3333 | 1838.3333 | 2227.1667 | 2278.3333 |  | 2324.535 | 2705.1167 | 2273 | 2340.6667 | 2310.1667 | 2841.6667 | 2176.3333 | 2398.3333 | 2338.3333 | 2196.6667 | 2523.2833 | 2183.3333 | 2027.2333 | 2545.3333 | 2501.6667 | 2286.3333 | 2270
| Median |  | 2290 | 2460 | 2283 | 2447 | 1830 | 2229 | 2245 |  | 2351.745 | 2695.95 | 2254 | 2348.5 | 2302.5 | 2855 | 2176.5 | 2430 | 2340 | 2190 | 2524.8 | 2185 | 2028.6 | 2556 | 2490 | 2279.55 | 2265
| Std Dev. |  | 21.6025 | 19.4079 | 32.1585 | 27.3764 | 67.0572 | 14.8515 | 59.8052 |  | 88.3979 | 46.3512 | 62.482 | 23.3723 | 35.3124 | 66.458 | 21.5839 | 74.1395 | 22.286 | 25.8199 | 17.0775 | 46.7618 | 15.9451 | 33.1461 | 53.8207 | 20.4452 | 20.9762
| Rel.Std.Dev. |  | 0.9379% | 0.7895% | 1.415% | 1.121% | 3.648% | 0.6668% | 2.625% |  | 3.803% | 1.713% | 2.749% | 0.9985% | 1.529% | 2.339% | 0.9918% | 3.091% | 0.9531% | 1.175% | 0.6768% | 2.142% | 0.7865% | 1.302% | 2.151% | 0.8942% | 0.9241%
| PDM3 |  | -0.6977% | 5.985% | -2.041% | 5.252% | -20.74% | -3.981% | -1.776% |  | 0.2164% | 16.62% | -2.005% | 0.9118% | -0.4031% | 22.51% | -6.173% | 3.398% | 0.8112% | -5.296% | 8.785% | -5.871% | -12.6% | 9.736% | 7.853% | -1.431% | -2.135%
| Z-Score (Absolute) |  | 0.02 | 1.01 | 0.23 | 0.9 | 3.13 | 0.53 | 0.19 | NA | 0.12 | 2.66 | 0.23 | 0.23 | 0.02 | 3.58 | 0.87 | 0.61 | 0.21 | 0.74 | 1.45 | 0.83 | 1.87 | 1.6 | 1.3 | 0.14 | 0.25
| Analytical results for B in OREAS 282 (Indicative Value 10 ppm) | [26 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.10 | 2.13 | 2.16 | 2.19 | 2.20 | 2.23 | 2.27 | 2.29 | 2.30 | 2.31 | [12 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.4g | 0.5g | 30g | 0.5g | 0.5g | 0.5g | 0.5g | [12 empty] | 1
| 1 | 1 | 10 | 15 | < 10 | 10 | 17 | 22 |  | 10 | 8 | < 20 | 10 | < 10 | 10 | [12 empty] | 1
| 1 | 2 | 10 | 16 | < 10 | 10 | 17 | 20 |  | 10 | 10 | < 20 | 10 | < 10 | 10 | [12 empty] | 3
| 1 | 3 | 10 | 15 | < 10 | 10 | 18 | 21 |  | 10 | 10 | < 20 | 10 | 10 | 10 | [12 empty] | 16
| 1 | 4 | 10 | 15 | < 10 | 10 | 17 | 20 |  | 10 | 10 | < 20 | 10 | < 10 | 10 | [12 empty] | 10
| 1 | 5 | 10 | 15 | < 10 | 11 | 17 | 21 |  | 10 | 10 | < 20 | 10 | < 10 | 10 | [12 empty] | 13
| 1 | 6 | 10 | 17 | < 10 | 10 | 17 | 20 |  | 10 | 9 | < 20 | 10 | < 10 | 10
| Mean |  | 10 | 15.5 |  | 10.1667 | 17.1667 | 20.6667 |  | 10 | 9.5 |  | 10 | 10 | 10
| Median |  | 10 | 15 |  | 10 | 17 | 20.5 |  | 10 | 10 |  | 10 | 10 | 10
| Std Dev. |  | 0 | 0.83666 |  | 0.40825 | 0.40825 | 0.8165 |  | 0 | 0.83666 |  | 0 |  | 0
| Rel.Std.Dev. |  | 0% | 5.398% |  | 4.016% | 2.378% | 3.951% |  | 0% | 8.807% |  | 0% |  | 0%
| PDM3 |  | 0% | 55% |  | 1.667% | 71.67% | 106.7% |  | 0% | -5% |  | 0% | 0% | 0%
| Z-Score (Absolute) |  | 0 | 11.13 | 10.11 | 0.34 | 14.5 | 21.58 | NA | 0 | 1.01 | 0 | 0 | 8.43 | 0
| Analytical results for Ba in OREAS 282 (Certified Value 29.8 ppm) | [26 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.25g | 0.25g | 25g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 30 | 40 | 32 | 32.4 | 27 | 26 | 33.29 | 30 | 31 | 25.87 | 31 | 32.2 | 29 | 30 | 30 | 30 | 33.4 | 26.8 | 31 | 30 | 29 | 30 |  |  |  | 1
| 1 | 2 | 30 | 39 | 33 | 31.8 | 27 | 26 | 33.12 | 29 | 30 | 26.88 | 31 | 30.7 | 28 | 27 | 30 | 32 | 35.1 | 26.5 | 31 | 30 | 27 | 30 |  |  |  | 21
| 1 | 3 | 30 | 38 | 31 | 32.6 | 26 | 26 | 32.68 | 29 | 31 | 26.53 | 31 | 33.7 | 29 | 29 | 30 | 32 | 32.6 | 26.3 | 31 | 30 | 30 | 30 |  |  |  | 16
| 1 | 4 | 30 | 40 | 29 | 32.8 | 27 | 27 | 33.22 | 29 | 30 | 26.56 | 31 | 29.3 | 28 | 26 | 30 | 32 | 32.6 | 26.7 | 31 | 30 | 29 | 30 |  |  |  | 29.7765
| 1 | 5 | 30 | 39 | 31 | 33.1 | 27 | 27 | 32.98 | 29 | 30 | 25.47 | 30 | 30.7 | 29 | 30 | 30 | 32 | 33.5 | 26.3 | 31 | 30 | 29 | 30 |  |  |  | 74
| 1 | 6 | 30 | 41 | 32 | 31.3 | 26 | 26 | 33.44 | 29 | 30 | 26.22 | 30 | 32.9 | 29 | 28 | 30 | 31 | 33 | 27.3 | 31 | 30 | 28 | 30
| Mean |  | 30 | 39.5 | 31.3333 | 32.3333 | 26.6667 | 26.3333 | 33.1217 | 29.1667 | 30.3333 | 26.255 | 30.6667 | 31.5833 | 28.6667 | 28.3333 | 30 | 31.5 | 33.3667 | 26.65 | 31 | 30 | 28.6667 | 30
| Median |  | 30 | 39.5 | 31.5 | 32.5 | 27 | 26 | 33.17 | 29 | 30 | 26.375 | 31 | 31.45 | 29 | 28.5 | 30 | 32 | 33.2 | 26.6 | 31 | 30 | 29 | 30
| Std Dev. |  | 0 | 1.0488 | 1.3663 | 0.66833 | 0.5164 | 0.5164 | 0.26626 | 0.40825 | 0.5164 | 0.51477 | 0.5164 | 1.6351 | 0.5164 | 1.633 | 0 | 0.83666 | 0.93095 | 0.37815 | 0 | 0 | 1.0328 | 0
| Rel.Std.Dev. |  | 0% | 2.655% | 4.36% | 2.067% | 1.936% | 1.961% | 0.8039% | 1.4% | 1.702% | 1.961% | 1.684% | 5.177% | 1.801% | 5.764% | 0% | 2.656% | 2.79% | 1.419% | 0% | 0% | 3.603% | 0%
| PDM3 |  | 0.7507% | 32.66% | 5.228% | 8.587% | -10.44% | -11.56% | 11.23% | -2.048% | 1.87% | -11.83% | 2.99% | 6.068% | -3.727% | -4.847% | 0.7507% | 5.788% | 12.06% | -10.5% | 4.109% | 0.7507% | -3.727% | 0.7507%
| Z-Score (Absolute) |  | 0.13 | 3.71 | 0.4 | 0.81 | 1.48 | 1.62 | 1.13 | 0.47 | 0 | 1.65 | 0.13 | 0.51 | 0.67 | 0.81 | NA | 0.47 | 1.23 | 1.49 | 0.27 | NA | 0.67 | NA
Results from laboratories 20, 29 and 31 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 282 (Certified Value 0.67 ppm) | [26 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.29 | 2.30 | 2.31 | [6 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.25g | 0.25g | 25g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.64 | < 5 | 0.6 | 0.5 | 0.7 | < 1 | 0.66 | 0.56 | 0.6 | 0.663 | 0.74 | 0.8 | 0.6 | 0.8 | 0.7 | 0.77 | 0.76 | 0.61 | 0.65 | [6 empty] | 1
| 1 | 2 | 0.68 | < 5 | 0.7 | 0.5 | 0.8 | < 1 | 0.66 | 0.66 | 0.6 | 0.654 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.78 | 0.66 | 0.58 | 0.63 | [6 empty] | 22
| 1 | 3 | 0.66 | < 5 | 0.6 | 0.5 | 0.7 | < 1 | 0.65 | 0.62 | 0.7 | 0.67 | 0.68 | 0.8 | 0.7 | 0.7 | 0.71 | 0.8 | 0.75 | 0.62 | 0.64 | [6 empty] | 16
| 1 | 4 | 0.67 | < 5 | 0.7 | 0.5 | 0.7 | < 1 | 0.67 | 0.68 | 0.6 | 0.661 | 0.72 | 0.7 | 0.7 | 0.8 | 0.73 | 0.78 | 0.73 | 0.62 | 0.58 | [6 empty] | 0.66693
| 1 | 5 | 0.65 | < 5 | 0.7 | 0.5 | 0.7 | < 1 | 0.65 | 0.68 | 0.7 | 0.652 | 0.7 | 0.8 | 0.7 | 0.8 | 0.74 | 0.8 | 0.68 | 0.61 | 0.62 | [6 empty] | 75
| 1 | 6 | 0.65 | < 5 | 0.6 | 0.5 | 0.7 | < 1 | 0.62 | 0.64 | 0.6 | 0.658 | 0.66 | 0.8 | 0.7 | 0.8 | 0.75 | 0.8 | 0.71 | 0.62 | 0.72
| Mean |  | 0.65833 |  | 0.65 | 0.5 | 0.71667 |  | 0.65167 | 0.64 | 0.63333 | 0.65967 | 0.7 | 0.76667 | 0.68333 | 0.76667 | 0.72167 | 0.78833 | 0.715 | 0.61 | 0.64
| Median |  | 0.655 |  | 0.65 | 0.5 | 0.7 |  | 0.655 | 0.65 | 0.6 | 0.6595 | 0.7 | 0.8 | 0.7 | 0.8 | 0.72 | 0.79 | 0.72 | 0.615 | 0.635
| Std Dev. |  | 0.01472 |  | 0.054772 | 0 | 0.040825 |  | 0.017224 | 0.045607 | 0.05164 | 0.006532 | 0.028284 | 0.05164 | 0.040825 | 0.05164 | 0.02137 | 0.013292 | 0.03937 | 0.015492 | 0.046043
| Rel.Std.Dev. |  | 2.236% |  | 8.427% | 0% | 5.696% |  | 2.643% | 7.126% | 8.154% | 0.9902% | 4.041% | 6.736% | 5.974% | 6.736% | 2.961% | 1.686% | 5.506% | 2.54% | 7.194%
| PDM3 |  | -1.288% |  | -2.538% | -25.03% | 7.458% |  | -2.288% | -4.037% | -5.037% | -1.088% | 4.959% | 14.96% | 2.46% | 14.96% | 8.208% | 18.2% | 7.208% | -8.536% | -4.037%
| Z-Score (Absolute) |  | 0.01 | 27.59 | NA | NA | NA | 2.38 | 0.11 | 0.28 | NA | 0.01 | 0.61 | NA | NA | NA | 0.94 | 1.94 | 0.84 | 0.73 | 0.28
Results from laboratories 4, 6, 7, 13, 16, 18 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 282 (Certified Value 0.91 ppm) | [26 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.94 | 0.9 | < 5 | 0.92 | 0.8 | 0.8 | 0.8 | 0.92 | 0.93 | 0.91 | 0.892 | 0.93 | 0.98 | 0.89 | 0.9 | 0.86 | 0.93 | 1.06 | 0.86 | < 5 | 0.97 | 0.89 | 0.95 |  |  | 1
| 1 | 2 | 0.93 | 0.88 | < 5 | 0.86 | 0.9 | 0.8 | 0.8 | 0.92 | 0.87 | 0.89 | 0.922 | 0.97 | 0.93 | 0.91 | 0.9 | 0.89 | 0.91 | 1.07 | 0.88 | < 5 | 0.92 | 0.87 | 0.88 |  |  | 23
| 1 | 3 | 0.92 | 0.88 | < 5 | 0.91 | 0.8 | 0.9 | 0.8 | 0.91 | 0.87 | 0.91 | 0.919 | 0.96 | 0.92 | 0.91 | 0.8 | 0.89 | 0.95 | 1 | 0.88 | < 5 | 0.93 | 0.92 | 0.9 |  |  | 16
| 1 | 4 | 0.91 | 0.92 | < 5 | 0.94 | 0.8 | 0.9 | 0.8 | 0.91 | 0.89 | 0.89 | 0.902 | 0.99 | 0.9 | 0.9 | 0.8 | 0.86 | 0.9 | 1.01 | 0.89 | < 5 | 0.95 | 0.9 | 0.87 |  |  | 0.90693
| 1 | 5 | 0.9 | 0.92 | < 5 | 0.86 | 0.8 | 0.9 | 0.8 | 0.92 | 0.89 | 0.9 | 0.893 | 0.95 | 0.92 | 0.91 | 0.9 | 0.89 | 0.93 | 1.02 | 0.84 | < 5 | 0.87 | 0.91 | 0.88 |  |  | 76
| 1 | 6 | 0.91 | 0.88 | < 5 | 0.92 | 0.8 | 0.9 | 0.8 | 0.85 | 0.9 | 0.89 | 0.906 | 0.95 | 0.93 | 0.93 | 0.9 | 0.88 | 0.94 | 1.02 | 0.88 | < 5 | 0.92 | 0.88 | 0.95
| Mean |  | 0.91833 | 0.89667 |  | 0.90167 | 0.81667 | 0.86667 | 0.8 | 0.905 | 0.89167 | 0.89833 | 0.90567 | 0.95833 | 0.93 | 0.90833 | 0.86667 | 0.87833 | 0.92667 | 1.03 | 0.87167 |  | 0.92667 | 0.895 | 0.905
| Median |  | 0.915 | 0.89 |  | 0.915 | 0.8 | 0.9 | 0.8 | 0.915 | 0.89 | 0.895 | 0.904 | 0.955 | 0.925 | 0.91 | 0.9 | 0.885 | 0.93 | 1.02 | 0.88 |  | 0.925 | 0.895 | 0.89
| Std Dev. |  | 0.01472 | 0.019664 |  | 0.033714 | 0.040825 | 0.05164 | 1.2162e-16 | 0.027386 | 0.022286 | 0.0098319 | 0.012691 | 0.020412 | 0.026833 | 0.013292 | 0.05164 | 0.01472 | 0.018619 | 0.028284 | 0.018348 |  | 0.033862 | 0.018708 | 0.036194
| Rel.Std.Dev. |  | 1.603% | 2.193% |  | 3.739% | 4.999% | 5.958% | 1.52e-14% | 3.026% | 2.499% | 1.094% | 1.401% | 2.13% | 2.885% | 1.463% | 5.958% | 1.676% | 2.009% | 2.746% | 2.105% |  | 3.654% | 2.09% | 3.999%
| PDM3 |  | 1.258% | -1.131% |  | -0.5799% | -9.952% | -4.439% | -11.79% | -0.2124% | -1.683% | -0.9474% | -0.1388% | 5.668% | 2.544% | 0.1552% | -4.439% | -3.153% | 2.177% | 13.57% | -3.888% |  | 2.177% | -1.315% | -0.2124%
| Z-Score (Absolute) |  | 0.41 | 0.26 | 49.64 | 0.1 | 2.75 | 1.19 | 3.27 | 0 | 0.41 | 0.21 | 0.02 | 1.66 | 0.78 | 0.1 | 1.19 | 0.83 | 0.67 | 3.89 | 1.04 | 49.64 | 0.67 | 0.31 | 0
| Analytical results for Ca in OREAS 282 (Certified Value 16.35 wt.%) | [26 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 17.55 | 17.8 | 15.41 | > 15 | 16.0425 | 17.9 | 17 | 17.849 | 14.34 | 15.08 | 22 | > 15 | 16.3 | 17.1 | 15.6 | 16.52 | 14.75 | > 10 | 18.3 | 16.34 | 16.3 |  |  |  |  | 1
| 1 | 2 | 17.1 | 17.8 | 15.05 | > 15 | 16.2405 | 17.6 | 17.09 | 17.828 | 14.18 | 15.54 | 22 | > 15 | 16 | 17 | 15.4 | 16.31 | 14.44 | > 10 | 18.15 | 15.97 | 16.55 |  |  |  |  | #N/A
| 1 | 3 | 17.15 | 17.1 | 14.92 | > 15 | 15.673 | 17.8 | 16.6 | 17.29 | 14.39 | 15.22 | 22.6 | > 15 | 16.2 | 16.95 | 15 | 16.46 | 14.54 | > 10 | 17.6 | 15.85 | 16.75 |  |  |  |  | 16
| 1 | 4 | 17.35 | 17.8 | 15.23 | > 15 | 15.8385 | 17.8 | 17.13 | 17.797 | 14.1 | 15.77 | 21.1 | > 15 | 17 | 17.1 | 15.2 | 16.11 | 14.54 | > 10 | 17.85 | 15.76 | 16.95 |  |  |  |  | 16.3536
| 1 | 5 | 17.4 | 17.3 | 14.83 | > 15 | 15.8805 | 17.8 | 17.18 | 17.995 | 14.32 | 15.3 | 21.7 | > 15 | 16.5 | 17.35 | 15.5 | 16.35 | 14.54 | > 10 | 17.1 | 15.86 | 16.7 |  |  |  |  | 77
| 1 | 6 | 17.25 | 17.6 | 15.18 | > 15 | 15.9085 | 17.6 | 16.89 | 17.56 | 14.12 | 15.59 | 22.1 | > 15 | 16.7 | 16.95 | 15.4 | 16.29 | 15.7 | > 10 | 17.65 | 15.68 | 16.5
| Mean |  | 17.3 | 17.5667 | 15.1033 |  | 15.9306 | 17.75 | 16.9817 | 17.7198 | 14.2417 | 15.4167 | 21.9167 |  | 16.45 | 17.075 | 15.35 | 16.34 | 14.7517 |  | 17.775 | 15.91 | 16.625
| Median |  | 17.3 | 17.7 | 15.115 |  | 15.8945 | 17.8 | 17.045 | 17.8125 | 14.25 | 15.42 | 22 |  | 16.4 | 17.05 | 15.4 | 16.33 | 14.54 |  | 17.75 | 15.855 | 16.625
| Std Dev. |  | 0.16733 | 0.30111 | 0.21314 |  | 0.19308 | 0.12247 | 0.2133 | 0.25312 | 0.12368 | 0.25913 | 0.49565 |  | 0.36194 | 0.15083 | 0.21679 | 0.14367 | 0.47558 |  | 0.42983 | 0.23238 | 0.22528
| Rel.Std.Dev. |  | 0.9672% | 1.714% | 1.411% |  | 1.212% | 0.69% | 1.256% | 1.428% | 0.8684% | 1.681% | 2.262% |  | 2.2% | 0.8833% | 1.412% | 0.8792% | 3.224% |  | 2.418% | 1.461% | 1.355%
| PDM3 |  | 5.787% | 7.418% | -7.645% |  | -2.587% | 8.539% | 3.84% | 8.354% | -12.91% | -5.729% | 34.02% |  | 0.5894% | 4.411% | -6.137% | -0.08324% | -9.796% |  | 8.692% | -2.713% | 1.659%
| Z-Score (Absolute) |  | 0.48 | 0.65 | 0.9 | NA | 0.38 | 0.76 | 0.28 | 0.74 | 1.44 | 0.7 | 3.37 | NA | 0.05 | 0.34 | 0.74 | 0.12 | 1.12 | NA | 0.78 | 0.39 | 0.05
| Analytical results for Cd in OREAS 282 (Certified Value 1.09 ppm) | [26 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.15 | 1.05 | 14 | 1.14 | 1 | 1 | 1.1 | 1.08 | 1.12 | 0.99 | 1.001 | 1.14 | 1.19 | 1.11 | 1.1 | 1.05 | 1.12 | 1.21 | 1.07 | 1 | 1.1 | 1.05 | 1.11 |  |  | 1
| 1 | 2 | 1.16 | 1 | 15 | 1.13 | 1 | 1 | 1.1 | 1.094 | 1.09 | 0.97 | 1.038 | 1.11 | 1.19 | 1.15 | 1 | 1.08 | 1.02 | 1.25 | 1.14 | 1 | 1.01 | 1.04 | 1.04 |  |  | 24
| 1 | 3 | 1.13 | 1 | 14 | 1.13 | 1 | 1 | 1.1 | 1.039 | 1.09 | 1.02 | 1.044 | 1.12 | 0.97 | 1.09 | 1 | 1.03 | 1.1 | 1.23 | 1.11 | < 1 | 1.03 | 1.07 | 1.08 |  |  | 16
| 1 | 4 | 1.18 | 1 | 14 | 1.12 | 1 | 1 | 1.1 | 1.102 | 1.11 | 1.02 | 1.021 | 1.18 | 1.13 | 1.09 | 1 | 1.05 | 1.08 | 1.23 | 1.12 | 1 | 1.05 | 1.07 | 1.06 |  |  | 1.0933
| 1 | 5 | 1.15 | 1.05 | 15 | 1.15 | 1 | 1 | 1.1 | 1.085 | 1.13 | 0.96 | 1.012 | 1.09 | 1.22 | 1.12 | 1 | 1.08 | 1.09 | 1.19 | 1.08 | < 1 | 1.01 | 1.08 | 1.14 |  |  | 78
| 1 | 6 | 1.11 | 1 | 15 | 1.09 | 1 | 1.1 | 1.1 | 0.986 | 1.15 | 0.95 | 1.002 | 1.13 | 1.25 | 1.11 | 1 | 1.04 | 1.11 | 1.26 | 0.98 | 1 | 1.05 | 1.03 | 1.14
| Mean |  | 1.1467 | 1.0167 | 14.5 | 1.1267 | 1 | 1.0167 | 1.1 | 1.0643 | 1.115 | 0.985 | 1.0197 | 1.1283 | 1.1583 | 1.1117 | 1.0167 | 1.055 | 1.0867 | 1.2283 | 1.0833 | 1 | 1.0417 | 1.0567 | 1.095
| Median |  | 1.15 | 1 | 14.5 | 1.13 | 1 | 1 | 1.1 | 1.0825 | 1.115 | 0.98 | 1.0165 | 1.125 | 1.19 | 1.11 | 1 | 1.05 | 1.095 | 1.23 | 1.095 | 1 | 1.04 | 1.06 | 1.095
| Std Dev. |  | 0.024221 | 0.02582 | 0.54772 | 0.020656 | 0 | 0.040825 | 2.4324e-16 | 0.044157 | 0.023452 | 0.030166 | 0.018162 | 0.030605 | 0.10048 | 0.022286 | 0.040825 | 0.020736 | 0.03559 | 0.025626 | 0.056804 | 0 | 0.033714 | 0.019664 | 0.041833
| Rel.Std.Dev. |  | 2.112% | 2.54% | 3.777% | 1.833% | 0% | 4.016% | 2.211e-14% | 4.149% | 2.103% | 3.063% | 1.781% | 2.712% | 8.675% | 2.005% | 4.016% | 1.966% | 3.275% | 2.086% | 5.243% | 0% | 3.237% | 1.861% | 3.82%
| PDM3 |  | 4.884% | -7.007% | 1226% | 3.054% | -8.532% | -7.007% | 0.6152% | -2.647% | 1.987% | -9.904% | -6.733% | 3.207% | 5.951% | 1.682% | -7.007% | -3.501% | -0.6044% | 12.35% | -0.9093% | -8.532% | -4.72% | -3.348% | 0.1578%
| Z-Score (Absolute) |  | 1.03 | 1.36 | 246.6 | 0.66 | NA | NA | NA | 0.49 | 0.44 | 1.95 | 1.31 | 0.69 | 1.24 | 0.38 | NA | 0.66 | 0.08 | 2.53 | 0.14 | NA | 0.9 | 0.63 | 0.08
Results from laboratories 7, 8, 9, 19 and 28 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for Ce in OREAS 282 (Certified Value 33.8 ppm) | [26 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.28 | 2.29 | 2.30 | 2.31 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 1
|  |  | 0.5g | 04g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 35 | 36.9 | 30.42 | 31.5 | 32.8 | 29 | 34.508 | 33.91 | 31.65 | 32.95 | 36.31 | 36.9 | 32.62 | 31.7 | 34.3 | 36.8 | 36 | 36.5 | 34.91 | 35.4 | [5 empty] | 1
| 1 | 2 | 34.6 | 36.2 | 30.58 | 37 | 32 | 29 | 34.5 | 32.1 | 31 | 33.53 | 37.52 | 36.7 | 33.26 | 30.5 | 34.7 | 36.4 | 33 | 34.4 | 32.61 | 32.3 | [5 empty] | 25
| 1 | 3 | 34.3 | 36.5 | 30.58 | 29 | 34.6 | 28 | 34.457 | 32.24 | 31.79 | 33.23 | 36.29 | 38.1 | 32.79 | 29.6 | 34.3 | 36.8 | 34 | 34.6 | 35.31 | 34 | [5 empty] | 16
| 1 | 4 | 35.3 | 37.1 | 31.57 | 32.5 | 32.6 | 29 | 34.298 | 32.71 | 31.05 | 33.14 | 37.5 | 34.8 | 32.39 | 30 | 33.7 | 36.3 | 34 | 35.2 | 33.05 | 31.5 | [5 empty] | 33.8419
| 1 | 5 | 34.5 | 36.9 | 31.1 | 32 | 33.4 | 29 | 34.951 | 32.37 | 31.46 | 32.82 | 36.83 | 36 | 33.57 | 31.6 | 34.6 | 37 | 33 | 33.6 | 33.37 | 34.9 | [5 empty] | 79
| 1 | 6 | 33.2 | 36.7 | 31.29 | 32 | 33.2 | 29 | 35.049 | 32.12 | 30.76 | 32.98 | 36.11 | 38 | 33.63 | 30.6 | 34.6 | 36.1 | 35 | 35 | 34.4 | 35.3
| Mean |  | 34.4833 | 36.7167 | 30.9233 | 32.3333 | 33.1 | 28.8333 | 34.6272 | 32.575 | 31.285 | 33.1083 | 36.76 | 36.75 | 33.0433 | 30.6667 | 34.3667 | 36.5667 | 34.1667 | 34.8833 | 33.9417 | 33.9
| Median |  | 34.55 | 36.8 | 30.84 | 32 | 33 | 29 | 34.504 | 32.305 | 31.255 | 33.06 | 36.57 | 36.8 | 33.025 | 30.55 | 34.45 | 36.6 | 34 | 34.8 | 33.885 | 34.45
| Std Dev. |  | 0.72503 | 0.32506 | 0.46324 | 2.6013 | 0.88318 | 0.40825 | 0.30016 | 0.69099 | 0.40766 | 0.2523 | 0.62865 | 1.247 | 0.51744 | 0.84301 | 0.36697 | 0.35024 | 1.169 | 0.96833 | 1.0877 | 1.6456
| Rel.Std.Dev. |  | 2.103% | 0.8853% | 1.498% | 8.045% | 2.668% | 1.416% | 0.8668% | 2.121% | 1.303% | 0.7621% | 1.71% | 3.393% | 1.566% | 2.749% | 1.068% | 0.9578% | 3.422% | 2.776% | 3.205% | 4.854%
| PDM3 |  | 1.895% | 8.495% | -8.624% | -4.458% | -2.192% | -14.8% | 2.32% | -3.744% | -7.556% | -2.168% | 8.623% | 8.593% | -2.36% | -9.383% | 1.551% | 8.051% | 0.9595% | 3.077% | 0.2947% | 0.1716%
| Z-Score (Absolute) |  | 0.33 | 1.63 | 1.75 | 0.93 | 0.48 | 2.97 | 0.41 | 0.79 | 1.54 | 0.47 | 1.66 | 1.65 | 0.51 | 1.9 | 0.26 | 1.55 | 0.14 | 0.56 | 0.01 | 0.01
| Analytical results for Co in OREAS 282 (Certified Value 3.61 ppm) | [26 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 3.7 | 5 | 3 | 3.8 | 3.9 | 3 | 2.8 | 4.8 | 3.36 | 3.4 | 3.4 | 2.84 | 3.6 | 4.4 | 3.8 | 4.5 | 3.8 | 3.7 | 3.9 | 3.3 | 4 | 4.2 | 3.3 | 3.6 |  | 1
| 1 | 2 | 3.8 | 4 | 3 | 3.8 | 3.7 | 3 | 3 | 5 | 3.3 | 3.4 | 3.3 | 2.98 | 3.7 | 4.3 | 3.9 | 4.2 | 3.8 | 3.7 | 3.9 | 3.4 | 4 | 3.9 | 3.1 | 3.4 |  | 26
| 1 | 3 | 3.7 | 5 | 3 | 3.8 | 3.9 | 3 | 3 | 5.1 | 3.27 | 3.4 | 3.4 | 2.99 | 3.6 | 4.3 | 3.9 | 4.4 | 3.7 | 3.6 | 3.9 | 3.3 | 4 | 4 | 3.4 | 3.5 |  | 16
| 1 | 4 | 3.8 | 5 | 3 | 3.8 | 3.8 | 3 | 3.2 | 5.1 | 3.34 | 3.6 | 3.3 | 2.79 | 3.7 | 4.2 | 3.8 | 4 | 3.7 | 3.7 | 3.6 | 3.5 | 4 | 4 | 3.3 | 3.3 |  | 3.6111
| 1 | 5 | 3.7 | 5 | 3 | 3.8 | 3.7 | 3 | 3 | 4.8 | 3.32 | 3.4 | 3.3 | 2.81 | 3.6 | 4.2 | 3.9 | 4.4 | 3.8 | 3.5 | 3.5 | 3.2 | 4 | 3.9 | 3.3 | 3.5 |  | 80
| 1 | 6 | 3.6 | 4 | 3 | 3.8 | 3.7 | 3 | 3 | 5.2 | 2.98 | 3.5 | 3.3 | 2.75 | 3.6 | 4.4 | 3.7 | 4.2 | 3.7 | 3.6 | 3.7 | 3.6 | 4 | 4.1 | 3.2 | 3.7
| Mean |  | 3.7167 | 4.6667 | 3 | 3.8 | 3.7833 | 3 | 3 | 5 | 3.2617 | 3.45 | 3.3333 | 2.86 | 3.6333 | 4.3 | 3.8333 | 4.2833 | 3.75 | 3.6333 | 3.75 | 3.3833 | 4 | 4.0167 | 3.2667 | 3.5
| Median |  | 3.7 | 5 | 3 | 3.8 | 3.75 | 3 | 3 | 5.05 | 3.31 | 3.4 | 3.3 | 2.825 | 3.6 | 4.3 | 3.85 | 4.3 | 3.75 | 3.65 | 3.8 | 3.35 | 4 | 4 | 3.3 | 3.5
| Std Dev. |  | 0.075277 | 0.5164 | 0 | 4.8648e-16 | 0.098319 | 0 | 0.12649 | 0.16733 | 0.14148 | 0.083666 | 0.05164 | 0.10119 | 0.05164 | 0.089443 | 0.08165 | 0.18348 | 0.054772 | 0.08165 | 0.17607 | 0.1472 | 0 | 0.1169 | 0.10328 | 0.14142
| Rel.Std.Dev. |  | 2.025% | 11.07% | 0% | 1.28e-14% | 2.599% | 0% | 4.216% | 3.347% | 4.338% | 2.425% | 1.549% | 3.538% | 1.421% | 2.08% | 2.13% | 4.284% | 1.461% | 2.247% | 4.695% | 4.351% | 0% | 2.91% | 3.162% | 4.041%
| PDM3 |  | 2.923% | 29.23% | -16.92% | 5.23% | 4.769% | -16.92% | -16.92% | 38.46% | -9.677% | -4.462% | -7.693% | -20.8% | 0.6151% | 19.08% | 6.154% | 18.62% | 3.846% | 0.6151% | 3.846% | -6.308% | 10.77% | 11.23% | -9.539% | -3.077%
| Z-Score (Absolute) |  | 0.11 | NA | NA | 0.33 | 0.28 | NA | 1.76 | 3.46 | 1.08 | 0.59 | 0.89 | 2.13 | 0.11 | 1.63 | 0.41 | 1.59 | 0.2 | 0.11 | 0.2 | 0.76 | NA | 0.89 | 1.07 | 0.46
Results from laboratories 2, 3, 7 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 282 (Certified Value 23.8 ppm) | [26 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 22 | 25 | 21 | 24 | 25 | 20 | 27 | 24.5 | 24 | 24.65 | 25 | 33 | 24 | 46 | 21 | 26 | 25.8 | 21.1 | 25 | 25 | 24 | 22 |  |  |  | 1
| 1 | 2 | 22 | 25 | 22 | 23 | 25 | 20 | 27 | 24.5 | 25 | 25.06 | 25 | 33 | 23 | 25 | 21 | 25 | 26 | 21.5 | 24 | 24 | 23 | 22 |  |  |  | 27
| 1 | 3 | 22 | 25 | 21 | 23 | 25 | 20 | 27 | 24.1 | 25 | 25.01 | 25 | 34 | 23 | 35 | 21 | 25 | 25.6 | 21.2 | 24 | 23 | 25 | 22 |  |  |  | 16
| 1 | 4 | 22 | 30 | 21 | 23 | 25 | 20 | 28 | 24.2 | 23 | 24.78 | 25 | 31 | 23 | 22 | 21 | 25 | 24.7 | 21.1 | 24 | 24 | 24 | 22 |  |  |  | 23.7639
| 1 | 5 | 23 | 30 | 21 | 23 | 25 | 20 | 28 | 24.8 | 23 | 24.15 | 24 | 32 | 23 | 25 | 21 | 26 | 24.4 | 20.8 | 24 | 23 | 24 | 22 |  |  |  | 81
| 1 | 6 | 22 | 30 | 22 | 23 | 25 | 20 | 27 | 24.4 | 22 | 24.84 | 25 | 33 | 23 | 39 | 20 | 25 | 25.8 | 21.1 | 24 | 24 | 24 | 22
| Mean |  | 22.1667 | 27.5 | 21.3333 | 23.1667 | 25 | 20 | 27.3333 | 24.4167 | 23.6667 | 24.7483 | 24.8333 | 32.6667 | 23.1667 | 32 | 20.8333 | 25.3333 | 25.3833 | 21.1333 | 24.1667 | 23.8333 | 24 | 22
| Median |  | 22 | 27.5 | 21 | 23 | 25 | 20 | 27 | 24.45 | 23.5 | 24.81 | 25 | 33 | 23 | 30 | 21 | 25 | 25.7 | 21.1 | 24 | 24 | 24 | 22
| Std Dev. |  | 0.40825 | 2.7386 | 0.5164 | 0.40825 | 0 | 0 | 0.5164 | 0.24833 | 1.2111 | 0.32945 | 0.40825 | 1.0328 | 0.40825 | 9.5079 | 0.40825 | 0.5164 | 0.66458 | 0.22509 | 0.40825 | 0.75277 | 0.63246 | 0
| Rel.Std.Dev. |  | 1.842% | 9.959% | 2.421% | 1.762% | 0% | 0% | 1.889% | 1.017% | 5.117% | 1.331% | 1.644% | 3.162% | 1.762% | 29.71% | 1.96% | 2.038% | 2.618% | 1.065% | 1.689% | 3.158% | 2.635% | 0%
| PDM3 |  | -6.721% | 15.72% | -10.23% | -2.513% | 5.201% | -15.84% | 15.02% | 2.747% | -0.4094% | 4.142% | 4.5% | 37.46% | -2.513% | 34.66% | -12.33% | 6.604% | 6.814% | -11.07% | 1.695% | 0.292% | 0.9933% | -7.423%
| Z-Score (Absolute) |  | 1.16 | 1.93 | 1.64 | 0.58 | 0.48 | NA | 1.83 | 0.14 | 0.29 | 0.34 | 0.39 | 4.91 | 0.58 | 4.53 | 1.93 | 0.67 | 0.7 | 1.75 | 0 | 0.19 | 0.1 | 1.25
Results from laboratory 8 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 282 (Certified Value 4.23 ppm) | [26 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.29 | 2.30 | 2.31 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 04g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 3.96 | 3.78 | 3.93 | 5.28 | 3 | 3.2 | 3.4 | 4.09 | 4.21 | 5.14 | 5.05 | 5.61 | 4.96 | 4.17 | 3.74 | 3.66 | 5.48 | 4.26 | 4.3 | 3.99 | [5 empty] | 1
| 1 | 2 | 4.02 | 3.64 | 4.06 | 5.12 | 3 | 3.2 | 3.6 | 4.1 | 4.06 | 4.96 | 5.12 | 5.67 | 4.73 | 4.27 | 3.6 | 3.47 | 5.3 | 4.02 | 4.43 | 3.64 | [5 empty] | 28
| 1 | 3 | 3.95 | 3.66 | 3.9 | 5.15 | 2.5 | 3.4 | 3.4 | 3.99 | 3.95 | 5.17 | 5.02 | 5.58 | 5.05 | 4.21 | 3.56 | 3.56 | 5.66 | 4.05 | 4.46 | 3.82 | [5 empty] | 16
| 1 | 4 | 4.12 | 3.78 | 4.18 | 5.17 | 3 | 3.4 | 3.4 | 4.04 | 4.14 | 5.03 | 5.08 | 5.57 | 4.7 | 4.22 | 3.53 | 3.47 | 5.46 | 4.07 | 4.36 | 3.59 | [5 empty] | 4.229
| 1 | 5 | 4 | 3.7 | 4.47 | 5.07 | 3 | 3.2 | 3.5 | 4.02 | 4.37 | 4.92 | 4.99 | 5.38 | 4.73 | 4.23 | 3.8 | 3.62 | 5.6 | 3.74 | 4.37 | 3.84 | [5 empty] | 82
| 1 | 6 | 4.02 | 3.58 | 3.92 | 5.29 | 3 | 3.2 | 3.6 | 3.78 | 4.4 | 4.88 | 5.07 | 5.56 | 4.8 | 4.22 | 3.7 | 3.5 | 5.64 | 3.97 | 4.25 | 4
| Mean |  | 4.0117 | 3.69 | 4.0767 | 5.18 | 2.9167 | 3.2667 | 3.4833 | 4.0033 | 4.1883 | 5.0167 | 5.055 | 5.5617 | 4.8283 | 4.22 | 3.655 | 3.5467 | 5.5233 | 4.0183 | 4.3617 | 3.8133
| Median |  | 4.01 | 3.68 | 3.995 | 5.16 | 3 | 3.2 | 3.45 | 4.03 | 4.175 | 4.995 | 5.06 | 5.575 | 4.765 | 4.22 | 3.65 | 3.53 | 5.54 | 4.035 | 4.365 | 3.83
| Std Dev. |  | 0.0608 | 0.07975 | 0.22042 | 0.088091 | 0.20412 | 0.10328 | 0.098319 | 0.11708 | 0.17543 | 0.11843 | 0.045935 | 0.097451 | 0.14359 | 0.032249 | 0.10766 | 0.080416 | 0.13706 | 0.16845 | 0.078337 | 0.17131
| Rel.Std.Dev. |  | 1.516% | 2.161% | 5.407% | 1.701% | 6.999% | 3.162% | 2.823% | 2.924% | 4.189% | 2.361% | 0.9087% | 1.752% | 2.974% | 0.7642% | 2.945% | 2.267% | 2.482% | 4.192% | 1.796% | 4.492%
| PDM3 |  | -5.14% | -12.75% | -3.603% | 22.49% | -31.03% | -22.76% | -17.63% | -5.337% | -0.9628% | 18.62% | 19.53% | 31.51% | 14.17% | -0.214% | -13.57% | -16.14% | 30.6% | -4.983% | 3.136% | -9.83%
| Z-Score (Absolute) |  | 0.05 | 0.54 | 0.04 | 1.71 | 1.71 | 1.18 | 0.85 | 0.07 | 0.21 | 1.46 | 1.52 | 2.29 | 1.18 | 0.26 | 0.59 | 0.76 | 2.23 | 0.04 | 0.47 | 0.35
| Analytical results for Cu in OREAS 282 (Certified Value 18.6 ppm) | [26 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 18.8 | 19 | 17 | 19.2 | 13 | 24 | 12.6 | 17.55 | 17 | 15 | 18.64 | 17 | 20.6 | 19.4 | 21 | 17.9 | 18 | 18.98 | 22.2 | 17.42 | 19 | 20.1 | 18.3 | 20.3 |  | 1
| 1 | 2 | 18.6 | 19 | 18 | 19.4 | 13 | 24 | 12.3 | 17.45 | 16.8 | 14 | 18.47 | 18 | 20.2 | 18.9 | 20 | 18.2 | 19 | 19.49 | 21.4 | 17.62 | 20 | 19 | 18.1 | 18.7 |  | 29
| 1 | 3 | 18.2 | 19.5 | 17 | 19 | 13 | 22 | 12 | 17.17 | 17 | 15 | 19.07 | 17 | 21.1 | 19.3 | 21 | 18.1 | 19 | 18.88 | 21.6 | 17.62 | 18 | 19.4 | 17.8 | 19.5 |  | 16
| 1 | 4 | 18.8 | 19.5 | 17 | 19.1 | 13 | 23 | 12.5 | 17.47 | 17.3 | 16 | 18.59 | 18 | 19.8 | 18.7 | 18 | 17.9 | 18 | 17.7 | 21.9 | 17.23 | 19 | 19.2 | 18.1 | 18.5 |  | 18.5941
| 1 | 5 | 18.6 | 19.5 | 17 | 19.2 | 13 | 22 | 12.1 | 17.79 | 17 | 15 | 18.11 | 17 | 19.8 | 19.1 | 23 | 18.3 | 18 | 17.7 | 22.4 | 17.23 | 19 | 18.6 | 17.7 | 19.7 |  | 83
| 1 | 6 | 18 | 20.5 | 18 | 19.1 | 13 | 24 | 11.9 | 17.37 | 16.9 | 16 | 18.43 | 18 | 20.6 | 19.1 | 20 | 18.3 | 18 | 18.28 | 21.9 | 16.6 | 20 | 19.5 | 17.9 | 20.5
| Mean |  | 18.5 | 19.5 | 17.3333 | 19.1667 | 13 | 23.1667 | 12.2333 | 17.4667 | 17 | 15.1667 | 18.5517 | 17.5 | 20.35 | 19.0833 | 20.5 | 18.1167 | 18.3333 | 18.505 | 21.9 | 17.2867 | 19.1667 | 19.3 | 17.9833 | 19.5333
| Median |  | 18.6 | 19.5 | 17 | 19.15 | 13 | 23.5 | 12.2 | 17.46 | 17 | 15 | 18.53 | 17.5 | 20.4 | 19.1 | 20.5 | 18.15 | 18 | 18.58 | 21.9 | 17.325 | 19 | 19.3 | 18 | 19.6
| Std Dev. |  | 0.32863 | 0.54772 | 0.5164 | 0.13663 | 0 | 0.98319 | 0.28048 | 0.20452 | 0.16733 | 0.75277 | 0.31448 | 0.54772 | 0.51284 | 0.25626 | 1.6432 | 0.18348 | 0.5164 | 0.73255 | 0.36878 | 0.37893 | 0.75277 | 0.50596 | 0.22286 | 0.81404
| Rel.Std.Dev. |  | 1.776% | 2.809% | 2.979% | 0.7128% | 0% | 4.244% | 2.293% | 1.171% | 0.9843% | 4.963% | 1.695% | 3.13% | 2.52% | 1.343% | 8.015% | 1.013% | 2.817% | 3.959% | 1.684% | 2.192% | 3.928% | 2.622% | 1.239% | 4.167%
| PDM3 |  | -0.5062% | 4.872% | -6.781% | 3.079% | -30.09% | 24.59% | -34.21% | -6.064% | -8.573% | -18.43% | -0.2284% | -5.884% | 9.443% | 2.631% | 10.25% | -2.568% | -1.403% | -0.4793% | 17.78% | -7.032% | 3.079% | 3.796% | -3.285% | 5.051%
| Z-Score (Absolute) |  | 0 | 0.66 | 0.78 | 0.44 | 3.65 | 3.09 | 4.16 | 0.69 | 1 | 2.21 | 0.03 | 0.66 | 1.23 | 0.39 | 1.32 | 0.26 | 0.11 | 0 | 2.25 | 0.81 | 0.44 | 0.53 | 0.34 | 0.68
| Analytical results for Dy in OREAS 282 (Certified Value 2.43 ppm) | [26 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | [17 empty] | 3
| 1 | 1 | 2.56 | 2.4 | 2.16 | 2.568 | 2.58 | 2.38 | 2.8 | 2.19 | [17 empty] | 1
| 1 | 2 | 2.52 | 2.7 | 2.14 | 2.566 | 2.31 | 2.45 | 2.7 | 2.18 | [17 empty] | 30
| 1 | 3 | 2.49 | 2.4 | 2.29 | 2.534 | 2.42 | 2.44 | 3 | 2.16 | [17 empty] | 16
| 1 | 4 | 2.61 | 2.5 | 2.2 | 2.546 | 2.43 | 2.41 | 2.6 | 2.01 | [17 empty] | 2.433
| 1 | 5 | 2.56 | 2.5 | 2.23 | 2.553 | 2.47 | 2.31 | 2.5 | 2.25 | [17 empty] | 84
| 1 | 6 | 2.5 | 2.4 | 2.18 | 2.597 | 2.34 | 2.28 | 2.8 | 2.07
| Mean |  | 2.54 | 2.4833 | 2.2 | 2.5607 | 2.425 | 2.3783 | 2.7333 | 2.1433
| Median |  | 2.54 | 2.45 | 2.19 | 2.5595 | 2.425 | 2.395 | 2.75 | 2.17
| Std Dev. |  | 0.045166 | 0.1169 | 0.054037 | 0.021851 | 0.096488 | 0.06969 | 0.17512 | 0.08756
| Rel.Std.Dev. |  | 1.778% | 4.708% | 2.456% | 0.8533% | 3.979% | 2.93% | 6.407% | 4.085%
| PDM3 |  | 4.398% | 2.069% | -9.577% | 5.247% | -0.3288% | -2.247% | 12.34% | -11.91%
| Z-Score (Absolute) |  | 0.6 | 0.2 | 1.78 | 0.75 | 0.2 | 0.53 | 1.96 | 2.18
| Analytical results for Er in OREAS 282 (Certified Value 1.26 ppm) | [26 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | [17 empty] | 3
| 1 | 1 | 1.36 | 1.2 | 1.18 | 1.314 | 1.33 | 1.22 | 1.4 | 1.19 | [17 empty] | 1
| 1 | 2 | 1.36 | 1.4 | 1.17 | 1.301 | 1.23 | 1.36 | 1.4 | 1.17 | [17 empty] | 31
| 1 | 3 | 1.36 | 1.2 | 1.26 | 1.299 | 1.26 | 1.34 | 1.4 | 1.19 | [17 empty] | 16
| 1 | 4 | 1.4 | 1.25 | 1.19 | 1.291 | 1.27 | 1.26 | 1.3 | 1.12 | [17 empty] | 1.2584
| 1 | 5 | 1.35 | 1.25 | 1.21 | 1.3 | 1.32 | 1.21 | 1.3 | 1.23 | [17 empty] | 85
| 1 | 6 | 1.35 | 1.25 | 1.17 | 1.312 | 1.27 | 1.27 | 1.4 | 1.13
| Mean |  | 1.3633 | 1.2583 | 1.1967 | 1.3028 | 1.28 | 1.2767 | 1.3667 | 1.1717
| Median |  | 1.36 | 1.25 | 1.185 | 1.3005 | 1.27 | 1.265 | 1.4 | 1.18
| Std Dev. |  | 0.018619 | 0.073598 | 0.034448 | 0.0086583 | 0.037947 | 0.061536 | 0.05164 | 0.041191
| Rel.Std.Dev. |  | 1.366% | 5.849% | 2.879% | 0.6646% | 2.965% | 4.82% | 3.779% | 3.516%
| PDM3 |  | 8.342% | -0.001892% | -4.902% | 3.534% | 1.72% | 1.455% | 8.607% | -6.889%
| Z-Score (Absolute) |  | 2.23 | 0.47 | 2.06 | 0.67 | 0.09 | 0 | NA | 2.71
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 282 (Certified Value 0.6 ppm) | [26 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | [17 empty] | 3
| 1 | 1 | 0.66 | 0.6 | 0.568 | 0.653 | 0.66 | 0.513 | 0.7 | 0.56 | [17 empty] | 1
| 1 | 2 | 0.65 | 0.6 | 0.548 | 0.661 | 0.63 | 0.528 | 0.7 | 0.57 | [17 empty] | 32
| 1 | 3 | 0.644 | 0.6 | 0.604 | 0.654 | 0.62 | 0.524 | 0.7 | 0.54 | [17 empty] | 16
| 1 | 4 | 0.656 | 0.6 | 0.584 | 0.653 | 0.62 | 0.517 | 0.7 | 0.53 | [17 empty] | 0.5985
| 1 | 5 | 0.644 | 0.6 | 0.6 | 0.653 | 0.62 | 0.512 | 0.7 | 0.59 | [17 empty] | 86
| 1 | 6 | 0.65 | 0.6 | 0.572 | 0.66 | 0.64 | 0.518 | 0.7 | 0.54
| Mean |  | 0.65067 | 0.6 | 0.57933 | 0.65567 | 0.63167 | 0.51867 | 0.7 | 0.555
| Median |  | 0.65 | 0.6 | 0.578 | 0.6535 | 0.625 | 0.5175 | 0.7 | 0.55
| Std Dev. |  | 0.0064083 | 0 | 0.021078 | 0.0037771 | 0.016021 | 0.0062503 | 1.2162e-16 | 0.022583
| Rel.Std.Dev. |  | 0.9849% | 0% | 3.638% | 0.5761% | 2.536% | 1.205% | 1.737e-14% | 4.069%
| PDM3 |  | 8.716% | 0.2506% | -3.202% | 9.552% | 5.542% | -13.34% | 16.96% | -7.268%
| Z-Score (Absolute) |  | 0.64 | NA | 0.37 | 0.71 | 0.37 | 1.23 | NA | 0.71
Results from laboratories 7 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 282 (Certified Value 0.998 wt.%) | [26 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1 | 1.09 | 0.97 | 1.05 | 1.02 | 0.9235 | 0.98 | 0.82 | 1.004 | 1.01 | 0.9654 | 1.02 | 1.26 | 0.97 | 1.15 | 0.97 | 0.79 | 0.98 | 0.94 | 1.04 | 1.12 | 1.04 | 0.97 |  |  | 1
| 1 | 2 | 0.98 | 1.02 | 0.97 | 1.04 | 1.02 | 0.935 | 1.03 | 0.82 | 1.006 | 0.98 | 0.9683 | 1.05 | 1.25 | 0.97 | 1.19 | 0.97 | 0.79 | 0.96 | 0.96 | 1.04 | 1.07 | 1.02 | 0.96 |  |  | #N/A
| 1 | 3 | 0.97 | 1.06 | 0.96 | 1.04 | 1.03 | 0.9315 | 1.02 | 0.8 | 0.98 | 1.01 | 0.9827 | 1.02 | 1.3 | 0.96 | 1.2 | 0.97 | 0.79 | 0.96 | 0.96 | 1.05 | 1.06 | 1.02 | 0.98 |  |  | 16
| 1 | 4 | 0.99 | 1.13 | 0.97 | 1.04 | 1.02 | 0.928 | 1.02 | 0.84 | 1 | 0.97 | 0.9621 | 1.05 | 1.19 | 0.94 | 1.16 | 0.97 | 0.81 | 0.95 | 0.93 | 1.05 | 1.08 | 1.01 | 0.98 |  |  | 0.99785
| 1 | 5 | 1 | 1.08 | 0.97 | 1.04 | 1.01 | 0.9215 | 1.03 | 0.83 | 1.011 | 0.98 | 0.9567 | 1.03 | 1.25 | 0.96 | 1.16 | 0.98 | 0.81 | 0.96 | 0.93 | 1.02 | 1.04 | 1.02 | 0.97 |  |  | 87
| 1 | 6 | 0.98 | 1.05 | 0.98 | 1.05 | 1.04 | 0.9325 | 0.97 | 0.81 | 0.992 | 0.96 | 0.9642 | 1.04 | 1.26 | 0.96 | 1.19 | 0.96 | 0.79 | 0.97 | 0.99 | 1.06 | 1.07 | 1.01 | 0.97
| Mean |  | 0.98667 | 1.0717 | 0.97 | 1.0433 | 1.0233 | 0.92867 | 1.0083 | 0.82 | 0.99883 | 0.985 | 0.96657 | 1.035 | 1.2517 | 0.96 | 1.175 | 0.97 | 0.79667 | 0.96333 | 0.95167 | 1.0433 | 1.0733 | 1.02 | 0.97167
| Median |  | 0.985 | 1.07 | 0.97 | 1.04 | 1.02 | 0.92975 | 1.02 | 0.82 | 1.002 | 0.98 | 0.9648 | 1.035 | 1.255 | 0.96 | 1.175 | 0.97 | 0.79 | 0.96 | 0.95 | 1.045 | 1.07 | 1.02 | 0.97
| Std Dev. |  | 0.012111 | 0.037639 | 0.0063246 | 0.005164 | 0.010328 | 0.0053166 | 0.026394 | 0.014142 | 0.011215 | 0.020736 | 0.008804 | 0.013784 | 0.035449 | 0.010954 | 0.020736 | 0.0063246 | 0.010328 | 0.010328 | 0.023166 | 0.013663 | 0.026583 | 0.010954 | 0.0075277
| Rel.Std.Dev. |  | 1.227% | 3.512% | 0.652% | 0.4949% | 1.009% | 0.5725% | 2.618% | 1.725% | 1.123% | 2.105% | 0.9109% | 1.332% | 2.832% | 1.141% | 1.765% | 0.652% | 1.296% | 1.072% | 2.434% | 1.31% | 2.477% | 1.074% | 0.7747%
| PDM3 |  | -1.12% | 7.398% | -2.791% | 4.559% | 2.554% | -6.933% | 1.051% | -17.82% | 0.09899% | -1.287% | -3.135% | 3.723% | 25.44% | -3.793% | 17.75% | -2.791% | -20.16% | -3.459% | -4.628% | 4.559% | 7.565% | 2.22% | -2.624%
| Z-Score (Absolute) |  | 0 | 1.64 | 0.32 | 1.09 | 0.71 | 1.12 | 0.42 | 3.21 | 0.23 | 0.03 | 0.39 | 0.93 | 5.11 | 0.51 | 3.63 | 0.32 | 3.66 | 0.45 | 0.67 | 1.09 | 1.67 | 0.64 | 0.29
| Analytical results for Ga in OREAS 282 (Certified Value 2.99 ppm) | [26 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 2.89 | 3 | 2 | 2.6 | 3.5 | 2 | 4 | 2.9 | 3 | 3.8 | 3.42 | 3.74 | 3.95 | 3 | 2.9 | 2.83 | 4.5 | 2.8 | 2.3 | < 2 | 3.24 | 2.69 | 2.98 |  |  | 1
| 1 | 2 | 2.93 | 3 | 2 | 2.5 | 3.4 | 3 | 4 | 2.9 | 2.63 | 3.6 | 3.53 | 3.69 | 3.91 | 2.9 | 3.1 | 2.8 | 4.7 | 2.8 | 2.4 | < 2 | 2.91 | 2.48 | 2.74 |  |  | 33
| 1 | 3 | 2.87 | 3 | 2 | 2.5 | 3.6 | 2 | 4 | 2.7 | 2.79 | 3.8 | 3.49 | 3.7 | 4.11 | 3 | 3 | 2.77 | 4.5 | 2.9 | 2.2 | < 2 | 3.07 | 2.78 | 2.86 |  |  | 16
| 1 | 4 | 3 | 3 | 2 | 2.6 | 3.5 | 2 | 4 | 2.8 | 3.09 | 3.3 | 3.44 | 3.62 | 3.8 | 3 | 2.7 | 2.77 | 4.6 | 2.7 | 2.4 | < 2 | 2.97 | 2.69 | 2.73 |  |  | 2.9864
| 1 | 5 | 2.94 | 3 | 2 | 2.5 | 3.6 | 2 | 4 | 2.8 | 3.02 | 3.5 | 3.41 | 3.61 | 3.95 | 3 | 3.2 | 2.8 | 4.4 | 2.8 | 2.3 | < 2 | 2.88 | 2.59 | 2.93 |  |  | 88
| 1 | 6 | 2.85 | 2.8 | 2 | 2.6 | 3.7 | 2 | 4 | 2.7 | 3.08 | 3.4 | 3.39 | 3.53 | 4 | 2.9 | 2.9 | 2.74 | 4.6 | 2.6 | 2.6 | < 2 | 3.02 | 2.63 | 3.05
| Mean |  | 2.9133 | 2.9667 | 2 | 2.55 | 3.55 | 2.1667 | 4 | 2.8 | 2.935 | 3.5667 | 3.4467 | 3.6483 | 3.9533 | 2.9667 | 2.9667 | 2.785 | 4.55 | 2.7667 | 2.3667 |  | 3.015 | 2.6433 | 2.8817
| Median |  | 2.91 | 3 | 2 | 2.55 | 3.55 | 2 | 4 | 2.8 | 3.01 | 3.55 | 3.43 | 3.655 | 3.95 | 3 | 2.95 | 2.785 | 4.55 | 2.8 | 2.35 |  | 2.995 | 2.66 | 2.895
| Std Dev. |  | 0.05465 | 0.08165 | 0 | 0.054772 | 0.10488 | 0.40825 | 0 | 0.089443 | 0.18469 | 0.20656 | 0.053166 | 0.076267 | 0.10211 | 0.05164 | 0.17512 | 0.031464 | 0.10488 | 0.10328 | 0.13663 |  | 0.13035 | 0.1027 | 0.12952
| Rel.Std.Dev. |  | 1.876% | 2.752% | 0% | 2.148% | 2.954% | 18.84% | 0% | 3.194% | 6.293% | 5.791% | 1.543% | 2.09% | 2.583% | 1.741% | 5.903% | 1.13% | 2.305% | 3.733% | 5.773% |  | 4.323% | 3.885% | 4.495%
| PDM3 |  | -2.446% | -0.6599% | -33.03% | -14.61% | 18.87% | -27.45% | 33.94% | -6.241% | -1.72% | 19.43% | 15.41% | 22.17% | 32.38% | -0.6599% | -0.6599% | -6.743% | 52.36% | -7.357% | -20.75% |  | 0.9586% | -11.49% | -3.506%
| Z-Score (Absolute) |  | 0.1 | 0.04 | NA | 1.1 | 1.64 | NA | NA | 0.41 | 0.04 | 1.69 | 1.36 | 1.91 | 2.75 | 0.04 | 0.04 | 0.45 | 4.39 | 0.51 | 1.6 | 5.35 | 0.18 | 0.84 | 0.19
Results from laboratories 3, 7 and 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 282 (Certified Value 2.92 ppm) | [26 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | [17 empty] | 3
| 1 | 1 | 3.18 | 2.7 | 2.6 | 3.027 | 3.03 | 2.81 | 3.3 | 2.74 | [17 empty] | 1
| 1 | 2 | 3.15 | 3.1 | 2.61 | 3.03 | 2.88 | 2.98 | 3.2 | 2.74 | [17 empty] | 34
| 1 | 3 | 3.16 | 2.7 | 2.77 | 3.012 | 2.9 | 2.97 | 3.4 | 2.71 | [17 empty] | 16
| 1 | 4 | 3.2 | 2.8 | 2.63 | 3.016 | 2.92 | 2.84 | 3 | 2.6 | [17 empty] | 2.9224
| 1 | 5 | 3.22 | 2.8 | 2.68 | 3.036 | 2.92 | 2.88 | 3.3 | 2.82 | [17 empty] | 89
| 1 | 6 | 3.16 | 2.8 | 2.63 | 2.752 | 2.85 | 2.91 | 3.5 | 2.65
| Mean |  | 3.1783 | 2.8167 | 2.6533 | 2.9788 | 2.9167 | 2.8983 | 3.2833 | 2.71
| Median |  | 3.17 | 2.8 | 2.63 | 3.0215 | 2.91 | 2.895 | 3.3 | 2.725
| Std Dev. |  | 0.027142 | 0.1472 | 0.063456 | 0.11148 | 0.061536 | 0.068532 | 0.17224 | 0.076942
| Rel.Std.Dev. |  | 0.854% | 5.226% | 2.392% | 3.742% | 2.11% | 2.365% | 5.246% | 2.839%
| PDM3 |  | 8.759% | -3.617% | -9.206% | 1.933% | -0.1946% | -0.822% | 12.35% | -7.267%
| Z-Score (Absolute) |  | 1.27 | 0.42 | 1.19 | 0.33 | 0.04 | 0.04 | 1.76 | 0.92
| Analytical results for Ge in OREAS 282 (Certified Value < 0.1 ppm) | [26 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.09 | 2.10 | 2.16 | 2.18 | 2.20 | 2.29 | 2.30 | 2.31 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [14 empty] | 2
|  |  | 0.5g | 0.25g | 25g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | [14 empty] | 2
| 1 | 1 | 0.06 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | 0.06 | < 0.05 | < 0.05 | < 0.05 | [14 empty] | 1
| 1 | 2 | 0.06 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | 0.05 | < 0.05 | < 0.05 | < 0.05 | [14 empty] | 35
| 1 | 3 | 0.06 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | 0.06 | < 0.05 | < 0.05 | 0.05 | [14 empty] | 16
| 1 | 4 | 0.07 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.1 | 0.06 | < 0.05 | 0.05 | 0.05 | [14 empty] | < 0.1
| 1 | 5 | 0.06 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | 0.07 | < 0.05 | < 0.05 | < 0.05 | [14 empty] | 90
| 1 | 6 | 0.07 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | < 0.1 | 0.05 | < 0.05 | < 0.05 | < 0.05
| Mean |  | 0.063333 | [5 empty] | 0.1 | 0.058333 |  | 0.05 | 0.05
| Median |  | 0.06 | [5 empty] | 0.1 | 0.06 |  | 0.05 | 0.05
| Std Dev. |  | 0.005164 | [6 empty] | 0.0075277 |  |  | 0
| Rel.Std.Dev. |  | 8.154% | [6 empty] | 12.9% |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 1.08 | 0 | 0 | 0 | 2.02 | 0 | 0.67 | 0.67 | 2.02 | 1.69 | 1.35
| Analytical results for Hf in OREAS 282 (Certified Value 0.35 ppm) | [26 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.29 | 2.30 | 2.31 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 04g | 0.25g | 0.25g | 25g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.35 | 0.2 | 0.33 | 0.48 | < 0.5 | 0.3 | 0.34 | 0.38 | 0.33 | 0.218 | 0.35 | 0.5 | 0.36 | 0.39 | 0.38 | 0.41 | 0.38 | 0.35 | 0.33 | [6 empty] | 1
| 1 | 2 | 0.33 | 0.2 | 0.31 | 0.45 | < 0.5 | 0.3 | 0.36 | 0.39 | 0.3 | 0.239 | 0.37 | 0.5 | 0.35 | 0.36 | 0.38 | 0.39 | 0.39 | 0.34 | 0.31 | [6 empty] | 19
| 1 | 3 | 0.36 | 0.2 | 0.33 | 0.45 | < 0.5 | 0.3 | 0.34 | 0.37 | 0.32 | 0.247 | 0.36 | 0.5 | 0.35 | 0.37 | 0.38 | 0.4 | 0.39 | 0.37 | 0.32 | [6 empty] | 16
| 1 | 4 | 0.35 | 0.2 | 0.31 | 0.43 | < 0.5 | 0.3 | 0.35 | 0.4 | 0.31 | 0.221 | 0.36 | 0.5 | 0.36 | 0.35 | 0.38 | 0.41 | 0.36 | 0.36 | 0.29 | [6 empty] | 0.35474
| 1 | 5 | 0.35 | 0.2 | 0.31 | 0.43 | < 0.5 | 0.3 | 0.34 | 0.37 | 0.32 | 0.212 | 0.35 | 0.5 | 0.36 | 0.38 | 0.38 | 0.4 | 0.35 | 0.34 | 0.31 | [6 empty] | 91
| 1 | 6 | 0.34 | 0.2 | 0.34 | 0.46 | < 0.5 | 0.3 | 0.33 | 0.36 | 0.3 | 0.227 | 0.37 | 0.5 | 0.35 | 0.37 | 0.36 | 0.44 | 0.36 | 0.35 | 0.33
| Mean |  | 0.34667 | 0.2 | 0.32167 | 0.45 |  | 0.3 | 0.34333 | 0.37833 | 0.31333 | 0.22733 | 0.36 | 0.5 | 0.355 | 0.37 | 0.37667 | 0.40833 | 0.37167 | 0.35167 | 0.315
| Median |  | 0.35 | 0.2 | 0.32 | 0.45 |  | 0.3 | 0.34 | 0.375 | 0.315 | 0.224 | 0.36 | 0.5 | 0.355 | 0.37 | 0.38 | 0.405 | 0.37 | 0.35 | 0.315
| Std Dev. |  | 0.010328 | 3.0405e-17 | 0.013292 | 0.018974 |  | 0 | 0.010328 | 0.01472 | 0.012111 | 0.013307 | 0.0089443 | 0 | 0.0054772 | 0.014142 | 0.008165 | 0.017224 | 0.017224 | 0.01169 | 0.015166
| Rel.Std.Dev. |  | 2.979% | 1.52e-14% | 4.132% | 4.216% |  | 0% | 3.008% | 3.891% | 3.865% | 5.853% | 2.485% | 0% | 1.543% | 3.822% | 2.168% | 4.218% | 4.634% | 3.324% | 4.815%
| PDM3 |  | -2.277% | -43.62% | -9.324% | 26.85% |  | -15.43% | -3.216% | 6.65% | -11.67% | -35.92% | 1.482% | 40.95% | 0.07228% | 4.301% | 6.18% | 15.11% | 4.771% | -0.8674% | -11.2%
| Z-Score (Absolute) |  | 0.19 | NA | 0.88 | 2.7 | 2.88 | NA | 0.28 | 0.7 | 1.12 | 3.52 | 0.19 | NA | 0.05 | 0.47 | 0.65 | 1.53 | 0.51 | 0.05 | 1.07
Results from laboratories 2, 9 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 282 (Certified Value 40.3 ppm) | [26 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.09 | 2.10 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 0.25g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 40.5 | 40.2 | 34 | 37 | 38.78 | 39.339 | 43.49 | 42.14 | 42 | > 10 | 40.64 | 43.6 | 40.4 | 41.247 | 37.992 | 43 | 42.4 | 40.224 | 39.1 | [6 empty] | 1
| 1 | 2 | 40 | 37.8 | 34 | 38 | 38.87 | 38.661 | 43.31 | 41.18 | 42.72 | > 10 | 40.87 | 40.7 | 40.9 | 40.335 | 38.06 | 44 | 39.9 | 41.049 | 36.2 | [6 empty] | 3
| 1 | 3 | 40.7 | 37.8 | 33 | 38 | 37.54 | 39.115 | 43.52 | 42.01 | 41.7 | > 10 | 40.97 | 41.6 | 41.1 | 40.349 | 38.462 | 44 | 40.2 | 41.378 | 36.6 | [6 empty] | 16
| 1 | 4 | 39.7 | 40.1 | 33 | 37 | 38.42 | 38.477 | 43.59 | 41.04 | 43.33 | > 10 | 40.06 | 40.7 | 40.9 | 40.732 | 37.434 | 44 | 39.7 | 40.683 | 36.9 | [6 empty] | 40.3023
| 1 | 5 | 40.5 | 39 | 33 | 38 | 38.55 | 37.952 | 44.06 | 41.98 | 42.56 | > 10 | 41.02 | 43.2 | 40.7 | 40.677 | 37.916 | 43 | 38.2 | 40.246 | 36.9 | [6 empty] | 92
| 1 | 6 | 40 | 38.2 | 33 | 38 | 38.93 | 39.11 | 43.59 | 40.71 | 42.17 | > 10 | 41.98 | 43.2 | 40.8 | 40.341 | 40 | 44 | 39.8 | 38.893 | 39.1
| Mean |  | 40.2333 | 38.85 | 33.3333 | 37.6667 | 38.515 | 38.7757 | 43.5933 | 41.51 | 42.4133 |  | 40.9233 | 42.1667 | 40.8 | 40.6135 | 38.3107 | 43.6667 | 40.0333 | 40.4122 | 37.4667
| Median |  | 40.25 | 38.6 | 33 | 38 | 38.665 | 38.8855 | 43.555 | 41.58 | 42.365 |  | 40.92 | 42.4 | 40.85 | 40.513 | 38.026 | 44 | 39.85 | 40.4645 | 36.9
| Std Dev. |  | 0.38816 | 1.0986 | 0.5164 | 0.5164 | 0.51563 | 0.51479 | 0.25065 | 0.60623 | 0.58219 |  | 0.62555 | 1.3277 | 0.23664 | 0.3581 | 0.89055 | 0.5164 | 1.3545 | 0.86978 | 1.291
| Rel.Std.Dev. |  | 0.9648% | 2.828% | 1.549% | 1.371% | 1.339% | 1.328% | 0.575% | 1.46% | 1.373% |  | 1.529% | 3.149% | 0.58% | 0.8817% | 2.325% | 1.183% | 3.383% | 2.152% | 3.446%
| PDM3 |  | -0.1711% | -3.604% | -17.29% | -6.54% | -4.435% | -3.788% | 8.166% | 2.997% | 5.238% |  | 1.541% | 4.626% | 1.235% | 0.7722% | -4.942% | 8.348% | -0.6674% | 0.2726% | -7.036%
| Z-Score (Absolute) |  | 0.04 | 0.59 | 2.81 | 1.07 | 0.73 | 0.62 | 1.31 | 0.48 | 0.84 | NA | 0.24 | 0.74 | 0.19 | 0.12 | 0.81 | 1.34 | 0.12 | 0.04 | 1.15
| Analytical results for Ho in OREAS 282 (Certified Value 0.45 ppm) | [26 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | [17 empty] | 3
| 1 | 1 | 0.49 | 0.45 | 0.425 | 0.479 | 0.46 | 0.435 | 0.5 | 0.43 | [17 empty] | 1
| 1 | 2 | 0.475 | 0.5 | 0.42 | 0.484 | 0.43 | 0.44 | 0.5 | 0.42 | [17 empty] | 20
| 1 | 3 | 0.485 | 0.45 | 0.455 | 0.485 | 0.45 | 0.439 | 0.5 | 0.42 | [17 empty] | 16
| 1 | 4 | 0.495 | 0.45 | 0.435 | 0.476 | 0.43 | 0.421 | 0.5 | 0.41 | [17 empty] | 0.44986
| 1 | 5 | 0.47 | 0.45 | 0.445 | 0.476 | 0.46 | 0.432 | 0.5 | 0.44 | [17 empty] | 93
| 1 | 6 | 0.48 | 0.45 | 0.435 | 0.483 | 0.45 | 0.424 | 0.5 | 0.41
| Mean |  | 0.4825 | 0.45833 | 0.43583 | 0.4805 | 0.44667 | 0.43183 | 0.5 | 0.42167
| Median |  | 0.4825 | 0.45 | 0.435 | 0.481 | 0.45 | 0.4335 | 0.5 | 0.42
| Std Dev. |  | 0.0093541 | 0.020412 | 0.012813 | 0.0040373 | 0.013663 | 0.0078337 | 0 | 0.01169
| Rel.Std.Dev. |  | 1.939% | 4.454% | 2.94% | 0.8402% | 3.059% | 1.814% | 0% | 2.772%
| PDM3 |  | 7.256% | 1.884% | -3.117% | 6.812% | -0.7092% | -4.007% | 11.15% | -6.267%
| Z-Score (Absolute) |  | 1.63 | 0.53 | 0.49 | 1.54 | 0 | 0.67 | NA | 1.14
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 282 (Certified Value 0.056 ppm) | [26 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.29 | 2.30 | 2.31 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.061 | 0.05 | < 5 | < 0.02 | < 0.1 | < 0.5 | 0.04 | 0.06 | 0.051 | 0.05 | 0.052 | 0.06 | 0.07 | 0.06 | 0.06 | 0.056 | < 0.05 | 0.06 | 0.057 | 0.057 | [5 empty] | 1
| 1 | 2 | 0.058 | 0.055 | < 5 | < 0.02 | < 0.1 | < 0.5 | 0.05 | 0.06 | 0.049 | 0.05 | 0.057 | 0.06 | 0.06 | 0.06 | 0.05 | 0.055 | < 0.05 | 0.062 | 0.056 | 0.054 | [5 empty] | 21
| 1 | 3 | 0.057 | 0.055 | < 5 | < 0.02 | < 0.1 | < 0.5 | 0.05 | 0.05 | 0.052 | 0.05 | 0.054 | 0.06 | 0.07 | 0.06 | 0.05 | 0.058 | 0.05 | 0.056 | 0.061 | 0.056 | [5 empty] | 16
| 1 | 4 | 0.06 | 0.055 | < 5 | < 0.02 | < 0.1 | < 0.5 | 0.05 | 0.06 | 0.052 | 0.05 | 0.055 | 0.06 | 0.07 | 0.06 | 0.05 | 0.062 | 0.05 | 0.055 | 0.056 | 0.053 | [5 empty] | 0.055625
| 1 | 5 | 0.059 | 0.055 | < 5 | < 0.02 | < 0.1 | < 0.5 | 0.05 | 0.06 | 0.049 | 0.05 | 0.051 | 0.06 | 0.07 | 0.06 | 0.05 | 0.057 | < 0.05 | 0.051 | 0.058 | 0.055 | [5 empty] | 94
| 1 | 6 | 0.056 | 0.055 | < 5 | < 0.02 | < 0.1 | < 0.5 | 0.05 | 0.05 | 0.053 | 0.05 | 0.049 | 0.06 | 0.07 | 0.06 | 0.05 | 0.059 | 0.05 | 0.065 | 0.051 | 0.063
| Mean |  | 0.0585 | 0.054167 |  |  |  |  | 0.048333 | 0.056667 | 0.051 | 0.05 | 0.053 | 0.06 | 0.068333 | 0.06 | 0.051667 | 0.057833 | 0.05 | 0.058167 | 0.0565 | 0.056333
| Median |  | 0.0585 | 0.055 |  |  |  |  | 0.05 | 0.06 | 0.0515 | 0.05 | 0.053 | 0.06 | 0.07 | 0.06 | 0.05 | 0.0575 | 0.05 | 0.058 | 0.0565 | 0.0555
| Std Dev. |  | 0.0018708 | 0.0020412 |  |  |  |  | 0.0040825 | 0.005164 | 0.0016733 | 7.6012e-18 | 0.0028983 | 0 | 0.0040825 | 0 | 0.0040825 | 0.0024833 | 8.4984e-18 | 0.0051153 | 0.0032711 | 0.003559
| Rel.Std.Dev. |  | 3.198% | 3.768% |  |  |  |  | 8.447% | 9.113% | 3.281% | 1.52e-14% | 5.468% | 0% | 5.974% | 0% | 7.902% | 4.294% | 1.7e-14% | 8.794% | 5.79% | 6.318%
| PDM3 |  | 5.169% | -2.622% |  |  |  |  | -13.11% | 1.873% | -8.315% | -10.11% | -4.719% | 7.865% | 22.85% | 7.865% | -7.116% | 3.97% | -10.11% | 4.569% | 1.573% | 1.273%
| Z-Score (Absolute) |  | 0.34 | 0.36 | 395.46 | 7.51 | 1.04 | 31.33 | 1.31 | 0.04 | 0.88 | 1.04 | 0.55 | 0.58 | 1.93 | 0.58 | 0.77 | 0.23 | 3.06 | 0.28 | 0.01 | 0.01
| Analytical results for K in OREAS 282 (Certified Value 0.412 wt.%) | [26 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.4 | 0.4 | 0.43 | 0.45 | 0.44 | 0.3645 | 0.42 | 0.42 | 0.401 | 0.477 | 0.4366 | 0.4384 | 0.44 | 0.42 | 0.445 | 0.38 | 0.465 | 0.41 | 0.36 | 0.41 | 0.43 | 0.42 | 0.39 |  |  | 1
| 1 | 2 | 0.39 | 0.39 | 0.41 | 0.43 | 0.43 | 0.3805 | 0.41 | 0.42 | 0.407 | 0.483 | 0.4392 | 0.4583 | 0.43 | 0.41 | 0.391 | 0.38 | 0.452 | 0.42 | 0.38 | 0.41 | 0.41 | 0.38 | 0.38 |  |  | #N/A
| 1 | 3 | 0.39 | 0.4 | 0.42 | 0.46 | 0.43 | 0.364 | 0.41 | 0.41 | 0.395 | 0.467 | 0.4378 | 0.4425 | 0.46 | 0.41 | 0.432 | 0.37 | 0.455 | 0.41 | 0.36 | 0.41 | 0.4 | 0.43 | 0.38 |  |  | 16
| 1 | 4 | 0.4 | 0.4 | 0.42 | 0.43 | 0.43 | 0.3715 | 0.4 | 0.43 | 0.406 | 0.474 | 0.4367 | 0.4381 | 0.42 | 0.41 | 0.366 | 0.38 | 0.454 | 0.42 | 0.37 | 0.4 | 0.41 | 0.42 | 0.39 |  |  | 0.4121
| 1 | 5 | 0.4 | 0.4 | 0.41 | 0.43 | 0.45 | 0.3755 | 0.4 | 0.43 | 0.408 | 0.48 | 0.4314 | 0.434 | 0.44 | 0.42 | 0.426 | 0.38 | 0.459 | 0.42 | 0.37 | 0.4 | 0.39 | 0.42 | 0.38 |  |  | 95
| 1 | 6 | 0.4 | 0.39 | 0.42 | 0.44 | 0.44 | 0.3635 | 0.41 | 0.42 | 0.399 | 0.47 | 0.4322 | 0.443 | 0.43 | 0.41 | 0.416 | 0.37 | 0.451 | 0.42 | 0.39 | 0.41 | 0.4 | 0.41 | 0.38
| Mean |  | 0.39667 | 0.39667 | 0.41833 | 0.44 | 0.43667 | 0.36992 | 0.40833 | 0.42167 | 0.40267 | 0.47517 | 0.43565 | 0.44238 | 0.43667 | 0.41333 | 0.41267 | 0.37667 | 0.456 | 0.41667 | 0.37167 | 0.40667 | 0.40667 | 0.41333 | 0.38333
| Median |  | 0.4 | 0.4 | 0.42 | 0.435 | 0.435 | 0.368 | 0.41 | 0.42 | 0.4035 | 0.4755 | 0.43665 | 0.44045 | 0.435 | 0.41 | 0.421 | 0.38 | 0.4545 | 0.42 | 0.37 | 0.41 | 0.405 | 0.42 | 0.38
| Std Dev. |  | 0.005164 | 0.005164 | 0.0075277 | 0.012649 | 0.008165 | 0.0070881 | 0.0075277 | 0.0075277 | 0.005164 | 0.006047 | 0.0031367 | 0.0084644 | 0.013663 | 0.005164 | 0.029173 | 0.005164 | 0.0052154 | 0.005164 | 0.01169 | 0.005164 | 0.013663 | 0.017512 | 0.005164
| Rel.Std.Dev. |  | 1.302% | 1.302% | 1.799% | 2.875% | 1.87% | 1.916% | 1.844% | 1.785% | 1.282% | 1.273% | 0.72% | 1.913% | 3.129% | 1.249% | 7.069% | 1.371% | 1.144% | 1.239% | 3.145% | 1.27% | 3.36% | 4.237% | 1.347%
| PDM3 |  | -3.744% | -3.744% | 1.514% | 6.771% | 5.963% | -10.24% | -0.9129% | 2.323% | -2.288% | 15.31% | 5.716% | 7.35% | 5.963% | 0.3004% | 0.1386% | -8.597% | 10.65% | 1.109% | -9.811% | -1.317% | -1.317% | 0.3004% | -6.979%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.2 | 1.08 | 0.94 | 1.76 | 0.2 | 0.34 | 0.43 | 2.5 | 0.9 | 1.18 | 0.94 | 0 | 0.03 | 1.48 | 1.73 | 0.13 | 1.69 | 0.27 | 0.27 | 0 | 1.21
| Analytical results for La in OREAS 282 (Certified Value 18.2 ppm) | [26 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 18.8 | 20.1 | 18 | 17.9 | 18.8 | 16 | 19.2 | 15 | 18.666 | 17.6 | 16.48 | 17.96 | 19.33 | 21.9 | 17.5 | 17.2 | 18.9 | 19.6 | 19 | 16.5 | 19 | 19.2 | 19.5 | 18.5 |  | 1
| 1 | 2 | 19 | 19.9 | 19 | 17.8 | 18.1 | 18 | 19 | 15 | 18.45 | 16.8 | 16.1 | 18.22 | 19.76 | 21.2 | 17.8 | 16.5 | 19.3 | 19.5 | 19.5 | 16.4 | 19 | 18.2 | 19.6 | 17.7 |  | 22
| 1 | 3 | 18.5 | 20.1 | 17 | 17.7 | 18.8 | 15.5 | 20.8 | 15 | 18.56 | 16.8 | 16.44 | 17.96 | 19.34 | 22.6 | 17.6 | 15.9 | 19.3 | 19.3 | 18.4 | 15.8 | 19 | 18.5 | 19.8 | 18.7 |  | 16
| 1 | 4 | 19.2 | 20.1 | 18 | 17.9 | 19 | 16.5 | 20 | 16 | 18.366 | 17.1 | 16.09 | 18.11 | 19.88 | 20.4 | 17.5 | 16.3 | 18.8 | 18.9 | 18.7 | 15.8 | 19 | 18.2 | 19.3 | 18 |  | 18.2275
| 1 | 5 | 18.6 | 20.2 | 17 | 18.2 | 19 | 16 | 20.2 | 16 | 18.646 | 17 | 16.25 | 17.82 | 19.45 | 21.5 | 18 | 17 | 19.4 | 19.5 | 18.4 | 16.2 | 19 | 17.7 | 19.1 | 18.5 |  | 96
| 1 | 6 | 18 | 19.9 | 18 | 18.1 | 18.6 | 16 | 19.8 | 15 | 18.903 | 16.6 | 16.15 | 18.01 | 19.22 | 22.5 | 17.8 | 16.7 | 19.3 | 19.4 | 18.6 | 17.2 | 19 | 18.1 | 19 | 19.9
| Mean |  | 18.6833 | 20.05 | 17.8333 | 17.9333 | 18.7167 | 16.3333 | 19.8333 | 15.3333 | 18.5985 | 16.9833 | 16.2517 | 18.0133 | 19.4967 | 21.6833 | 17.7 | 16.6 | 19.1667 | 19.3667 | 18.7667 | 16.3167 | 19 | 18.3167 | 19.3833 | 18.55
| Median |  | 18.7 | 20.1 | 18 | 17.9 | 18.8 | 16 | 19.9 | 15 | 18.603 | 16.9 | 16.2 | 17.985 | 19.395 | 21.7 | 17.7 | 16.6 | 19.3 | 19.45 | 18.65 | 16.3 | 19 | 18.2 | 19.4 | 18.5
| Std Dev. |  | 0.42151 | 0.12247 | 0.75277 | 0.18619 | 0.33714 | 0.8756 | 0.66232 | 0.5164 | 0.18821 | 0.34881 | 0.17151 | 0.13794 | 0.26357 | 0.83287 | 0.2 | 0.47329 | 0.25033 | 0.25033 | 0.42269 | 0.52313 | 0 | 0.50365 | 0.30605 | 0.75829
| Rel.Std.Dev. |  | 2.256% | 0.6108% | 4.221% | 1.038% | 1.801% | 5.361% | 3.339% | 3.368% | 1.012% | 2.054% | 1.055% | 0.7658% | 1.352% | 3.841% | 1.13% | 2.851% | 1.306% | 1.293% | 2.252% | 3.206% | 0% | 2.75% | 1.579% | 4.088%
| PDM3 |  | 2.501% | 9.999% | -2.163% | -1.614% | 2.684% | -10.39% | 8.81% | -15.88% | 2.035% | -6.826% | -10.84% | -1.175% | 6.963% | 18.96% | -2.894% | -8.929% | 5.152% | 6.25% | 2.958% | -10.48% | 4.238% | 0.4891% | 6.341% | 1.769%
| Z-Score (Absolute) |  | 0.07 | 1.25 | NA | 0.57 | 0.1 | 1.95 | 1.06 | NA | 0 | 1.39 | 2.02 | 0.5 | 0.77 | 2.65 | 0.77 | 1.72 | 0.49 | 0.66 | 0.14 | 1.96 | NA | 0.24 | 0.67 | 0.04
Results from laboratories 3, 9 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 282 (Certified Value 12.3 ppm) | [26 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 11.3 | 12.8 | 12 | 12 | 16.8 | 11 | < 10 | 15.3 | 12.55 | 13 | 11.7 | 12.15 | 12.7 | 15.7 | 13 | 11.8 | 15.4 | 13 | 13.5 | 11.6 | 12.2 |  |  |  |  | 1
| 1 | 2 | 11.8 | 11.6 | 12 | 9 | 16.4 | 12 | < 10 | 13.2 | 12.38 | 12.5 | 11.7 | 12.08 | 13.3 | 16 | 12 | 12.1 | 15.2 | 13 | 12.3 | 11.1 | 12.2 |  |  |  |  | #N/A
| 1 | 3 | 11.5 | 11.8 | 11 | 11 | 16.2 | 11 | 10 | 14.1 | 12.24 | 12.6 | 11.8 | 12.43 | 12.9 | 16.5 | 13 | 11.9 | 15.1 | 12 | 12.6 | 12.1 | 12.7 |  |  |  |  | 16
| 1 | 4 | 11.8 | 12.4 | 11 | 11 | 16.7 | 11 | 10 | 13.5 | 12.47 | 12.5 | 11.9 | 12.02 | 13.3 | 15.3 | 12 | 11.9 | 15.6 | 12 | 12.9 | 11.9 | 12.3 |  |  |  |  | 12.2831
| 1 | 5 | 11.8 | 12.2 | 11 | 12 | 16.7 | 11 | 10 | 12.4 | 12.97 | 12.6 | 11.5 | 12.12 | 12.6 | 15.7 | 12 | 12 | 15.3 | 12 | 12.5 | 11.4 | 12.9 |  |  |  |  | 97
| 1 | 6 | 11.7 | 11.8 | 11 | 10 | 16.9 | 11 | 10 | 11.8 | 11.9 | 12.6 | 11.5 | 12.01 | 12.9 | 16 | 12 | 11.7 | 15.1 | 12 | 12.7 | 11.5 | 13.6
| Mean |  | 11.65 | 12.1 | 11.3333 | 10.8333 | 16.6167 | 11.1667 | 10 | 13.3833 | 12.4183 | 12.6333 | 11.6833 | 12.135 | 12.95 | 15.8667 | 12.3333 | 11.9 | 15.2833 | 12.3333 | 12.75 | 11.6 | 12.65
| Median |  | 11.75 | 12 | 11 | 11 | 16.7 | 11 | 10 | 13.35 | 12.425 | 12.6 | 11.7 | 12.1 | 12.9 | 15.85 | 12 | 11.9 | 15.25 | 12 | 12.65 | 11.55 | 12.5
| Std Dev. |  | 0.20736 | 0.45166 | 0.5164 | 1.169 | 0.26394 | 0.40825 | 0 | 1.2416 | 0.35392 | 0.18619 | 0.16021 | 0.1545 | 0.29496 | 0.40332 | 0.5164 | 0.14142 | 0.19408 | 0.5164 | 0.41833 | 0.35777 | 0.54681
| Rel.Std.Dev. |  | 1.78% | 3.733% | 4.556% | 10.79% | 1.588% | 3.656% | 0% | 9.277% | 2.85% | 1.474% | 1.371% | 1.273% | 2.278% | 2.542% | 4.187% | 1.188% | 1.27% | 4.187% | 3.281% | 3.084% | 4.323%
| PDM3 |  | -5.154% | -1.49% | -7.732% | -11.8% | 35.28% | -9.089% | -18.59% | 8.958% | 1.101% | 2.852% | -4.883% | -1.205% | 5.43% | 29.18% | 0.4093% | -3.119% | 24.43% | 0.4093% | 3.802% | -5.561% | 2.987%
| Z-Score (Absolute) |  | 0.9 | 0.49 | NA | NA | 3.66 | NA | NA | 0.69 | 0.2 | 0 | 0.87 | 0.46 | 0.29 | 2.97 | NA | 0.67 | 2.44 | NA | 0.11 | 0.95 | 0.02
Results from laboratories 3, 4, 7, 8, 18 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 282 (Certified Value 0.15 ppm) | [26 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.18 | 2.19 | [14 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 04g | 0.25g | 0.25g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.25g | 0.4g | [14 empty] | 3
| 1 | 1 | 0.165 | 0.14 | 0.19 | 0.2 | 0.125 | 0.166 | 0.16 | 0.157 | 0.2 | 0.16 | 0.14 | [14 empty] | 1
| 1 | 2 | 0.165 | 0.12 | 0.17 | 0.2 | 0.13 | 0.16 | 0.15 | 0.163 | 0.2 | 0.16 | 0.15 | [14 empty] | 23
| 1 | 3 | 0.16 | 0.14 | 0.21 | 0.2 | 0.145 | 0.161 | 0.15 | 0.162 | 0.2 | 0.16 | 0.15 | [14 empty] | 16
| 1 | 4 | 0.165 | 0.13 | 0.19 | 0.2 | 0.14 | 0.164 | 0.15 | 0.158 | 0.2 | 0.16 | 0.13 | [14 empty] | 0.15352
| 1 | 5 | 0.165 | 0.12 | 0.2 | 0.2 | 0.14 | 0.167 | 0.15 | 0.152 | 0.2 | 0.16 | 0.14 | [14 empty] | 98
| 1 | 6 | 0.16 | 0.14 | 0.22 | 0.2 | 0.135 | 0.164 | 0.15 | 0.159 | 0.2 | 0.16 | 0.14
| Mean |  | 0.16333 | 0.13167 | 0.19667 | 0.2 | 0.13583 | 0.16367 | 0.15167 | 0.1585 | 0.2 | 0.16 | 0.14167
| Median |  | 0.165 | 0.135 | 0.195 | 0.2 | 0.1375 | 0.164 | 0.15 | 0.1585 | 0.2 | 0.16 | 0.14
| Std Dev. |  | 0.002582 | 0.0098319 | 0.017512 | 3.0405e-17 | 0.0073598 | 0.0027325 | 0.0040825 | 0.003937 | 3.0405e-17 | 0 | 0.0075277
| Rel.Std.Dev. |  | 1.581% | 7.467% | 8.904% | 1.52e-14% | 5.418% | 1.67% | 2.692% | 2.484% | 1.52e-14% | 0% | 5.314%
| PDM3 |  | 6.39% | -14.24% | 28.1% | 30.27% | -11.52% | 6.607% | -1.21% | 3.241% | 30.27% | 4.218% | -7.723%
| Z-Score (Absolute) |  | 0.48 | 2.65 | 3.77 | NA | 2.24 | 0.51 | 0.67 | 0 | NA | 0.15 | 1.66
Results from laboratories 7 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 282 (Certified Value 2.8 wt.%) | [26 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 2.89 | 2.8 | 2.93 | 2.77 | 3.1 | 2.7985 | 2.72 | 2.79 | 2.727 | 2.79 | 2.8244 | 2.54 | 3.49 | 2.83 | 3.07 | 2.79 | 3.03 | 2.8 | 2.57 | 2.95 | 3.28 | 2.79 | 2.75 |  |  | 1
| 1 | 2 | 2.86 | 2.81 | 2.95 | 2.76 | 3.02 | 2.849 | 2.73 | 2.8 | 2.779 | 2.69 | 2.8278 | 2.61 | 3.49 | 2.85 | 2.81 | 2.79 | 3.06 | 2.8 | 2.61 | 2.91 | 3.13 | 2.75 | 2.74 |  |  | #N/A
| 1 | 3 | 2.84 | 2.73 | 2.85 | 2.76 | 3.09 | 2.782 | 2.76 | 2.72 | 2.675 | 2.85 | 2.8363 | 2.55 | 3.74 | 2.81 | 3.01 | 2.77 | 3.05 | 2.75 | 2.63 | 2.95 | 3.09 | 2.73 | 2.76 |  |  | 16
| 1 | 4 | 2.87 | 2.8 | 2.93 | 2.73 | 3.09 | 2.792 | 2.82 | 2.83 | 2.726 | 2.68 | 2.8215 | 2.63 | 3.33 | 2.75 | 2.78 | 2.79 | 3.06 | 2.77 | 2.57 | 2.95 | 3.14 | 2.74 | 2.78 |  |  | 2.8034
| 1 | 5 | 2.91 | 2.75 | 2.88 | 2.76 | 3.06 | 2.7975 | 2.8 | 2.82 | 2.755 | 2.7 | 2.8141 | 2.57 | 3.43 | 2.79 | 3.04 | 2.84 | 3.01 | 2.76 | 2.56 | 2.91 | 3.03 | 2.74 | 2.76 |  |  | 99
| 1 | 6 | 2.86 | 2.78 | 2.87 | 2.72 | 3.07 | 2.7755 | 2.71 | 2.77 | 2.706 | 2.84 | 2.8223 | 2.61 | 3.67 | 2.81 | 2.96 | 2.76 | 3.05 | 2.77 | 2.73 | 2.96 | 3.09 | 2.72 | 2.75
| Mean |  | 2.8717 | 2.7783 | 2.9017 | 2.75 | 3.0717 | 2.7991 | 2.7567 | 2.7883 | 2.728 | 2.7583 | 2.8244 | 2.585 | 3.525 | 2.8067 | 2.945 | 2.79 | 3.0433 | 2.775 | 2.6117 | 2.9383 | 3.1267 | 2.745 | 2.7567
| Median |  | 2.865 | 2.79 | 2.905 | 2.76 | 3.08 | 2.7947 | 2.745 | 2.795 | 2.7265 | 2.745 | 2.8234 | 2.59 | 3.49 | 2.81 | 2.985 | 2.79 | 3.05 | 2.77 | 2.59 | 2.95 | 3.11 | 2.74 | 2.755
| Std Dev. |  | 0.024833 | 0.031885 | 0.040208 | 0.02 | 0.029269 | 0.026049 | 0.045019 | 0.039707 | 0.036409 | 0.077825 | 0.0073767 | 0.036742 | 0.15281 | 0.034448 | 0.12211 | 0.027568 | 0.019664 | 0.020736 | 0.064005 | 0.022286 | 0.084538 | 0.02429 | 0.013663
| Rel.Std.Dev. |  | 0.8648% | 1.148% | 1.386% | 0.7273% | 0.9529% | 0.9306% | 1.633% | 1.424% | 1.335% | 2.821% | 0.2612% | 1.421% | 4.335% | 1.227% | 4.146% | 0.9881% | 0.6461% | 0.7473% | 2.451% | 0.7585% | 2.704% | 0.8849% | 0.4956%
| PDM3 |  | 2.434% | -0.8949% | 3.505% | -1.906% | 9.569% | -0.1547% | -1.668% | -0.5382% | -2.69% | -1.608% | 0.7484% | -7.791% | 25.74% | 0.1158% | 5.05% | -0.4787% | 8.558% | -1.014% | -6.84% | 4.812% | 11.53% | -2.084% | -1.668%
| Z-Score (Absolute) |  | 1.22 | 0.17 | 1.67 | 0.6 | 4.22 | 0.14 | 0.5 | 0.02 | 0.93 | 0.47 | 0.52 | 3.07 | 11.01 | 0.25 | 2.32 | 0 | 3.8 | 0.22 | 2.67 | 2.22 | 5.04 | 0.67 | 0.5
| Analytical results for Mn in OREAS 282 (Certified Value 0.024 wt.%) | [26 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.0227 | 0.0277 | 0.0208 | 0.0229 | 0.0237 | 0.024 | 0.023 | 0.0245 | 0.02528 | 0.0232 | 0.0249 | 0.0234 | 0.0294 | 0.0238 | 0.0261 | 0.0225 | 0.0207 | 0.0258 | 0.0212 | 0.026 | 0.0267 | 0.0233 | 0.0222 |  |  | 1
| 1 | 2 | 0.0222 | 0.0259 | 0.0221 | 0.0238 | 0.0234 | 0.024 | 0.022 | 0.0246 | 0.02523 | 0.0225 | 0.0248 | 0.024 | 0.0294 | 0.0238 | 0.0243 | 0.0227 | 0.0209 | 0.0256 | 0.0218 | 0.0261 | 0.0256 | 0.023 | 0.0223 |  |  | 24
| 1 | 3 | 0.0221 | 0.0274 | 0.021 | 0.0229 | 0.0236 | 0.024 | 0.023 | 0.0239 | 0.02463 | 0.0231 | 0.0246 | 0.0234 | 0.0303 | 0.0236 | 0.0258 | 0.0223 | 0.0204 | 0.0256 | 0.0215 | 0.0262 | 0.0254 | 0.0229 | 0.022 |  |  | 16
| 1 | 4 | 0.0224 | 0.0291 | 0.0212 | 0.0227 | 0.0242 | 0.024 | 0.023 | 0.0252 | 0.02485 | 0.0224 | 0.0236 | 0.0242 | 0.0281 | 0.0231 | 0.0241 | 0.0224 | 0.0207 | 0.0253 | 0.0211 | 0.026 | 0.0256 | 0.0229 | 0.0228 |  |  | 0.023668
| 1 | 5 | 0.0227 | 0.0276 | 0.0214 | 0.0231 | 0.0234 | 0.024 | 0.023 | 0.0248 | 0.02522 | 0.0228 | 0.0233 | 0.0236 | 0.0292 | 0.0232 | 0.0262 | 0.023 | 0.021 | 0.0257 | 0.0211 | 0.0259 | 0.0245 | 0.0229 | 0.0226 |  |  | 100
| 1 | 6 | 0.0224 | 0.0271 | 0.0217 | 0.0236 | 0.0237 | 0.0235 | 0.022 | 0.0242 | 0.02505 | 0.0222 | 0.0239 | 0.024 | 0.0299 | 0.0235 | 0.0259 | 0.0223 | 0.0207 | 0.0255 | 0.0228 | 0.0265 | 0.0255 | 0.0228 | 0.0226
| Mean |  | 0.022417 | 0.027467 | 0.021367 | 0.023167 | 0.023667 | 0.023917 | 0.022667 | 0.024533 | 0.025043 | 0.0227 | 0.024183 | 0.023767 | 0.029383 | 0.0235 | 0.0254 | 0.022533 | 0.020733 | 0.025583 | 0.021583 | 0.026117 | 0.02555 | 0.022967 | 0.022417
| Median |  | 0.0224 | 0.0275 | 0.0213 | 0.023 | 0.02365 | 0.024 | 0.023 | 0.02455 | 0.025135 | 0.02265 | 0.02425 | 0.0238 | 0.0294 | 0.02355 | 0.02585 | 0.02245 | 0.0207 | 0.0256 | 0.02135 | 0.02605 | 0.02555 | 0.0229 | 0.02245
| Std Dev. |  | 0.00024833 | 0.0010328 | 0.0004761 | 0.00043665 | 0.00029439 | 0.00020412 | 0.0005164 | 0.00045461 | 0.00025703 | 0.0004 | 0.00067355 | 0.00034448 | 0.00074677 | 0.00029665 | 0.00094234 | 0.00027325 | 0.00020656 | 0.00017224 | 0.00065549 | 0.0002137 | 0.00070071 | 0.00017512 | 0.00029944
| Rel.Std.Dev. |  | 1.108% | 3.76% | 2.228% | 1.885% | 1.244% | 0.8535% | 2.278% | 1.853% | 1.026% | 1.762% | 2.785% | 1.449% | 2.541% | 1.262% | 3.71% | 1.213% | 0.9963% | 0.6733% | 3.037% | 0.8182% | 2.743% | 0.7625% | 1.336%
| PDM3 |  | -5.289% | 16.05% | -9.725% | -2.12% | -0.007682% | 1.049% | -4.233% | 3.654% | 5.809% | -4.092% | 2.175% | 0.4148% | 24.15% | -0.7119% | 7.316% | -4.796% | -12.4% | 8.09% | -8.81% | 10.34% | 7.949% | -2.965% | -5.289%
| Z-Score (Absolute) |  | 0.67 | 2.05 | 1.24 | 0.27 | 0 | 0.13 | 0.54 | 0.47 | 0.74 | 0.52 | 0.28 | 0.05 | 3.08 | 0.09 | 0.94 | 0.61 | 1.58 | 1.03 | 1.12 | 1.32 | 1.02 | 0.38 | 0.67
| Analytical results for Mo in OREAS 282 (Certified Value 13.1 ppm) | [26 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 13.15 | 12.5 | 11 | 13 | 14 | 12.2 | 14.7 | 13.71 | 13.88 | 12.1 | 12.62 | 13.7 | 15.7 | 12.53 | 14.6 | 11.95 | 13.2 | 13.37 | 16.7 | 12.65 | 14 | 14 | 13.93 | 13.75 |  | 1
| 1 | 2 | 13.45 | 12.2 | 11 | 13 | 16 | 11 | 14.6 | 13.66 | 14.06 | 12.2 | 12.58 | 13.8 | 15.7 | 12.45 | 14.3 | 11.95 | 13.2 | 13.31 | 16.6 | 12.88 | 14 | 12.8 | 13.1 | 12.6 |  | 25
| 1 | 3 | 13.05 | 12.2 | 11 | 13 | 15 | 11.8 | 14.2 | 13.6 | 14.02 | 12.1 | 12.61 | 13.4 | 16.1 | 12.66 | 14.4 | 11.85 | 13.1 | 13.26 | 16.5 | 12.54 | 14 | 12.9 | 13.51 | 13.6 |  | 16
| 1 | 4 | 13.45 | 12.6 | 11 | 13 | 15 | 12.4 | 14.7 | 13.88 | 14.26 | 11.8 | 12.6 | 14.1 | 15 | 12.42 | 13.8 | 12.2 | 13.1 | 13.31 | 17.1 | 12.65 | 14 | 12.55 | 13.53 | 12.4 |  | 13.0859
| 1 | 5 | 13.25 | 12.7 | 11 | 13 | 14 | 12.2 | 14.6 | 13.72 | 14.31 | 12.2 | 12.51 | 13.6 | 15.4 | 12.63 | 15.3 | 12.05 | 12.9 | 13.18 | 16.7 | 12.31 | 14 | 12.5 | 13.55 | 13.35 |  | 101
| 1 | 6 | 12.75 | 12.4 | 11 | 13 | 15 | 11 | 14.8 | 13.74 | 13.99 | 11.8 | 12.44 | 13.6 | 16.1 | 12.7 | 14.2 | 11.9 | 12.9 | 12.96 | 16.4 | 12.3 | 14 | 13.1 | 13.82 | 14.05
| Mean |  | 13.1833 | 12.4333 | 11 | 13 | 14.8333 | 11.7667 | 14.6 | 13.7183 | 14.0867 | 12.0333 | 12.56 | 13.7 | 15.6667 | 12.565 | 14.4333 | 11.9833 | 13.0667 | 13.2317 | 16.6667 | 12.555 | 14 | 12.975 | 13.5733 | 13.2917
| Median |  | 13.2 | 12.45 | 11 | 13 | 15 | 12 | 14.65 | 13.715 | 14.04 | 12.1 | 12.59 | 13.65 | 15.7 | 12.58 | 14.35 | 11.95 | 13.1 | 13.285 | 16.65 | 12.595 | 14 | 12.85 | 13.54 | 13.475
| Std Dev. |  | 0.26583 | 0.20656 | 0 | 0 | 0.75277 | 0.62503 | 0.20976 | 0.093897 | 0.16561 | 0.18619 | 0.070711 | 0.23664 | 0.42269 | 0.11572 | 0.50067 | 0.12517 | 0.13663 | 0.14743 | 0.24221 | 0.22314 | 0 | 0.54932 | 0.28932 | 0.65682
| Rel.Std.Dev. |  | 2.016% | 1.661% | 0% | 0% | 5.075% | 5.312% | 1.437% | 0.6845% | 1.176% | 1.547% | 0.563% | 1.727% | 2.698% | 0.9209% | 3.469% | 1.045% | 1.046% | 1.114% | 1.453% | 1.777% | 0% | 4.234% | 2.132% | 4.942%
| PDM3 |  | 0.7447% | -4.987% | -15.94% | -0.6563% | 13.35% | -10.08% | 11.57% | 4.833% | 7.648% | -8.043% | -4.019% | 4.693% | 19.72% | -3.98% | 10.3% | -8.425% | -0.1468% | 1.114% | 27.36% | -4.057% | 6.986% | -0.8473% | 3.725% | 1.573%
| Z-Score (Absolute) |  | 0 | 0.81 | NA | 0.2 | NA | 1.53 | 1.53 | 0.58 | 0.98 | 1.24 | 0.67 | 0.56 | 2.69 | 0.67 | 1.35 | 1.3 | 0.13 | 0.05 | 3.77 | 0.68 | NA | 0.23 | 0.42 | 0.12
Results from laboratories 3, 7 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 282 (Indicative Value 0.02 wt.%) | [26 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.03 | 0.014 | 0.02 | 0.0147 | 0.02 | 0.021 | 0.02 | 0.0155 | 0.01 | 0.02 | 0.014 | 0.014 | 0.02 | 0.02 | 0.02 | 0.01 |  |  | 1
| 1 | 2 | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.0215 | 0.02 | 0.02 | 0.014 | 0.02 | 0.0153 | 0.02 | 0.02 | 0.02 | 0.0135 | 0.01 | 0.02 | 0.014 | 0.014 | 0.02 | 0.02 | 0.02 | 0.01 |  |  | 8
| 1 | 3 | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.023 | 0.02 | 0.02 | 0.014 | 0.02 | 0.0156 | 0.02 | 0.023 | 0.02 | 0.016 | 0.01 | 0.02 | 0.014 | 0.014 | 0.02 | 0.02 | 0.02 | 0.01 |  |  | 16
| 1 | 4 | 0.01 | 0.01 | 0.02 | 0.02 | 0.02 | 0.0205 | 0.02 | 0.02 | 0.014 | 0.02 | 0.0149 | 0.02 | 0.02 | 0.02 | 0.013 | 0.01 | 0.02 | 0.014 | 0.014 | 0.02 | 0.02 | 0.02 | 0.01 |  |  | 0.020126
| 1 | 5 | 0.01 | 0.01 | 0.02 | 0.02 | 0.02 | 0.0205 | 0.02 | 0.02 | 0.014 | 0.02 | 0.0148 | 0.02 | 0.021 | 0.02 | 0.0155 | 0.01 | 0.02 | 0.014 | 0.014 | 0.02 | 0.02 | 0.02 | 0.01 |  |  | 14
| 1 | 6 | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.0205 | 0.02 | 0.02 | 0.013 | 0.02 | 0.015 | 0.02 | 0.022 | 0.02 | 0.0145 | 0.01 | 0.02 | 0.014 | 0.014 | 0.02 | 0.02 | 0.02 | 0.01
| Mean |  | 0.01 | 0.016667 | 0.02 | 0.02 | 0.02 | 0.021 | 0.02 | 0.021667 | 0.013833 | 0.02 | 0.01505 | 0.02 | 0.021167 | 0.02 | 0.014667 | 0.01 | 0.02 | 0.014 | 0.014 | 0.02 | 0.02 | 0.02 | 0.01
| Median |  | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.0205 | 0.02 | 0.02 | 0.014 | 0.02 | 0.01495 | 0.02 | 0.021 | 0.02 | 0.015 | 0.01 | 0.02 | 0.014 | 0.014 | 0.02 | 0.02 | 0.02 | 0.01
| Std Dev. |  | 0 | 0.005164 | 0 | 0 | 0 | 0.0010954 | 0 | 0.0040825 | 0.00040825 | 0 | 0.00033912 | 0 | 0.001169 | 0 | 0.0012111 | 0 | 0 | 1.9003e-18 | 1.9003e-18 | 0 | 0 | 0 | 0
| Rel.Std.Dev. |  | 0% | 30.98% | 0% | 0% | 0% | 5.216% | 0% | 18.84% | 2.951% | 0% | 2.253% | 0% | 5.523% | 0% | 8.257% | 0% | 0% | 1.357e-14% | 1.357e-14% | 0% | 0% | 0% | 0%
| PDM3 |  | -50.31% | -17.19% | -0.627% | -0.627% | -0.627% | 4.342% | -0.627% | 7.654% | -31.27% | -0.627% | -25.22% | -0.627% | 5.17% | -0.627% | -27.13% | -50.31% | -0.627% | -30.44% | -30.44% | -0.627% | -0.627% | -0.627% | -50.31%
| Z-Score (Absolute) |  | 6.74 | 2.25 | 0 | 0 | 0 | 0.67 | 0 | 1.12 | 4.16 | 0 | 3.34 | 0 | 0.79 | 0 | 3.6 | 6.74 | 0 | 4.05 | 4.05 | 0 | 0 | 0 | 6.74
| Analytical results for Nb in OREAS 282 (Certified Value 0.17 ppm) | [26 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.16 | 0.4 | < 5 | 0.21 | 0.14 | < 1 | < 0.1 | 0.14 | 0.1 | 0.3 | 0.198 | 0.2 | 0.2 | 0.22 | 0.3 | 0.11 | < 0.2 | < 1 | 0.18 | 0.16 | 0.25 |  |  |  |  | 1
| 1 | 2 | 0.15 | 0.2 | < 5 | 0.15 | 0.16 | < 1 | 0.1 | 0.14 | 0.11 | 0.3 | 0.21 | 0.23 | 0.2 | 0.22 | 0.3 | 0.1 | < 0.2 | < 1 | 0.18 | 0.14 | 0.25 |  |  |  |  | 26
| 1 | 3 | 0.16 | 0.4 | < 5 | 0.22 | 0.17 | < 1 | < 0.1 | 0.12 | 0.09 | 0.3 | 0.196 | 0.22 | 0.2 | 0.23 | 0.3 | 0.11 | < 0.2 | < 1 | 0.18 | 0.17 | 0.25 |  |  |  |  | 16
| 1 | 4 | 0.15 | < 0.2 | < 5 | 0.16 | 0.27 | < 1 | < 0.1 | 0.13 | 0.08 | 0.3 | 0.202 | 0.22 | 0.2 | 0.24 | 0.2 | 0.1 | < 0.2 | < 1 | 0.17 | 0.13 | 0.24 |  |  |  |  | 0.17204
| 1 | 5 | 0.17 | < 0.2 | < 5 | 0.19 | 0.24 | < 1 | < 0.1 | 0.12 | 0.1 | 0.3 | 0.192 | 0.22 | 0.2 | 0.21 | 0.3 | 0.1 | < 0.2 | < 1 | 0.14 | 0.11 | 0.26 |  |  |  |  | 102
| 1 | 6 | 0.17 | 0.2 | < 5 | 0.21 | 0.1 | < 1 | < 0.1 | 0.12 | 0.09 | 0.3 | 0.199 | 0.22 | 0.3 | 0.22 | 0.3 | 0.09 | < 0.2 | < 1 | 0.16 | 0.17 | 0.27
| Mean |  | 0.16 | 0.3 |  | 0.19 | 0.18 |  | 0.1 | 0.12833 | 0.095 | 0.3 | 0.1995 | 0.21833 | 0.21667 | 0.22333 | 0.28333 | 0.10167 |  |  | 0.16833 | 0.14667 | 0.25333
| Median |  | 0.16 | 0.3 |  | 0.2 | 0.165 |  | 0.1 | 0.125 | 0.095 | 0.3 | 0.1985 | 0.22 | 0.2 | 0.22 | 0.3 | 0.1 |  |  | 0.175 | 0.15 | 0.25
| Std Dev. |  | 0.0089443 | 0.11547 |  | 0.028983 | 0.063561 |  |  | 0.0098319 | 0.010488 | 0 | 0.0061237 | 0.0098319 | 0.040825 | 0.010328 | 0.040825 | 0.0075277 |  |  | 0.016021 | 0.024221 | 0.010328
| Rel.Std.Dev. |  | 5.59% | 38.49% |  | 15.25% | 35.31% |  |  | 7.661% | 11.04% | 0% | 3.07% | 4.503% | 18.84% | 4.624% | 14.41% | 7.404% |  |  | 9.517% | 16.51% | 4.077%
| PDM3 |  | -6.999% | 74.38% |  | 10.44% | 4.626% |  | -41.87% | -25.41% | -44.78% | 74.38% | 15.96% | 26.91% | 25.94% | 29.81% | 64.69% | -40.91% |  |  | -2.155% | -14.75% | 47.25%
| Z-Score (Absolute) |  | 0.47 | 0.67 | 35.95 | 0 | 0.16 | 4.82 | NA | 0.96 | 1.48 | NA | 0.15 | 0.44 | NA | 0.52 | NA | 1.37 | 1.4 | 4.82 | 0.34 | 0.67 | 0.99
Results from laboratories 9, 13, 16 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 282 (Certified Value 16.2 ppm) | [26 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 1
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | [17 empty] | 2
| 1 | 1 | 17.3 | 14.5 | 14 | 17.005 | 18.09 | 17.32 | 18 | 14.1 | [17 empty] | 1
| 1 | 2 | 16.7 | 16.5 | 14 | 16.738 | 17.4 | 18.13 | 18.1 | 13.6 | [17 empty] | 27
| 1 | 3 | 17 | 14 | 15 | 16.652 | 17.61 | 17.1 | 18.2 | 13.3 | [17 empty] | 16
| 1 | 4 | 17.1 | 15 | 14.4 | 16.813 | 17.74 | 17.77 | 17.1 | 13.3 | [17 empty] | 16.245
| 1 | 5 | 16.9 | 14.5 | 14.8 | 16.845 | 17.46 | 17.28 | 17.5 | 14.2 | [17 empty] | 103
| 1 | 6 | 16.7 | 15 | 14.5 | 16.895 | 17.22 | 17.89 | 18.9 | 13.6
| Mean |  | 16.95 | 14.9167 | 14.45 | 16.8247 | 17.5867 | 17.5817 | 17.9667 | 13.6833
| Median |  | 16.95 | 14.75 | 14.45 | 16.829 | 17.535 | 17.545 | 18.05 | 13.6
| Std Dev. |  | 0.23452 | 0.8612 | 0.40866 | 0.12275 | 0.30421 | 0.40563 | 0.6186 | 0.38687
| Rel.Std.Dev. |  | 1.384% | 5.773% | 2.828% | 0.7296% | 1.73% | 2.307% | 3.443% | 2.827%
| PDM3 |  | 4.34% | -8.177% | -11.05% | 3.569% | 8.259% | 8.228% | 10.6% | -15.77%
| Z-Score (Absolute) |  | 0.05 | 1.49 | 1.85 | 0.05 | 0.53 | 0.53 | 0.82 | 2.43
| Analytical results for Ni in OREAS 282 (Certified Value 21 ppm) | [26 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 21 | 24 | 18 | 19 | 21.2 | 21 | 23 | 27.2 | 21.2 | 20.1 | 20 | 18.88 | 22 | 25.4 | 20 | 38 | 20.6 | 22 | 22.4 | 18.1 | 23 | 26.2 | 20.9 | 20.9 |  | 1
| 1 | 2 | 21.1 | 22 | 19 | 20 | 21.9 | 21 | 22 | 28.7 | 21.2 | 19.4 | 19 | 18.98 | 22.3 | 24.8 | 19 | 28 | 21.1 | 22 | 22.4 | 18 | 22 | 24.7 | 21.5 | 19.8 |  | 28
| 1 | 3 | 20.6 | 22 | 18 | 19 | 21.4 | 20 | 22 | 28.2 | 20.8 | 19.9 | 20 | 19.07 | 21.7 | 25.3 | 20 | 34 | 20.6 | 22 | 23.4 | 18 | 22 | 25.6 | 21.6 | 20.6 |  | 16
| 1 | 4 | 21.4 | 23 | 18 | 19 | 20.9 | 20 | 23 | 28.6 | 21 | 20.8 | 19 | 18.67 | 22.2 | 24.1 | 20 | 26 | 20.7 | 22 | 21.3 | 18 | 22 | 25.1 | 21.1 | 19.8 |  | 21.0368
| 1 | 5 | 21 | 23 | 18 | 20 | 21 | 20 | 23 | 28.1 | 21.5 | 20.2 | 20 | 18.21 | 21.7 | 25.1 | 20 | 26 | 21.1 | 21 | 20.3 | 17.8 | 22 | 24.1 | 21.3 | 20.5 |  | 104
| 1 | 6 | 20.2 | 22 | 18 | 20 | 20.1 | 20 | 20 | 27.3 | 19.4 | 20.3 | 19 | 18.83 | 21.9 | 25.8 | 19 | 34 | 20.7 | 21 | 21.6 | 17.9 | 24 | 25 | 20.6 | 21.6
| Mean |  | 20.8833 | 22.6667 | 18.1667 | 19.5 | 21.0833 | 20.3333 | 22.1667 | 28.0167 | 20.85 | 20.1167 | 19.5 | 18.7733 | 21.9667 | 25.0833 | 19.6667 | 31 | 20.8 | 21.6667 | 21.9 | 17.9667 | 22.5 | 25.1167 | 21.1667 | 20.5333
| Median |  | 21 | 22.5 | 18 | 19.5 | 21.1 | 20 | 22.5 | 28.15 | 21.1 | 20.15 | 19.5 | 18.855 | 21.95 | 25.2 | 20 | 31 | 20.7 | 22 | 22 | 18 | 22 | 25.05 | 21.2 | 20.55
| Std Dev. |  | 0.42151 | 0.8165 | 0.40825 | 0.54772 | 0.59805 | 0.5164 | 1.169 | 0.63692 | 0.74766 | 0.46224 | 0.54772 | 0.30768 | 0.25033 | 0.58452 | 0.5164 | 5.02 | 0.23664 | 0.5164 | 1.0733 | 0.10328 | 0.83666 | 0.72503 | 0.37771 | 0.68605
| Rel.Std.Dev. |  | 2.018% | 3.602% | 2.247% | 2.809% | 2.837% | 2.54% | 5.274% | 2.273% | 3.586% | 2.298% | 2.809% | 1.639% | 1.14% | 2.33% | 2.626% | 16.19% | 1.138% | 2.383% | 4.901% | 0.5748% | 3.718% | 2.887% | 1.784% | 3.341%
| PDM3 |  | -0.7295% | 7.748% | -13.64% | -7.305% | 0.2213% | -3.344% | 5.371% | 33.18% | -0.8879% | -4.374% | -7.305% | -10.76% | 4.42% | 19.24% | -6.513% | 47.36% | -1.126% | 2.994% | 4.103% | -14.59% | 6.955% | 19.39% | 0.6174% | -2.393%
| Z-Score (Absolute) |  | 0.05 | 0.91 | 1.52 | 0.8 | 0.05 | 0.35 | 0.64 | 3.79 | 0.07 | 0.47 | 0.8 | 1.19 | 0.53 | 2.21 | 0.71 | 5.4 | 0.1 | 0.37 | 0.49 | 1.63 | 0.82 | 2.23 | 0.1 | 0.24
| Analytical results for P in OREAS 282 (Certified Value 0.031 wt.%) | [26 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.031 | 0.034 | 0.0255 | 0.03 | 0.034 | 0.0145 | 0.032 | 0.04 | 0.0342 | 0.029 | 0.0275 | 0.0292 | 0.026 | 0.03 | 0.0285 | 0.031 | 0.028 | 0.032 | 0.029 | 0.032 | 0.037 | 0.0339 | 0.03 |  |  | 1
| 1 | 2 | 0.03 | 0.032 | 0.0261 | 0.03 | 0.034 | 0.0135 | 0.032 | 0.04 | 0.0331 | 0.029 | 0.028 | 0.0301 | 0.025 | 0.03 | 0.03 | 0.031 | 0.0289 | 0.033 | 0.028 | 0.0314 | 0.035 | 0.0318 | 0.031 |  |  | #N/A
| 1 | 3 | 0.03 | 0.03 | 0.0249 | 0.03 | 0.035 | 0.0145 | 0.032 | 0.039 | 0.0318 | 0.029 | 0.0276 | 0.0294 | 0.027 | 0.03 | 0.029 | 0.031 | 0.0282 | 0.032 | 0.028 | 0.0315 | 0.035 | 0.0352 | 0.031 |  |  | 16
| 1 | 4 | 0.031 | 0.034 | 0.026 | 0.03 | 0.036 | 0.0135 | 0.032 | 0.041 | 0.0333 | 0.029 | 0.0271 | 0.0302 | 0.026 | 0.03 | 0.031 | 0.032 | 0.029 | 0.031 | 0.029 | 0.0325 | 0.035 | 0.0344 | 0.031 |  |  | 0.030541
| 1 | 5 | 0.031 | 0.032 | 0.0246 | 0.03 | 0.034 | 0.014 | 0.032 | 0.041 | 0.0337 | 0.029 | 0.0277 | 0.0295 | 0.026 | 0.03 | 0.0285 | 0.031 | 0.0289 | 0.033 | 0.028 | 0.0319 | 0.034 | 0.0348 | 0.031 |  |  | 105
| 1 | 6 | 0.031 | 0.03 | 0.0255 | 0.03 | 0.033 | 0.014 | 0.034 | 0.039 | 0.0324 | 0.029 | 0.0276 | 0.0301 | 0.025 | 0.03 | 0.029 | 0.031 | 0.028 | 0.03 | 0.03 | 0.0321 | 0.035 | 0.0336 | 0.031
| Mean |  | 0.030667 | 0.032 | 0.025433 | 0.03 | 0.034333 | 0.014 | 0.032333 | 0.04 | 0.033083 | 0.029 | 0.027583 | 0.02975 | 0.025833 | 0.03 | 0.029333 | 0.031167 | 0.0285 | 0.031833 | 0.028667 | 0.0319 | 0.035167 | 0.03395 | 0.030833
| Median |  | 0.031 | 0.032 | 0.0255 | 0.03 | 0.034 | 0.014 | 0.032 | 0.04 | 0.0332 | 0.029 | 0.0276 | 0.0298 | 0.026 | 0.03 | 0.029 | 0.031 | 0.02855 | 0.032 | 0.0285 | 0.03195 | 0.035 | 0.03415 | 0.031
| Std Dev. |  | 0.0005164 | 0.0017889 | 0.00059217 | 0 | 0.0010328 | 0.00044721 | 0.0008165 | 0.00089443 | 0.00087044 | 0 | 0.00029269 | 0.00043243 | 0.00075277 | 0 | 0.00098319 | 0.00040825 | 0.00048166 | 0.001169 | 0.0008165 | 0.00040497 | 0.00098319 | 0.0012029 | 0.00040825
| Rel.Std.Dev. |  | 1.684% | 5.59% | 2.328% | 0% | 3.008% | 3.194% | 2.525% | 2.236% | 2.631% | 0% | 1.061% | 1.454% | 2.914% | 0% | 3.352% | 1.31% | 1.69% | 3.672% | 2.848% | 1.269% | 2.796% | 3.543% | 1.324%
| PDM3 |  | 0.4106% | 4.776% | -16.72% | -1.772% | 12.42% | -54.16% | 5.868% | 30.97% | 8.323% | -5.047% | -9.685% | -2.591% | -15.41% | -1.772% | -3.955% | 2.048% | -6.684% | 4.231% | -6.138% | 4.449% | 15.14% | 11.16% | 0.9563%
| Z-Score (Absolute) |  | 0 | 0.54 | 2.12 | 0.27 | 1.48 | 6.74 | 0.67 | 3.78 | 0.98 | 0.67 | 1.25 | 0.37 | 1.96 | 0.27 | 0.54 | 0.2 | 0.88 | 0.47 | 0.81 | 0.5 | 1.82 | 1.33 | 0.07
| Analytical results for Pb in OREAS 282 (Certified Value 17.5 ppm) | [26 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 17.7 | 16 | 16 | 17.3 | 19.6 | 15 | 15 | 15.4 | 17.7 | 16.6 | 16.9 | 17.9 | 19.4 | 17.6 | 18 | 16.5 | 17.4 | 19.71 | 16.42 | 20 | 18.3 | 18.3 | 18 |  |  | 1
| 1 | 2 | 17.3 | 16 | 17 | 17 | 20 | 17 | 15 | 15.6 | 17.7 | 16 | 16.9 | 20.3 | 18.8 | 17.7 | 17 | 17 | 17.7 | 19.42 | 17.24 | 21 | 17.1 | 18.7 | 17.2 |  |  | 29
| 1 | 3 | 17.4 | 16 | 16 | 17.1 | 20.6 | 15 | 15 | 15.1 | 17.5 | 15.9 | 16.4 | 18.8 | 19.3 | 17.8 | 17 | 16.8 | 17.8 | 18.83 | 16.93 | 20 | 17.6 | 18.7 | 17 |  |  | 16
| 1 | 4 | 17.2 | 16 | 16 | 17.8 | 20.8 | 15 | 15 | 15.8 | 17.6 | 16.2 | 16.1 | 18.5 | 18.2 | 17.5 | 17 | 17.2 | 17.7 | 18.73 | 17.14 | 20 | 17.2 | 18.3 | 16.9 |  |  | 17.4767
| 1 | 5 | 17.3 | 16 | 16 | 17.1 | 19.7 | 15 | 15 | 15.7 | 17.4 | 16.2 | 15.7 | 19.5 | 18.8 | 17.7 | 18 | 16.6 | 17.9 | 18.98 | 16.52 | 19 | 16.4 | 18.4 | 17 |  |  | 106
| 1 | 6 | 16.9 | 16 | 17 | 17.5 | 21 | 15 | 16 | 15.4 | 17.6 | 16.2 | 15.6 | 20.1 | 19.4 | 17.9 | 18 | 17.5 | 17.5 | 19.04 | 17 | 20 | 17.2 | 17.7 | 18.2
| Mean |  | 17.3 | 16 | 16.3333 | 17.3 | 20.2833 | 15.3333 | 15.1667 | 15.5 | 17.5833 | 16.1833 | 16.2667 | 19.1833 | 18.9833 | 17.7 | 17.5 | 16.9333 | 17.6667 | 19.1183 | 16.875 | 20 | 17.3 | 18.35 | 17.3833
| Median |  | 17.3 | 16 | 16 | 17.2 | 20.3 | 15 | 15 | 15.5 | 17.6 | 16.2 | 16.25 | 19.15 | 19.05 | 17.7 | 17.5 | 16.9 | 17.7 | 19.01 | 16.965 | 20 | 17.2 | 18.35 | 17.1
| Std Dev. |  | 0.26077 | 0 | 0.5164 | 0.30332 | 0.5947 | 0.8165 | 0.40825 | 0.25298 | 0.1169 | 0.24014 | 0.56804 | 0.94322 | 0.47504 | 0.14142 | 0.54772 | 0.37771 | 0.18619 | 0.37424 | 0.33321 | 0.63246 | 0.6261 | 0.36742 | 0.56716
| Rel.Std.Dev. |  | 1.507% | 0% | 3.162% | 1.753% | 2.932% | 5.325% | 2.692% | 1.632% | 0.6649% | 1.484% | 3.492% | 4.917% | 2.502% | 0.799% | 3.13% | 2.231% | 1.054% | 1.958% | 1.975% | 3.162% | 3.619% | 2.002% | 3.263%
| PDM3 |  | -1.011% | -8.449% | -6.542% | -1.011% | 16.06% | -12.26% | -13.22% | -11.31% | 0.6103% | -7.4% | -6.924% | 9.765% | 8.621% | 1.278% | 0.1335% | -3.109% | 1.087% | 9.393% | -3.443% | 14.44% | -1.011% | 4.997% | -0.534%
| Z-Score (Absolute) |  | 0.11 | NA | NA | 0.11 | 3.85 | NA | NA | 2.5 | 0.27 | 1.59 | 1.48 | 2.39 | 2.12 | 0.42 | NA | 0.6 | 0.38 | 2.3 | 0.67 | NA | 0.11 | 1.28 | 0
Results from laboratories 2, 3, 7, 8, 19 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 282 (Indicative Value < 10 ppb) | [26 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.13 | 2.15 | [18 empty] | ppb
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [18 empty] | 1
|  |  | 04g | 25g | 40g | 0.5g | 25g | 01g | 01g | [18 empty] | 1
| 1 | 1 | < 5 | 550 | < 10 | < 1 | < 5 | < 10 | < 10 | [18 empty] | 1
| 1 | 2 | < 5 | 500 | < 10 | < 1 | < 5 | < 10 | < 10 | [18 empty] | 1
| 1 | 3 | < 5 | 550 | < 10 | 1 | < 5 | < 10 | < 10 | [18 empty] | 16
| 1 | 4 | < 5 | 550 | < 10 | < 1 | < 5 | < 10 | < 10 | [18 empty] | < 10
| 1 | 5 | < 5 | 500 | < 10 | 3 | < 5 | < 10 | < 10 | [18 empty] | 13
| 1 | 6 | < 5 | 550 | < 10 | 2 | < 5 | < 10 | < 10
| Mean |  |  | 533.3333 |  | 2
| Median |  |  | 550 |  | 2
| Std Dev. |  |  | 25.8199 |  | 1
| Rel.Std.Dev. |  |  | 4.841% |  | 50%
PDM3
| Z-Score (Absolute) |  | 0.67 | 142.5 | 0 | 1.01 | 0.67 | 0 | 0
| Analytical results for Pr in OREAS 282 (Certified Value 4.21 ppm) | [26 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | [17 empty] | 2
| 1 | 1 | 4.71 | 3.7 | 3.51 | 4.594 | 4.44 | 4.15 | 4.9 | 4 | [17 empty] | 1
| 1 | 2 | 4.6 | 4.1 | 3.43 | 4.563 | 4.21 | 4.19 | 4.9 | 3.84 | [17 empty] | 30
| 1 | 3 | 4.61 | 3.6 | 3.7 | 4.458 | 4.31 | 4.2 | 4.9 | 3.78 | [17 empty] | 16
| 1 | 4 | 4.7 | 3.8 | 3.55 | 4.519 | 4.29 | 4.24 | 4.5 | 3.77 | [17 empty] | 4.2067
| 1 | 5 | 4.7 | 3.7 | 3.6 | 4.569 | 4.31 | 4.11 | 4.7 | 4.04 | [17 empty] | 107
| 1 | 6 | 4.63 | 3.7 | 3.52 | 4.611 | 4.21 | 4.21 | 5.1 | 3.85
| Mean |  | 4.6583 | 3.7667 | 3.5517 | 4.5523 | 4.295 | 4.1833 | 4.8333 | 3.88
| Median |  | 4.665 | 3.7 | 3.535 | 4.566 | 4.3 | 4.195 | 4.9 | 3.845
| Std Dev. |  | 0.050365 | 0.17512 | 0.091524 | 0.05582 | 0.084794 | 0.046332 | 0.20656 | 0.11367
| Rel.Std.Dev. |  | 1.081% | 4.649% | 2.577% | 1.226% | 1.974% | 1.108% | 4.274% | 2.93%
| PDM3 |  | 10.73% | -10.46% | -15.57% | 8.215% | 2.098% | -0.5566% | 14.89% | -7.767%
| Z-Score (Absolute) |  | 0.73 | 0.82 | 1.19 | 0.54 | 0.1 | 0.1 | 1.03 | 0.62
| Analytical results for Pt in OREAS 282 (Indicative Value < 5 ppb) | [26 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.13 | 2.15 | [18 empty] | ppb
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [18 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 01g | 01g | [18 empty] | 2
| 1 | 1 | 1 | 10 | < 10 | < 1 | < 5 | < 5 | < 5 | [18 empty] | 1
| 1 | 2 | < 1 | 10 | < 10 | < 1 | < 5 | < 5 | < 5 | [18 empty] | 1
| 1 | 3 | < 1 | < 10 | < 10 | < 1 | < 5 | < 5 | < 5 | [18 empty] | 16
| 1 | 4 | < 1 | 10 | < 10 | < 1 | < 5 | < 5 | < 5 | [18 empty] | < 5
| 1 | 5 | < 1 | 10 | < 10 | 1 | < 5 | < 5 | < 5 | [18 empty] | 14
| 1 | 6 | < 1 | 10 | < 10 | < 1 | < 5 | < 5 | < 5
| Mean |  | 1 | 10 |  | 1
| Median |  | 1 | 10 |  | 1
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 2.35 | 0.88 | 0.67 | 0 | 0 | 0
| Analytical results for Rb in OREAS 282 (Certified Value 39 ppm) | [26 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.29 | 2.30 | 2.31 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.5g | 04g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 39.3 | 34.6 | 37.1 | 33 | 41.8 | 36 | 38.58 | 42.03 | 47.77 | 47.82 | 52.07 | 37.3 | 40.6 | 36.2 | 36.6 | 51.4 | 39.8 | 40.5 | 38 | [6 empty] | 1
| 1 | 2 | 39.8 | 34.2 | 38.2 | 38.5 | 41.8 | 36 | 38.71 | 40.43 | 46.22 | 48.83 | 54.01 | 36 | 41 | 38.3 | 36.3 | 49.7 | 36.9 | 41.7 | 35.5 | [6 empty] | 31
| 1 | 3 | 38.6 | 34.2 | 36.8 | 30 | 42.3 | 34.8 | 37.37 | 39.16 | 47.67 | 47.51 | 51.93 | 37.5 | 41.2 | 37.6 | 35.9 | 49.9 | 37.3 | 42.5 | 36.1 | [6 empty] | 16
| 1 | 4 | 39.9 | 35.2 | 38.2 | 35 | 40.5 | 36.4 | 37.89 | 39.76 | 44.84 | 47.99 | 51.89 | 35.3 | 41.5 | 35.2 | 35.7 | 49.3 | 37.8 | 41.7 | 35.4 | [6 empty] | 38.9758
| 1 | 5 | 39.5 | 35.2 | 37.5 | 35 | 43 | 36.1 | 38.43 | 42.61 | 45.02 | 47.21 | 50.91 | 36 | 42 | 38.5 | 36.6 | 49.4 | 35 | 41.9 | 36.5 | [6 empty] | 108
| 1 | 6 | 39.2 | 34.4 | 36.9 | 34.5 | 40.3 | 36 | 35.78 | 41.28 | 44.52 | 47.89 | 51.73 | 36.7 | 40.7 | 36.5 | 35.8 | 50.4 | 36.7 | 40.3 | 38.1
| Mean |  | 39.3833 | 34.6333 | 37.45 | 34.3333 | 41.6167 | 35.8833 | 37.7933 | 40.8783 | 46.0067 | 47.875 | 52.09 | 36.4667 | 41.1667 | 37.05 | 36.15 | 50.0167 | 37.25 | 41.4333 | 36.6
| Median |  | 39.4 | 34.5 | 37.3 | 34.75 | 41.8 | 36 | 38.16 | 40.855 | 45.62 | 47.855 | 51.91 | 36.35 | 41.1 | 37.05 | 36.1 | 49.8 | 37.1 | 41.7 | 36.3
| Std Dev. |  | 0.47081 | 0.46332 | 0.62849 | 2.7869 | 1.042 | 0.55287 | 1.1051 | 1.3335 | 1.4468 | 0.54844 | 1.0271 | 0.8501 | 0.52409 | 1.2973 | 0.40373 | 0.78337 | 1.5681 | 0.85479 | 1.1933
| Rel.Std.Dev. |  | 1.195% | 1.338% | 1.678% | 8.117% | 2.504% | 1.541% | 2.924% | 3.262% | 3.145% | 1.146% | 1.972% | 2.331% | 1.273% | 3.501% | 1.117% | 1.566% | 4.21% | 2.063% | 3.26%
| PDM3 |  | 1.045% | -11.14% | -3.915% | -11.91% | 6.776% | -7.934% | -3.034% | 4.881% | 18.04% | 22.83% | 33.65% | -6.438% | 5.621% | -4.941% | -7.25% | 28.33% | -4.428% | 6.305% | -6.096%
| Z-Score (Absolute) |  | 0.35 | 0.69 | 0.08 | 0.76 | 0.84 | 0.42 | 0 | 0.67 | 1.8 | 2.2 | 3.12 | 0.29 | 0.74 | 0.16 | 0.36 | 2.67 | 0.12 | 0.8 | 0.26
| Analytical results for Re in OREAS 282 (Certified Value 0.024 ppm) | [26 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.21 | 2.29 | 2.30 | 2.31 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.25g | 25g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.4g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.023 | 0.021 | < 0.05 | 0.027 | 0.022 | 0.022 | < 0.05 | 0.025 | 0.021 | 0.023 | < 0.05 | 0.022 | < 0.05 | 0.026 | 0.026 | 0.024 | [9 empty] | 1
| 1 | 2 | 0.024 | 0.021 | < 0.05 | 0.027 | 0.023 | 0.022 | < 0.05 | 0.026 | 0.022 | 0.023 | < 0.05 | 0.023 | < 0.05 | 0.024 | 0.022 | 0.021 | [9 empty] | 32
| 1 | 3 | 0.022 | 0.022 | < 0.05 | 0.026 | 0.025 | 0.021 | < 0.05 | 0.026 | 0.022 | 0.024 | < 0.05 | 0.022 | < 0.05 | 0.025 | 0.027 | 0.021 | [9 empty] | 16
| 1 | 4 | 0.025 | 0.02 | < 0.05 | 0.028 | 0.023 | 0.025 | < 0.05 | 0.029 | 0.022 | 0.022 | < 0.05 | 0.022 | < 0.05 | 0.025 | 0.025 | 0.022 | [9 empty] | 0.023547
| 1 | 5 | 0.023 | 0.022 | < 0.05 | 0.029 | 0.025 | 0.021 | < 0.05 | 0.023 | 0.023 | 0.022 | < 0.05 | 0.021 | < 0.05 | 0.021 | 0.025 | 0.024 | [9 empty] | 109
| 1 | 6 | 0.026 | 0.022 | < 0.05 | 0.029 | 0.023 | 0.021 | < 0.05 | 0.025 | 0.021 | 0.022 | < 0.05 | 0.023 | < 0.05 | 0.028 | 0.027 | 0.022
| Mean |  | 0.023833 | 0.021333 |  | 0.027667 | 0.0235 | 0.022 |  | 0.025667 | 0.021833 | 0.022667 |  | 0.022167 |  | 0.024833 | 0.025333 | 0.022333
| Median |  | 0.0235 | 0.0215 |  | 0.0275 | 0.023 | 0.0215 |  | 0.0255 | 0.022 | 0.0225 |  | 0.022 |  | 0.025 | 0.0255 | 0.022
| Std Dev. |  | 0.001472 | 0.0008165 |  | 0.0012111 | 0.0012247 | 0.0015492 |  | 0.0019664 | 0.00075277 | 0.0008165 |  | 0.00075277 |  | 0.0023166 | 0.0018619 | 0.0013663
| Rel.Std.Dev. |  | 6.176% | 3.827% |  | 4.377% | 5.212% | 7.042% |  | 7.661% | 3.448% | 3.602% |  | 3.396% |  | 9.329% | 7.35% | 6.118%
| PDM3 |  | 1.215% | -9.402% |  | 17.49% | -0.2005% | -6.571% |  | 9.001% | -7.279% | -3.74% |  | -5.863% |  | 5.462% | 7.585% | -5.155%
| Z-Score (Absolute) |  | 0.29 | 1.73 | 0.39 | 1.93 | 0.48 | 1.35 | 0.39 | 0.77 | 1.44 | 0.96 | 0.39 | 1.25 | 0.39 | 0.29 | 0.58 | 1.16
| Analytical results for S in OREAS 282 (Certified Value 0.708 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.75 | 0.745 | 0.69 | 0.77 | 0.91 | 0.665 | 0.725 | 0.91 | 0.71 | 0.67 | 0.6692 | 0.632 | 0.78 | 0.835 | 0.68 | 0.67 | 0.61 | 0.65 | 0.76 | 0.78 | 0.74 | 0.67 |  |  |  | 1
| 1 | 2 | 0.74 | 0.73 | 0.71 | 0.8 | 0.9 | 0.6775 | 0.75 | 0.91 | 0.69 | 0.64 | 0.6897 | 0.616 | 0.75 | 0.84 | 0.69 | 0.66 | 0.61 | 0.66 | 0.75 | 0.75 | 0.73 | 0.67 |  |  |  | 17
| 1 | 3 | 0.7 | 0.735 | 0.69 | 0.78 | 0.93 | 0.673 | 0.74 | 0.89 | 0.7 | 0.66 | 0.684 | 0.651 | 0.75 | 0.814 | 0.68 | 0.67 | 0.61 | 0.65 | 0.76 | 0.74 | 0.73 | 0.68 |  |  |  | 16
| 1 | 4 | 0.74 | 0.74 | 0.69 | 0.75 | 0.92 | 0.6785 | 0.74 | 0.92 | 0.7 | 0.69 | 0.6971 | 0.631 | 0.76 | 0.855 | 0.7 | 0.67 | 0.6 | 0.65 | 0.77 | 0.75 | 0.72 | 0.68 |  |  |  | 0.70784
| 1 | 5 | 0.72 | 0.75 | 0.68 | 0.77 | 0.91 | 0.6555 | 0.74 | 0.92 | 0.72 | 0.66 | 0.68 | 0.631 | 0.76 | 0.862 | 0.69 | 0.67 | 0.61 | 0.65 | 0.75 | 0.73 | 0.73 | 0.66 |  |  |  | 110
| 1 | 6 | 0.7 | 0.72 | 0.7 | 0.8 | 0.92 | 0.681 | 0.735 | 0.9 | 0.7 | 0.69 | 0.6798 | 0.613 | 0.76 | 0.861 | 0.69 | 0.67 | 0.61 | 0.66 | 0.77 | 0.74 | 0.72 | 0.66
| Mean |  | 0.725 | 0.73667 | 0.69333 | 0.77833 | 0.915 | 0.67175 | 0.73833 | 0.90833 | 0.70333 | 0.66833 | 0.6833 | 0.629 | 0.76 | 0.8445 | 0.68833 | 0.66833 | 0.60833 | 0.65333 | 0.76 | 0.74833 | 0.72833 | 0.67
| Median |  | 0.73 | 0.7375 | 0.69 | 0.775 | 0.915 | 0.67525 | 0.74 | 0.91 | 0.7 | 0.665 | 0.682 | 0.631 | 0.76 | 0.8475 | 0.69 | 0.67 | 0.61 | 0.65 | 0.76 | 0.745 | 0.73 | 0.67
| Std Dev. |  | 0.021679 | 0.010801 | 0.010328 | 0.019408 | 0.010488 | 0.0097506 | 0.008165 | 0.01169 | 0.010328 | 0.019408 | 0.0095272 | 0.013609 | 0.010954 | 0.018598 | 0.0075277 | 0.0040825 | 0.0040825 | 0.005164 | 0.0089443 | 0.017224 | 0.0075277 | 0.0089443
| Rel.Std.Dev. |  | 2.99% | 1.466% | 1.49% | 2.494% | 1.146% | 1.452% | 1.106% | 1.287% | 1.468% | 2.904% | 1.394% | 2.164% | 1.441% | 2.202% | 1.094% | 0.6108% | 0.6711% | 0.7904% | 1.177% | 2.302% | 1.034% | 1.335%
| PDM3 |  | 2.424% | 4.072% | -2.05% | 9.958% | 29.27% | -5.099% | 4.307% | 28.32% | -0.6373% | -5.582% | -3.467% | -11.14% | 7.368% | 19.31% | -2.756% | -5.582% | -14.06% | -7.701% | 7.368% | 5.72% | 2.895% | -5.346%
| Z-Score (Absolute) |  | 0.16 | 0.34 | 0.31 | 0.96 | 3.01 | 0.64 | 0.36 | 2.91 | 0.16 | 0.69 | 0.46 | 1.28 | 0.69 | 1.95 | 0.39 | 0.69 | 1.59 | 0.91 | 0.69 | 0.51 | 0.21 | 0.66
| Analytical results for Sb in OREAS 282 (Certified Value 10.5 ppm) | [26 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 10.85 | 10.2 | 11 | 10 | 8.5 | 6.8 | 9.6 | 11.03 | 9.43 | 10.91 | 7.11 | 9.6 | 12.1 | 10.32 | 13.3 | 8.38 | 11 | 11.57 | 10.96 | < 5 | 12.7 | 10.87 | 11.25 |  |  | 1
| 1 | 2 | 10.8 | 10.1 | 11 | 10 | 9.5 | 6.8 | 9.96 | 10.98 | 9.48 | 10.5 | 7.17 | 9.49 | 11.3 | 10.34 | 12.8 | 7.92 | 10.3 | 12.05 | 10.96 | < 5 | 12 | 10.33 | 10.6 |  |  | 33
| 1 | 3 | 10.7 | 10.1 | 10 | 11 | 8.5 | 7.6 | 9.71 | 10.6 | 8.83 | 11.03 | 7.06 | 9.41 | 12.2 | 10.48 | 12.4 | 8.33 | 10.8 | 11.59 | 10.77 | < 5 | 11.95 | 10.37 | 11.05 |  |  | 16
| 1 | 4 | 11 | 10.3 | 10 | 9 | 8.5 | 7.1 | 9.91 | 10.85 | 9.02 | 10.36 | 7.12 | 9.87 | 12.2 | 10.28 | 12.9 | 7.86 | 10 | 11.6 | 10.96 | < 5 | 12.15 | 10.75 | 10.3 |  |  | 10.4941
| 1 | 5 | 10.8 | 10.2 | 10 | 10 | 8 | 7.1 | 9.69 | 10.98 | 9.61 | 10.61 | 7.05 | 9.75 | 11.4 | 10.59 | 13.2 | 8.04 | 10.7 | 11.73 | 10.28 | < 5 | 11.55 | 10.6 | 10.95 |  |  | 111
| 1 | 6 | 10.55 | 10.2 | 11 | 12 | 8.5 | 8.2 | 10 | 10.44 | 9.54 | 10.25 | 7.02 | 9.49 | 12.3 | 10.34 | 12.8 | 7.72 | 10.2 | 11.58 | 10.4 | < 5 | 12 | 10.38 | 11.65
| Mean |  | 10.7833 | 10.1833 | 10.5 | 10.3333 | 8.5833 | 7.2667 | 9.8117 | 10.8133 | 9.3183 | 10.61 | 7.0883 | 9.6017 | 11.9167 | 10.3917 | 12.9 | 8.0417 | 10.5 | 11.6867 | 10.7217 |  | 12.0583 | 10.55 | 10.9667
| Median |  | 10.8 | 10.2 | 10.5 | 10 | 8.5 | 7.1 | 9.81 | 10.915 | 9.455 | 10.555 | 7.085 | 9.545 | 12.15 | 10.34 | 12.85 | 7.98 | 10.5 | 11.595 | 10.865 |  | 12 | 10.49 | 11
| Std Dev. |  | 0.15055 | 0.075277 | 0.54772 | 1.0328 | 0.4916 | 0.54283 | 0.16558 | 0.24031 | 0.31632 | 0.30679 | 0.054924 | 0.17646 | 0.4446 | 0.11839 | 0.32249 | 0.26415 | 0.38987 | 0.18747 | 0.30701 |  | 0.37339 | 0.22565 | 0.4761
| Rel.Std.Dev. |  | 1.396% | 0.7392% | 5.216% | 9.995% | 5.727% | 7.47% | 1.688% | 2.222% | 3.395% | 2.892% | 0.7749% | 1.838% | 3.731% | 1.139% | 2.5% | 3.285% | 3.713% | 1.604% | 2.863% |  | 3.096% | 2.139% | 4.341%
| PDM3 |  | 2.756% | -2.961% | 0.05651% | -1.532% | -18.21% | -30.75% | -6.503% | 3.042% | -11.2% | 1.105% | -32.45% | -8.504% | 13.56% | -0.9758% | 22.93% | -23.37% | 0.05651% | 11.36% | 2.169% |  | 14.91% | 0.533% | 4.503%
| Z-Score (Absolute) |  | 0.31 | 0.24 | NA | 0.1 | 1.7 | 2.9 | 0.58 | 0.34 | 1.03 | 0.15 | 3.06 | 0.77 | 1.34 | 0.05 | 2.24 | 2.19 | 0.05 | 1.13 | 0.25 | 7.25 | 1.47 | 0.1 | 0.48
Results from laboratory 3 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 282 (Certified Value 3.52 ppm) | [26 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 3
| 1 | 1 | 3.8 | 4 | < 5 | 3.4 | 3.2 | 4 | 3 | 3.4 | 3.55 | 3.2 | 3 | 3.21 | 3.6 | 3.8 | 3.6 | 4.5 | 3.6 | 3 | 3.9 | 2.9 | < 5 | 3.8 | 3.4 | 3.4 |  | 1
| 1 | 2 | 3.8 | 4 | < 5 | 3.7 | 3.1 | 4 | 3 | 3.4 | 3.6 | 3.4 | 3 | 3.21 | 3.6 | 3.7 | 3.7 | 4.3 | 3.6 | 3 | 4 | 2.9 | < 5 | 3.5 | 3.2 | 3.1 |  | 34
| 1 | 3 | 3.7 | 3 | < 5 | 3.4 | 3.2 | 4 | 3 | 3.3 | 3.47 | 3.2 | 3 | 3.19 | 3.6 | 4 | 3.7 | 4.3 | 3.5 | 3 | 3.9 | 2.9 | < 5 | 3.6 | 3.4 | 3.2 |  | 16
| 1 | 4 | 3.8 | 4 | < 5 | 3.9 | 3.3 | 4 | 3 | 3.4 | 3.51 | 3.4 | 3 | 3.17 | 3.6 | 3.6 | 3.6 | 4.1 | 3.6 | 3 | 4.1 | 2.9 | < 5 | 3.7 | 3.4 | 3 |  | 3.523
| 1 | 5 | 3.7 | 4 | < 5 | 3.6 | 3.2 | 4 | 3 | 3.4 | 3.58 | 3.3 | 3 | 3.18 | 3.7 | 3.8 | 3.7 | 4.6 | 3.6 | 3 | 3.7 | 2.9 | < 5 | 3.5 | 3.4 | 3.2 |  | 112
| 1 | 6 | 3.6 | 4 | < 5 | 3.5 | 3.2 | 4 | 3 | 3.3 | 3.48 | 3.3 | 3 | 3.22 | 3.5 | 4 | 3.6 | 4.2 | 3.6 | 3 | 4 | 3 | < 5 | 3.7 | 3.3 | 3.4
| Mean |  | 3.7333 | 3.8333 |  | 3.5833 | 3.2 | 4 | 3 | 3.3667 | 3.5317 | 3.3 | 3 | 3.1967 | 3.6 | 3.8167 | 3.65 | 4.3333 | 3.5833 | 3 | 3.9333 | 2.9167 |  | 3.6333 | 3.35 | 3.2167
| Median |  | 3.75 | 4 |  | 3.55 | 3.2 | 4 | 3 | 3.4 | 3.53 | 3.3 | 3 | 3.2 | 3.6 | 3.8 | 3.65 | 4.3 | 3.6 | 3 | 3.95 | 2.9 |  | 3.65 | 3.4 | 3.2
| Std Dev. |  | 0.08165 | 0.40825 |  | 0.19408 | 0.063246 | 0 | 0 | 0.05164 | 0.053448 | 0.089443 | 0 | 0.019664 | 0.063246 | 0.16021 | 0.054772 | 0.18619 | 0.040825 | 0 | 0.13663 | 0.040825 |  | 0.12111 | 0.083666 | 0.16021
| Rel.Std.Dev. |  | 2.187% | 10.65% |  | 5.416% | 1.976% | 0% | 0% | 1.534% | 1.513% | 2.71% | 0% | 0.6151% | 1.757% | 4.198% | 1.501% | 4.297% | 1.139% | 0% | 3.474% | 1.4% |  | 3.333% | 2.497% | 4.981%
| PDM3 |  | 5.969% | 8.808% |  | 1.711% | -9.169% | 13.54% | -14.85% | -4.439% | 0.2449% | -6.331% | -14.85% | -9.264% | 2.184% | 8.334% | 3.604% | 23% | 1.711% | -14.85% | 11.65% | -17.21% |  | 3.131% | -4.912% | -8.696%
| Z-Score (Absolute) |  | 0.59 | NA | 3 | 0.15 | 0.97 | NA | NA | 0.48 | 0 | 0.67 | NA | 0.98 | 0.2 | 0.83 | 0.34 | 2.33 | 0.15 | NA | 1.17 | 1.79 | 3 | 0.3 | 0.53 | 0.92
Results from laboratories 2, 7, 8, 13 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 282 (Certified Value 2.35 ppm) | [26 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 2.7 | 3 | < 5 | 3 | 3.3 | 2 | 2.5 | 2.7 | 2.47 | 1 | 2 | 2.25 | 2 | 2.6 | 3 | 2 | 2.2 | 2 | 2.1 | 2.6 | < 10 | 2.4 | 2.1 | 1.7 |  | 1
| 1 | 2 | 2.5 | 3 | < 5 | 3 | 3.5 | 2 | 2.5 | 2 | 2.39 | 2.4 | 2 | 2.19 | 2 | 2.7 | 3 | 2 | 2.4 | 2 | 2.1 | 2.8 | < 10 | 2.5 | 1.8 | 2 |  | 35
| 1 | 3 | 2.5 | 3 | < 5 | 3 | 3.4 | 2 | 2.5 | 1.7 | 2.27 | 1 | 2 | 2.28 | 2 | 2.9 | 3 | 3 | 2.5 | 2 | 2.2 | 2.7 | < 10 | 1.8 | 2.1 | 2 |  | 16
| 1 | 4 | 2.5 | 3 | < 5 | 3 | 3.3 | 2 | 2.5 | 3.2 | 2.46 | 3 | 2 | 2.18 | 2 | 2.5 | 3 | 2 | 2.6 | 2 | 2.1 | 2.2 | < 10 | 1.8 | 2 | 1.8 |  | 2.3454
| 1 | 5 | 2.7 | 3 | < 5 | 3 | 3.5 | 2 | 2.5 | 2.3 | 2.58 | < 0.5 | 2 | 2.21 | 2 | 2.6 | 3 | 3 | 2.2 | 2 | 2.2 | 2.4 | < 10 | 2 | 2.1 | 2.2 |  | 113
| 1 | 6 | 2.5 | 3 | < 5 | 4 | 3.1 | 2 | 2.5 | 3 | 2.31 | 0.7 | 2 | 2.22 | 2 | 2.8 | 3 | 2 | 2.5 | 2 | 2.3 | 2.3 | < 10 | 2 | 1.9 | 2.1
| Mean |  | 2.5667 | 3 |  | 3.1667 | 3.35 | 2 | 2.5 | 2.4833 | 2.4133 | 1.62 | 2 | 2.2217 | 2 | 2.6833 | 3 | 2.3333 | 2.4 | 2 | 2.1667 | 2.5 |  | 2.0833 | 2 | 1.9667
| Median |  | 2.5 | 3 |  | 3 | 3.35 | 2 | 2.5 | 2.5 | 2.425 | 1 | 2 | 2.215 | 2 | 2.65 | 3 | 2 | 2.45 | 2 | 2.15 | 2.5 |  | 2 | 2.05 | 2
| Std Dev. |  | 0.10328 | 0 |  | 0.40825 | 0.15166 | 0 | 0 | 0.58452 | 0.11396 | 1.0159 | 0 | 0.037639 | 0 | 0.1472 | 0 | 0.5164 | 0.16733 | 0 | 0.08165 | 0.23664 |  | 0.29944 | 0.12649 | 0.18619
| Rel.Std.Dev. |  | 4.024% | 0% |  | 12.89% | 4.527% | 0% | 0% | 23.54% | 4.722% | 62.71% | 0% | 1.694% | 0% | 5.486% | 0% | 22.13% | 6.972% | 0% | 3.768% | 9.466% |  | 14.37% | 6.325% | 9.467%
| PDM3 |  | 9.436% | 27.91% |  | 35.02% | 42.84% | -14.73% | 6.594% | 5.883% | 2.898% | -30.93% | -14.73% | -5.274% | -14.73% | 14.41% | 27.91% | -0.5127% | 2.33% | -14.73% | -7.619% | 6.594% |  | -11.17% | -14.73% | -16.15%
| Z-Score (Absolute) |  | 0.21 | NA | 0.04 | 1.76 | 2.23 | NA | 0.04 | 0 | 0.18 | 2.81 | NA | 0.67 | NA | 0.52 | NA | NA | 0.21 | NA | 0.82 | 0.04 | 6.49 | 1.03 | 1.25 | 1.33
Results from laboratories 2, 7, 13, 15, 18, 19 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Si in OREAS 282 (Indicative Value 0.014 wt.%) | [26 empty] | Indicative
Si | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.07 | [24 empty] | wt.%
|  |  | AR*MS | [24 empty] | 3
|  |  | 25g | [24 empty] | 3
| 1 | 1 | 0.0135 | [24 empty] | 1
| 1 | 2 | 0.0135 | [24 empty] | 7
| 1 | 3 | 0.019 | [24 empty] | 16
| 1 | 4 | 0.014 | [24 empty] | 0.014333
| 1 | 5 | 0.0135 | [24 empty] | 13
| 1 | 6 | 0.0125
| Mean |  | 0.014333
| Median |  | 0.0135
| Std Dev. |  | 0.0023381
| Rel.Std.Dev. |  | 16.31%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 282 (Certified Value 3.13 ppm) | [26 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | [17 empty] | 3
| 1 | 1 | 3.3 | 2.8 | 2.75 | 3.136 | 3.37 | 3.18 | 3.9 | 2.83 | [17 empty] | 1
| 1 | 2 | 3.3 | 3.1 | 2.71 | 3.139 | 3.28 | 3.27 | 3.4 | 2.82 | [17 empty] | 19
| 1 | 3 | 3.29 | 2.8 | 2.9 | 3.089 | 3.29 | 3.24 | 3.8 | 2.83 | [17 empty] | 16
| 1 | 4 | 3.33 | 2.9 | 2.83 | 3.062 | 3.25 | 3.21 | 3.7 | 2.89 | [17 empty] | 3.1308
| 1 | 5 | 3.31 | 2.8 | 2.93 | 3.111 | 3.21 | 3.14 | 3.5 | 2.97 | [17 empty] | 114
| 1 | 6 | 3.38 | 2.8 | 2.78 | 3.151 | 3.23 | 3.24 | 3.5 | 2.81
| Mean |  | 3.3183 | 2.8667 | 2.8167 | 3.1147 | 3.2717 | 3.2133 | 3.6333 | 2.8583
| Median |  | 3.305 | 2.8 | 2.805 | 3.1235 | 3.265 | 3.225 | 3.6 | 2.83
| Std Dev. |  | 0.033116 | 0.12111 | 0.086178 | 0.034098 | 0.056716 | 0.047188 | 0.19664 | 0.061455
| Rel.Std.Dev. |  | 0.998% | 4.225% | 3.06% | 1.095% | 1.734% | 1.468% | 5.412% | 2.15%
| PDM3 |  | 5.99% | -8.436% | -10.03% | -0.515% | 4.5% | 2.636% | 16.05% | -8.703%
| Z-Score (Absolute) |  | 0.46 | 0.89 | 1.04 | 0.15 | 0.32 | 0.15 | 1.4 | 0.91
| Analytical results for Sn in OREAS 282 (Certified Value 1.02 ppm) | [26 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1 | 1 | < 10 | 0.9 | 1.3 | 1 | 1 | 0.8 | 1.05 | 1.1 | 0.8 | 1.09 | 1.17 | 1.18 | 1 | 1 | 1 | 1.1 | < 20 | 1.1 | 1 | 1 |  |  |  | 1
| 1 | 2 | 1 | 1 | < 10 | 0.9 | 1.2 | 1 | 1 | 0.9 | 1.07 | 1 | 0.7 | 1.1 | 1.18 | 1.14 | 1 | 1 | 1 | 1.1 | < 20 | 1 | 1 | 0.9 |  |  |  | 20
| 1 | 3 | 1 | 1 | < 10 | 1 | 1.5 | 1 | 1 | 0.8 | 1.04 | 1 | 0.8 | 1.1 | 1.2 | 1.22 | 1 | 0.9 | 1 | 1.1 | < 20 | 1 | 1.1 | 1 |  |  |  | 16
| 1 | 4 | 1 | 1 | < 10 | 0.9 | 1.4 | 1 | 1 | 0.9 | 1.04 | 1 | 0.7 | 1.09 | 1.17 | 1.13 | 1 | 1 | 1 | 1 | < 20 | 1 | 1 | 0.9 |  |  |  | 1.0244
| 1 | 5 | 1 | 1 | < 10 | 1 | 1.3 | 1 | 1 | 0.8 | 1.05 | 1 | 0.7 | 1.02 | 1.13 | 1.15 | 1 | 1 | 1 | 1.1 | < 20 | 1 | 1 | 1 |  |  |  | 115
| 1 | 6 | 1 | 1 | < 10 | 1 | 1.3 | 1 | 1 | 0.8 | 1.07 | 1 | 0.7 | 1.05 | 1.14 | 1.2 | 0.9 | 1 | 0.9 | 1.1 | < 20 | 1.1 | 1.1 | 1
| Mean |  | 1 | 1 |  | 0.95 | 1.3333 | 1 | 1 | 0.83333 | 1.0533 | 1.0167 | 0.73333 | 1.075 | 1.165 | 1.17 | 0.98333 | 0.98333 | 0.98333 | 1.0833 |  | 1.0333 | 1.0333 | 0.96667
| Median |  | 1 | 1 |  | 0.95 | 1.3 | 1 | 1 | 0.8 | 1.05 | 1 | 0.7 | 1.09 | 1.17 | 1.165 | 1 | 1 | 1 | 1.1 |  | 1 | 1 | 1
| Std Dev. |  | 0 | 0 |  | 0.054772 | 0.10328 | 0 | 0 | 0.05164 | 0.013663 | 0.040825 | 0.05164 | 0.032711 | 0.025884 | 0.035777 | 0.040825 | 0.040825 | 0.040825 | 0.040825 |  | 0.05164 | 0.05164 | 0.05164
| Rel.Std.Dev. |  | 0% | 0% |  | 5.766% | 7.746% | 0% | 0% | 6.197% | 1.297% | 4.016% | 7.042% | 3.043% | 2.222% | 3.058% | 4.152% | 4.152% | 4.152% | 3.768% |  | 4.997% | 4.997% | 5.342%
| PDM3 |  | -2.378% | -2.378% |  | -7.259% | 30.16% | -2.378% | -2.378% | -18.65% | 2.829% | -0.7508% | -28.41% | 4.944% | 13.73% | 14.22% | -4.005% | -4.005% | -4.005% | 5.757% |  | 0.8763% | 0.8763% | -5.632%
| Z-Score (Absolute) |  | NA | NA | 45.85 | 0.96 | 3.47 | NA | NA | 2.31 | 0.23 | 0.19 | 3.47 | 0.48 | 1.52 | 1.58 | 0.58 | 0.58 | 0.58 | 0.58 | 103.65 | 0 | 0 | 0.77
Results from laboratories 1, 2, 7 and 8 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 282 (Certified Value 218 ppm) | [26 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 0
| 1 | 1 | 232 | 240 | 164 | 210.8 | 214 | 189 | 210 | 173 | 225.86 | 213.5 | 165.9 | 226 | 239.03 | 232 | 213 | 221 | 228 | 221 | 213.8 | 185.2 | 239 | 248 | 216.9 | 224 |  | 1
| 1 | 2 | 229 | 232 | 164 | 220 | 211 | 192 | 212 | 174 | 222.99 | 210.6 | 164.3 | 223 | 240.61 | 235 | 213 | 188 | 230 | 221 | 219.4 | 187.2 | 237 | 237 | 213.7 | 222 |  | 21
| 1 | 3 | 227 | 236 | 160 | 213.8 | 212 | 190 | 217 | 169 | 218.33 | 209.7 | 165.7 | 224 | 235.63 | 239 | 214 | 214 | 227 | 224 | 211.9 | 186.2 | 239 | 236 | 215.3 | 225 |  | 16
| 1 | 4 | 230 | 239 | 163 | 221.4 | 217 | 191 | 223 | 178 | 222.28 | 214.5 | 161.9 | 220 | 243.02 | 222 | 209 | 181 | 231 | 222 | 210 | 185.2 | 240 | 241 | 213.6 | 224 |  | 217.5887
| 1 | 5 | 233 | 243 | 159 | 217.6 | 211 | 190 | 216 | 175 | 224.22 | 214.7 | 164.2 | 221 | 234.22 | 233 | 214 | 214 | 233 | 224 | 208.9 | 185.2 | 240 | 231 | 214 | 224 |  | 116
| 1 | 6 | 228 | 237 | 164 | 213.7 | 213 | 191 | 204 | 171 | 221.12 | 213.1 | 161.2 | 221 | 237.4 | 242 | 214 | 205 | 229 | 228 | 205.3 | 227 | 242 | 237 | 213 | 222
| Mean |  | 229.8333 | 237.8333 | 162.3333 | 216.2167 | 213 | 190.5 | 213.6667 | 173.3333 | 222.4667 | 212.6833 | 163.8667 | 222.5 | 238.3183 | 233.8333 | 212.8333 | 203.8333 | 229.6667 | 223.3333 | 211.55 | 192.6667 | 239.5 | 238.3333 | 214.4167 | 223.5
| Median |  | 229.5 | 238 | 163.5 | 215.7 | 212.5 | 190.5 | 214 | 173.5 | 222.635 | 213.3 | 164.25 | 222 | 238.215 | 234 | 213.5 | 209.5 | 229.5 | 223 | 210.95 | 185.7 | 239.5 | 237 | 213.85 | 224
| Std Dev. |  | 2.3166 | 3.7639 | 2.2509 | 4.1136 | 2.2804 | 1.0488 | 6.532 | 3.1411 | 2.6009 | 2.0712 | 1.9377 | 2.2583 | 3.2475 | 6.9113 | 1.9408 | 15.9677 | 2.1602 | 2.6583 | 4.802 | 16.8388 | 1.6432 | 5.7155 | 1.4359 | 1.2247
| Rel.Std.Dev. |  | 1.008% | 1.583% | 1.387% | 1.903% | 1.071% | 0.5506% | 3.057% | 1.812% | 1.169% | 0.9738% | 1.182% | 1.015% | 1.363% | 2.956% | 0.9119% | 7.834% | 0.9406% | 1.19% | 2.27% | 8.74% | 0.6861% | 2.398% | 0.6697% | 0.548%
| PDM3 |  | 5.627% | 9.304% | -25.39% | -0.6306% | -2.109% | -12.45% | -1.803% | -20.34% | 2.242% | -2.254% | -24.69% | 2.257% | 9.527% | 7.466% | -2.185% | -6.322% | 5.551% | 2.64% | -2.775% | -11.45% | 10.07% | 9.534% | -1.458% | 2.717%
| Z-Score (Absolute) |  | 0.76 | 1.18 | 2.77 | 0.05 | 0.12 | 1.3 | 0.09 | 2.19 | 0.37 | 0.14 | 2.69 | 0.38 | 1.2 | 0.97 | 0.13 | 0.6 | 0.75 | 0.42 | 0.2 | 1.18 | 1.26 | 1.2 | 0.05 | 0.43
| Analytical results for Ta in OREAS 282 (Certified Value < 0.01 ppm) | [26 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.01 | < 0.005 | < 5 | < 0.05 | 0.08 | < 0.5 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | 0.006 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 10 | 0.01 | < 0.01 | < 0.01 |  |  |  |  | 1
| 1 | 2 | < 0.01 | < 0.005 | < 5 | < 0.05 | 0.08 | < 0.5 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | 0.0058 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 10 | 0.01 | < 0.01 | < 0.01 |  |  |  |  | 22
| 1 | 3 | < 0.01 | < 0.005 | < 5 | < 0.05 | 0.08 | < 0.5 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | 0.0061 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 10 | 0.01 | < 0.01 | < 0.01 |  |  |  |  | 16
| 1 | 4 | < 0.01 | < 0.005 | < 5 | < 0.05 | 0.12 | < 0.5 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | 0.0057 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 10 | 0.01 | < 0.01 | < 0.01 |  |  |  |  | < 0.01
| 1 | 5 | < 0.01 | < 0.005 | < 5 | < 0.05 | 0.19 | < 0.5 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | 0.0059 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 10 | 0.01 | < 0.01 | < 0.01 |  |  |  |  | 117
| 1 | 6 | < 0.01 | < 0.005 | < 5 | < 0.05 | 0.08 | < 0.5 | < 0.1 | < 0.01 | < 0.01 | < 0.05 | 0.006 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 10 | 0.01 | < 0.01 | < 0.01
| Mean | [5 empty] | 0.105 | [5 empty] | 0.0059167 | [7 empty] | 0.01
| Median | [5 empty] | 0.08 | [5 empty] | 0.00595 | [7 empty] | 0.01
| Std Dev. | [5 empty] | 0.044609 | [5 empty] | 0.0001472 | [7 empty] | 0
| Rel.Std.Dev. | [5 empty] | 42.49% | [5 empty] | 2.488% | [7 empty] | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.76 | 83.45 | 0 | 2.7 | 7.59 | 0.84 | 0.67 | 0.67 | 0 | 0.64 | 0.67 | 0 | 0 | 0 | 0.67 | 0 | 167.73 | 0.51 | 0.67 | 0.67
| Analytical results for Tb in OREAS 282 (Certified Value 0.42 ppm) | [26 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.02 | 2.04 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | 2.18 | 2.19 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [15 empty] | 2
|  |  | 04g | 0.25g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | 0.25g | 0.4g | [15 empty] | 3
| 1 | 1 | 0.45 | 0.46 | 0.4 | 0.37 | 0.436 | 0.42 | 0.436 | 0.4 | 0.41 | 0.39 | [15 empty] | 1
| 1 | 2 | 0.44 | 0.44 | 0.45 | 0.36 | 0.441 | 0.39 | 0.439 | 0.4 | 0.41 | 0.38 | [15 empty] | 23
| 1 | 3 | 0.445 | 0.48 | 0.4 | 0.385 | 0.431 | 0.39 | 0.435 | 0.5 | 0.41 | 0.38 | [15 empty] | 16
| 1 | 4 | 0.45 | 0.47 | 0.4 | 0.37 | 0.436 | 0.39 | 0.438 | 0.4 | 0.4 | 0.37 | [15 empty] | 0.41656
| 1 | 5 | 0.445 | 0.46 | 0.4 | 0.39 | 0.44 | 0.39 | 0.431 | 0.4 | 0.41 | 0.4 | [15 empty] | 118
| 1 | 6 | 0.445 | 0.47 | 0.45 | 0.365 | 0.434 | 0.38 | 0.434 | 0.5 | 0.41 | 0.37
| Mean |  | 0.44583 | 0.46333 | 0.41667 | 0.37333 | 0.43633 | 0.39333 | 0.4355 | 0.43333 | 0.40833 | 0.38167
| Median |  | 0.445 | 0.465 | 0.4 | 0.37 | 0.436 | 0.39 | 0.4355 | 0.4 | 0.41 | 0.38
| Std Dev. |  | 0.0037639 | 0.013663 | 0.02582 | 0.01169 | 0.0037238 | 0.013663 | 0.002881 | 0.05164 | 0.0040825 | 0.01169
| Rel.Std.Dev. |  | 0.8442% | 2.949% | 6.197% | 3.131% | 0.8534% | 3.474% | 0.6615% | 11.92% | 0.9998% | 3.063%
| PDM3 |  | 7.029% | 11.23% | 0.02667% | -10.38% | 4.748% | -5.575% | 4.548% | 4.028% | -1.974% | -8.376%
| Z-Score (Absolute) |  | 0.84 | 1.35 | 0 | 1.25 | 0.57 | 0.67 | 0.54 | NA | 0.24 | 1.01
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 282 (Certified Value 0.24 ppm) | [26 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 0.22 | 0.24 | < 5 | 0.22 | 0.25 | 0.2 | 0.2 | 0.2 | 0.24 | 0.29 | 0.24 | 0.216 | 0.28 | 0.26 | 0.21 | 0.2 | 0.22 | < 0.05 | 0.22 | 0.23 | < 10 | 0.26 | 0.33 | 0.26 |  | 1
| 1 | 2 | 0.23 | 0.26 | < 5 | 0.23 | 0.25 | 0.2 | < 0.2 | < 0.1 | 0.24 | 0.2 | 0.23 | 0.227 | 0.3 | 0.3 | 0.2 | 0.2 | 0.23 | < 0.05 | 0.27 | 0.21 | < 10 | 0.28 | 0.29 | 0.24 |  | 24
| 1 | 3 | 0.24 | 0.26 | < 5 | 0.21 | 0.26 | 0.2 | 0.2 | 0.1 | 0.23 | 0.31 | 0.24 | 0.22 | 0.27 | 0.27 | 0.2 | 0.2 | 0.22 | < 0.05 | 0.26 | 0.22 | < 10 | 0.26 | 0.37 | 0.24 |  | 16
| 1 | 4 | 0.25 | 0.22 | < 5 | 0.2 | 0.22 | 0.2 | 0.2 | 0.2 | 0.25 | 0.24 | 0.25 | 0.218 | 0.25 | 0.25 | 0.18 | 0.2 | 0.22 | < 0.05 | 0.25 | 0.22 | < 10 | 0.26 | 0.22 | 0.19 |  | 0.24286
| 1 | 5 | 0.23 | 0.22 | < 5 | 0.23 | 0.27 | 0.2 | 0.2 | 0.2 | 0.27 | 0.32 | 0.25 | 0.215 | 0.26 | 0.26 | 0.21 | 0.2 | 0.23 | < 0.05 | 0.25 | 0.22 | < 10 | 0.22 | 0.33 | 0.24 |  | 119
| 1 | 6 | 0.23 | 0.24 | < 5 | 0.23 | 0.25 | 0.2 | 0.2 | 0.2 | 0.21 | 0.29 | 0.23 | 0.216 | 0.22 | 0.26 | 0.19 | 0.2 | 0.22 | < 0.05 | 0.25 | 0.21 | < 10 | 0.25 | 0.29 | 0.25
| Mean |  | 0.23333 | 0.24 |  | 0.22 | 0.25 | 0.2 | 0.2 | 0.18 | 0.24 | 0.275 | 0.24 | 0.21867 | 0.26333 | 0.26667 | 0.19833 | 0.2 | 0.22333 |  | 0.25 | 0.21833 |  | 0.255 | 0.305 | 0.23667
| Median |  | 0.23 | 0.24 |  | 0.225 | 0.25 | 0.2 | 0.2 | 0.2 | 0.24 | 0.29 | 0.24 | 0.217 | 0.265 | 0.26 | 0.2 | 0.2 | 0.22 |  | 0.25 | 0.22 |  | 0.26 | 0.31 | 0.24
| Std Dev. |  | 0.010328 | 0.017889 |  | 0.012649 | 0.016733 | 3.0405e-17 | 0 | 0.044721 | 0.02 | 0.045935 | 0.0089443 | 0.0044572 | 0.027325 | 0.017512 | 0.01169 | 3.0405e-17 | 0.005164 |  | 0.016733 | 0.0075277 |  | 0.019748 | 0.051284 | 0.024221
| Rel.Std.Dev. |  | 4.426% | 7.454% |  | 5.75% | 6.693% | 1.52e-14% | 0% | 24.85% | 8.333% | 16.7% | 3.727% | 2.038% | 10.38% | 6.567% | 5.894% | 1.52e-14% | 2.312% |  | 6.693% | 3.448% |  | 7.744% | 16.81% | 10.23%
| PDM3 |  | -3.924% | -1.179% |  | -9.414% | 2.939% | -17.65% | -17.65% | -25.88% | -1.179% | 13.23% | -1.179% | -9.963% | 8.429% | 9.801% | -18.34% | -17.65% | -8.041% |  | 2.939% | -10.1% |  | 4.998% | 25.59% | -2.551%
| Z-Score (Absolute) |  | 0.22 | 0 | 73.74 | 0.65 | 0.33 | NA | NA | NA | 0 | 1.14 | 0 | 0.7 | 0.76 | 0.87 | 1.36 | NA | 0.54 | 7.01 | 0.33 | 0.71 | 155.31 | 0.49 | 2.12 | 0.11
Results from laboratories 7, 8, 9 and 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 282 (Certified Value 4.71 ppm) | [26 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 04g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 4.8 | 5.5 | 4.7 | 4.5 | 3.65 | 4.4 | 4.855 | 4.7 | 4.7 | 2.35 | 4.85 | 5.9 | 5 | 4.81 | 4.5 | 5.06 | 5.7 | 4.4 | < 10 | 5.1 | 5.7 | 4.7 |  |  |  | 1
| 1 | 2 | 4.8 | 5.52 | 4.5 | 5 | 3.35 | 4.4 | 4.833 | 4.5 | 4.54 | 5.35 | 4.94 | 5.9 | 5 | 4.55 | 4.6 | 5.01 | 5.7 | 4.6 | < 10 | 4.7 | 5.8 | 4.4 |  |  |  | 25
| 1 | 3 | 4.8 | 5.54 | 4.7 | 4.4 | 3.35 | 4.3 | 4.777 | 4.5 | 4.69 | 5.34 | 4.79 | 6 | 4.9 | 4.57 | 4.6 | 5 | 5.6 | 4.4 | < 10 | 4.7 | 5.9 | 4.5 |  |  |  | 16
| 1 | 4 | 4.7 | 5.58 | 4.7 | 4.5 | 3.5 | 4.6 | 4.796 | 4.5 | 4.55 | 5.28 | 4.97 | 5.5 | 4.9 | 4.55 | 4.6 | 5 | 5.6 | 4.6 | < 10 | 4.7 | 5.7 | 4.4 |  |  |  | 4.7125
| 1 | 5 | 4.8 | 5.6 | 4.5 | 4.6 | 3.6 | 4.5 | 4.817 | 4.5 | 4.6 | 5.25 | 4.8 | 5.8 | 5 | 4.71 | 4.5 | 5.04 | 5.3 | 4.3 | < 10 | 4.5 | 5.6 | 4.5 |  |  |  | 120
| 1 | 6 | 4.7 | 5.54 | 4.7 | 4.5 | 3.55 | 4.4 | 4.841 | 4.5 | 4.53 | 2.31 | 4.8 | 6.1 | 5 | 4.72 | 4.6 | 5.03 | 5.6 | 5 | < 10 | 4.6 | 5.6 | 4.8
| Mean |  | 4.7667 | 5.5467 | 4.6333 | 4.5833 | 3.5 | 4.4333 | 4.8198 | 4.5333 | 4.6017 | 4.3133 | 4.8583 | 5.8667 | 4.9667 | 4.6517 | 4.5667 | 5.0233 | 5.5833 | 4.55 |  | 4.7167 | 5.7167 | 4.55
| Median |  | 4.8 | 5.54 | 4.7 | 4.5 | 3.525 | 4.4 | 4.825 | 4.5 | 4.575 | 5.265 | 4.825 | 5.9 | 5 | 4.64 | 4.6 | 5.02 | 5.6 | 4.5 |  | 4.7 | 5.7 | 4.5
| Std Dev. |  | 0.05164 | 0.037238 | 0.10328 | 0.2137 | 0.12649 | 0.10328 | 0.029219 | 0.08165 | 0.076267 | 1.5368 | 0.078337 | 0.20656 | 0.05164 | 0.10998 | 0.05164 | 0.024221 | 0.1472 | 0.251 |  | 0.20412 | 0.1169 | 0.16432
| Rel.Std.Dev. |  | 1.083% | 0.6714% | 2.229% | 4.662% | 3.614% | 2.33% | 0.6062% | 1.801% | 1.657% | 35.63% | 1.612% | 3.521% | 1.04% | 2.364% | 1.131% | 0.4822% | 2.636% | 5.516% |  | 4.328% | 2.045% | 3.611%
| PDM3 |  | 1.15% | 17.7% | -1.68% | -2.741% | -25.73% | -5.924% | 2.278% | -3.802% | -2.352% | -8.47% | 3.095% | 24.49% | 5.394% | -1.291% | -3.094% | 6.596% | 18.48% | -3.448% |  | 0.08864% | 21.31% | -3.448%
| Z-Score (Absolute) |  | 0.34 | 3.58 | 0.21 | 0.42 | 4.91 | 1.04 | 0.56 | 0.63 | 0.34 | 1.54 | 0.72 | 4.91 | 1.17 | 0.13 | 0.49 | 1.41 | 3.73 | 0.56 | 1.31 | 0.13 | 4.28 | 0.56
| Analytical results for Ti in OREAS 282 (Certified Value 0.02 wt.%) | [26 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 3
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.022 | 0.015 | 0.02 | 0.02 | 0.0115 | 0.015 | 0.021 | 0.02 | 0.031 | 0.0289 | 0.0268 | 0.02 | 0.02 | 0.0215 | 0.019 | 0.0276 | 0.018 | 0.018 | 0.02 | 0.024 | 0.02 | 0.022 |  |  |  | 1
| 1 | 2 | 0.022 | 0.015 | 0.02 | 0.02 | 0.013 | 0.015 | 0.022 | 0.0203 | 0.03 | 0.029 | 0.0277 | 0.02 | 0.02 | 0.022 | 0.018 | 0.0269 | 0.02 | 0.018 | 0.02 | 0.023 | 0.021 | 0.021 |  |  |  | 26
| 1 | 3 | 0.022 | 0.015 | 0.02 | 0.02 | 0.012 | 0.015 | 0.021 | 0.0194 | 0.031 | 0.0294 | 0.0273 | 0.02 | 0.02 | 0.0215 | 0.018 | 0.0278 | 0.02 | 0.017 | 0.02 | 0.023 | 0.021 | 0.021 |  |  |  | 16
| 1 | 4 | 0.022 | 0.015 | 0.02 | 0.02 | 0.012 | 0.015 | 0.022 | 0.0195 | 0.03 | 0.0288 | 0.0272 | 0.02 | 0.02 | 0.0215 | 0.018 | 0.0272 | 0.018 | 0.018 | 0.02 | 0.023 | 0.021 | 0.022 |  |  |  | 0.019677
| 1 | 5 | 0.022 | 0.015 | 0.02 | 0.02 | 0.012 | 0.015 | 0.022 | 0.0199 | 0.03 | 0.0291 | 0.0267 | 0.02 | 0.02 | 0.022 | 0.019 | 0.0269 | 0.019 | 0.018 | 0.02 | 0.021 | 0.02 | 0.021 |  |  |  | 121
| 1 | 6 | 0.023 | 0.015 | 0.02 | 0.02 | 0.012 | 0.015 | 0.021 | 0.019 | 0.03 | 0.0292 | 0.0271 | 0.02 | 0.02 | 0.0225 | 0.017 | 0.0271 | 0.019 | 0.016 | 0.02 | 0.022 | 0.021 | 0.021
| Mean |  | 0.022167 | 0.015 | 0.02 | 0.02 | 0.012083 | 0.015 | 0.0215 | 0.019683 | 0.030333 | 0.029067 | 0.027133 | 0.02 | 0.02 | 0.021833 | 0.018167 | 0.02725 | 0.019 | 0.0175 | 0.02 | 0.022667 | 0.020667 | 0.021333
| Median |  | 0.022 | 0.015 | 0.02 | 0.02 | 0.012 | 0.015 | 0.0215 | 0.0197 | 0.03 | 0.02905 | 0.02715 | 0.02 | 0.02 | 0.02175 | 0.018 | 0.02715 | 0.019 | 0.018 | 0.02 | 0.023 | 0.021 | 0.021
| Std Dev. |  | 0.00040825 | 0 | 0 | 0 | 0.0004916 | 0 | 0.00054772 | 0.00047081 | 0.0005164 | 0.00021602 | 0.00036148 | 0 | 0 | 0.00040825 | 0.00075277 | 0.00037283 | 0.00089443 | 0.00083666 | 0 | 0.0010328 | 0.0005164 | 0.0005164
| Rel.Std.Dev. |  | 1.842% | 0% | 0% | 0% | 4.068% | 0% | 2.548% | 2.392% | 1.702% | 0.7432% | 1.332% | 0% | 0% | 1.87% | 4.144% | 1.368% | 4.708% | 4.781% | 0% | 4.556% | 2.499% | 2.421%
| PDM3 |  | 12.65% | -23.77% | 1.639% | 1.639% | -38.59% | -23.77% | 9.262% | 0.02989% | 54.15% | 47.72% | 37.89% | 1.639% | 1.639% | 10.96% | -7.678% | 38.48% | -3.443% | -11.07% | 1.639% | 15.19% | 5.027% | 8.415%
| Z-Score (Absolute) |  | 0.8 | 1.84 | 0 | 0 | 2.91 | 1.84 | 0.55 | 0.12 | 3.8 | 3.33 | 2.62 | 0 | 0 | 0.67 | 0.67 | 2.67 | 0.37 | 0.92 | 0 | 0.98 | 0.25 | 0.49
| Analytical results for Tl in OREAS 282 (Certified Value 25.3 ppm) | [26 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 27.4 | 21 | 16 | 25.57 | 23.05 | 21.5 | 21.9 | 26.45 | 27.52 | 27.11 | 17.67 | 25.1 | 25.1 | 26.4 | 26.1 | 24.2 | 26 | 27.44 | 23.77 | 25 | 27.1 | 26.54 | 26.6 |  |  | 1
| 1 | 2 | 26.9 | 20 | 19 | 24.91 | 25.15 | 21.4 | 22.3 | 26.51 | 26.55 | 26.53 | 17.99 | 25.89 | 24 | 26.44 | 24.7 | 24.6 | 25.8 | 27.32 | 24.58 | 29 | 25.3 | 25.55 | 25.1 |  |  | 27
| 1 | 3 | 27.1 | 20.1 | 16 | 25.64 | 22.95 | 23.5 | 21.5 | 26.58 | 25.93 | 27.13 | 17.87 | 25.58 | 25.3 | 26.9 | 24.7 | 24.8 | 25.2 | 26.65 | 23.97 | 27 | 25.8 | 25.42 | 25.7 |  |  | 16
| 1 | 4 | 26.6 | 20.8 | 17 | 26.44 | 23.45 | 24.4 | 22.6 | 26.28 | 26.73 | 26.66 | 17.83 | 25.89 | 23.9 | 25.97 | 25.1 | 24.3 | 25.2 | 26.46 | 24.48 | 25 | 25.3 | 26.24 | 25.4 |  |  | 25.337
| 1 | 5 | 26.8 | 21.1 | 16 | 25.26 | 23.35 | 23.8 | 22.4 | 26.53 | 26.23 | 26.82 | 17.55 | 25.14 | 24.6 | 26.27 | 25.5 | 24.5 | 27.3 | 26.46 | 23.36 | 31 | 24.3 | 26.08 | 25.5 |  |  | 122
| 1 | 6 | 26.6 | 20.3 | 17 | 26.18 | 23.6 | 22.4 | 22.2 | 26.4 | 26.14 | 26.33 | 27.62 | 25.28 | 24.9 | 26.81 | 26 | 24.5 | 25.8 | 26.56 | 24.9 | 25 | 25.3 | 25.58 | 26.9
| Mean |  | 26.9 | 20.55 | 16.8333 | 25.6667 | 23.5917 | 22.8333 | 22.15 | 26.4583 | 26.5167 | 26.7633 | 19.4217 | 25.48 | 24.6333 | 26.465 | 25.35 | 24.4833 | 25.8833 | 26.815 | 24.1767 | 27 | 25.5167 | 25.9017 | 25.8667
| Median |  | 26.85 | 20.55 | 16.5 | 25.605 | 23.4 | 22.95 | 22.25 | 26.48 | 26.39 | 26.74 | 17.85 | 25.43 | 24.75 | 26.42 | 25.3 | 24.5 | 25.8 | 26.605 | 24.225 | 26 | 25.3 | 25.83 | 25.6
| Std Dev. |  | 0.30984 | 0.47645 | 1.169 | 0.56737 | 0.80151 | 1.2533 | 0.3937 | 0.10759 | 0.56948 | 0.3196 | 4.0193 | 0.3595 | 0.5785 | 0.3453 | 0.61887 | 0.2137 | 0.7705 | 0.44496 | 0.57497 | 2.5298 | 0.91742 | 0.45009 | 0.7174
| Rel.Std.Dev. |  | 1.152% | 2.318% | 6.945% | 2.211% | 3.397% | 5.489% | 1.777% | 0.4067% | 2.148% | 1.194% | 20.7% | 1.411% | 2.348% | 1.305% | 2.441% | 0.8728% | 2.977% | 1.659% | 2.378% | 9.37% | 3.595% | 1.738% | 2.773%
| PDM3 |  | 6.169% | -18.89% | -33.56% | 1.301% | -6.888% | -9.881% | -12.58% | 4.426% | 4.656% | 5.629% | -23.35% | 0.5644% | -2.777% | 4.452% | 0.05131% | -3.369% | 2.156% | 5.833% | -4.58% | 6.564% | 0.7091% | 2.229% | 2.09%
| Z-Score (Absolute) |  | 0.9 | 3.24 | 5.67 | 0.1 | 1.26 | 1.75 | 2.2 | 0.61 | 0.65 | 0.81 | 3.98 | 0.02 | 0.58 | 0.62 | 0.11 | 0.67 | 0.24 | 0.85 | 0.87 | 0.97 | 0 | 0.25 | 0.23
| Analytical results for Tm in OREAS 282 (Certified Value 0.17 ppm) | [26 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.02 | 2.07 | 2.08 | 2.10 | 2.12 | 2.14 | 2.16 | [18 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [18 empty] | 2
|  |  | 04g | 25g | 40g | 0.5g | 25g | 0.25g | 0.5g | [18 empty] | 3
| 1 | 1 | 0.18 | 0.15 | 0.155 | 0.182 | 0.18 | 0.174 | 0.2 | [18 empty] | 1
| 1 | 2 | 0.175 | 0.2 | 0.15 | 0.182 | 0.17 | 0.184 | 0.2 | [18 empty] | 28
| 1 | 3 | 0.18 | 0.15 | 0.165 | 0.178 | 0.17 | 0.183 | 0.2 | [18 empty] | 16
| 1 | 4 | 0.175 | 0.2 | 0.16 | 0.178 | 0.17 | 0.175 | 0.2 | [18 empty] | 0.17192
| 1 | 5 | 0.18 | 0.15 | 0.155 | 0.185 | 0.17 | 0.181 | 0.2 | [18 empty] | 123
| 1 | 6 | 0.175 | 0.15 | 0.15 | 0.185 | 0.17 | 0.172 | 0.2
| Mean |  | 0.1775 | 0.16667 | 0.15583 | 0.18167 | 0.17167 | 0.17817 | 0.2
| Median |  | 0.1775 | 0.15 | 0.155 | 0.182 | 0.17 | 0.178 | 0.2
| Std Dev. |  | 0.0027386 | 0.02582 | 0.0058452 | 0.0031411 | 0.0040825 | 0.0051153 | 3.0405e-17
| Rel.Std.Dev. |  | 1.543% | 15.49% | 3.751% | 1.729% | 2.378% | 2.871% | 1.52e-14%
| PDM3 |  | 3.248% | -3.054% | -9.355% | 5.671% | -0.1454% | 3.635% | 16.34%
| Z-Score (Absolute) |  | 0.37 | 1 | 2.37 | 0.9 | 0.37 | 0.45 | NA
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 282 (Certified Value 6.17 ppm) | [26 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 04g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 6.25 | 6.3 | 6.22 | 5.8 | 4.95 | 4.4 | 6.277 | 6.05 | 5.72 | 4.93 | 6.11 | 6.8 | 6.31 | 6.14 | 5.65 | 6.5 | 6.9 | 6.06 | < 10 | 6.48 | 6.32 | 6.18 |  |  |  | 1
| 1 | 2 | 6.12 | 6.24 | 6.09 | 6.5 | 5 | 4.5 | 6.284 | 5.88 | 5.5 | 4.88 | 6.24 | 6.8 | 6.29 | 5.78 | 5.8 | 6.38 | 6.9 | 6.42 | < 10 | 6.05 | 6.04 | 5.85 |  |  |  | 29
| 1 | 3 | 6.17 | 6.26 | 6.17 | 5.8 | 5.2 | 4.3 | 6.176 | 5.77 | 5.71 | 4.92 | 6.14 | 7 | 6.3 | 5.78 | 5.84 | 6.49 | 6.8 | 6.17 | < 10 | 6.25 | 6.49 | 6.01 |  |  |  | 16
| 1 | 4 | 6.07 | 6.34 | 6.38 | 5.9 | 5.5 | 4.6 | 6.211 | 5.84 | 5.47 | 4.85 | 6.33 | 6.5 | 6.1 | 5.86 | 5.72 | 6.55 | 6.5 | 6.27 | < 10 | 6.11 | 6.36 | 5.97 |  |  |  | 6.1674
| 1 | 5 | 6.15 | 6.34 | 6.08 | 6 | 5.4 | 4.5 | 6.2 | 5.86 | 5.53 | 4.84 | 6.1 | 6.8 | 6.21 | 6.03 | 5.75 | 6.58 | 6.7 | 6.1 | < 10 | 5.74 | 6.34 | 5.93 |  |  |  | 124
| 1 | 6 | 5.97 | 6.26 | 6.33 | 5.9 | 4.95 | 4.5 | 6.216 | 5.78 | 5.45 | 4.83 | 6.13 | 7 | 6.29 | 5.97 | 5.73 | 6.5 | 6.6 | 6.55 | < 10 | 6.07 | 6.16 | 6.32
| Mean |  | 6.1217 | 6.29 | 6.2117 | 5.9833 | 5.1667 | 4.4667 | 6.2273 | 5.8633 | 5.5633 | 4.875 | 6.175 | 6.8167 | 6.25 | 5.9267 | 5.7483 | 6.5 | 6.7333 | 6.2617 |  | 6.1167 | 6.285 | 6.0433
| Median |  | 6.135 | 6.28 | 6.195 | 5.9 | 5.1 | 4.5 | 6.2135 | 5.85 | 5.515 | 4.865 | 6.135 | 6.8 | 6.29 | 5.915 | 5.74 | 6.5 | 6.75 | 6.22 |  | 6.09 | 6.33 | 5.99
| Std Dev. |  | 0.095167 | 0.043359 | 0.12352 | 0.26394 | 0.24014 | 0.10328 | 0.043486 | 0.10132 | 0.12061 | 0.042308 | 0.09094 | 0.18348 | 0.081731 | 0.14528 | 0.066156 | 0.068411 | 0.1633 | 0.19156 |  | 0.24427 | 0.15972 | 0.17432
| Rel.Std.Dev. |  | 1.555% | 0.6893% | 1.988% | 4.411% | 4.648% | 2.312% | 0.6983% | 1.728% | 2.168% | 0.8679% | 1.473% | 2.692% | 1.308% | 2.451% | 1.151% | 1.052% | 2.425% | 3.059% |  | 3.993% | 2.541% | 2.884%
| PDM3 |  | -0.7414% | 1.988% | 0.7179% | -2.984% | -16.23% | -27.58% | 0.972% | -4.93% | -9.794% | -20.96% | 0.1234% | 10.53% | 1.339% | -3.903% | -6.795% | 5.393% | 9.176% | 1.529% |  | -0.8224% | 1.907% | -2.011%
| Z-Score (Absolute) |  | 0.01 | 0.63 | 0.34 | 0.5 | 3.54 | 6.13 | 0.4 | 0.95 | 2.06 | 4.62 | 0.21 | 2.59 | 0.49 | 0.71 | 1.38 | 1.41 | 2.28 | 0.53 | 4.15 | 0.01 | 0.62 | 0.28
| Analytical results for V in OREAS 282 (Certified Value 60 ppm) | [26 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 0
|  |  | 0.5g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | 58 | 56 | 55 | 63.4 | 50 | 50 | 68 | 64 | 72 | 59 | 67 | 82 | 61 | 66 | 56 | 65 | 60 | 54 | 66 | 63 | 60 | 57 |  |  |  | 1
| 1 | 2 | 57 | 57 | 55 | 61.3 | 50 | 50 | 69 | 64.5 | 68 | 62 | 69 | 81 | 60 | 56 | 55 | 64 | 60 | 56 | 66 | 60 | 60 | 56 |  |  |  | 30
| 1 | 3 | 57 | 56 | 55 | 62.4 | 50 | 60 | 67 | 62.6 | 72 | 62 | 67 | 85 | 60 | 65 | 55 | 66 | 59 | 55 | 66 | 59 | 62 | 56 |  |  |  | 16
| 1 | 4 | 58 | 56 | 55 | 63.4 | 50 | 60 | 71 | 62.4 | 68 | 62 | 67 | 77 | 60 | 52 | 55 | 64 | 58 | 55 | 65 | 60 | 61 | 57 |  |  |  | 60.3558
| 1 | 5 | 58 | 54 | 55 | 61.6 | 50 | 60 | 69 | 65.1 | 68 | 63 | 66 | 81 | 61 | 64 | 57 | 64 | 60 | 55 | 64 | 57 | 59 | 57 |  |  |  | 125
| 1 | 6 | 58 | 57 | 56 | 61.5 | 50 | 60 | 67 | 59.5 | 66 | 59 | 68 | 83 | 58 | 60 | 54 | 64 | 59 | 57 | 66 | 59 | 59 | 57
| Mean |  | 57.6667 | 56 | 55.1667 | 62.2667 | 50 | 56.6667 | 68.5 | 63.0167 | 69 | 61.1667 | 67.3333 | 81.5 | 60 | 60.5 | 55.3333 | 64.5 | 59.3333 | 55.3333 | 65.5 | 59.6667 | 60.1667 | 56.6667
| Median |  | 58 | 56 | 55 | 62 | 50 | 60 | 68.5 | 63.3 | 68 | 62 | 67 | 81.5 | 60 | 62 | 55 | 64 | 59.5 | 55 | 66 | 59.5 | 60 | 57
| Std Dev. |  | 0.5164 | 1.0954 | 0.40825 | 0.95429 | 0 | 5.164 | 1.5166 | 2.0213 | 2.4495 | 1.7224 | 1.0328 | 2.6646 | 1.0954 | 5.5767 | 1.0328 | 0.83666 | 0.8165 | 1.0328 | 0.83666 | 1.9664 | 1.169 | 0.5164
| Rel.Std.Dev. |  | 0.8955% | 1.956% | 0.74% | 1.533% | 0% | 9.113% | 2.214% | 3.208% | 3.55% | 2.816% | 1.534% | 3.269% | 1.826% | 9.218% | 1.866% | 1.297% | 1.376% | 1.866% | 1.277% | 3.296% | 1.943% | 0.9113%
| PDM3 |  | -4.456% | -7.217% | -8.598% | 3.166% | -17.16% | -6.112% | 13.49% | 4.409% | 14.32% | 1.343% | 11.56% | 35.03% | -0.5896% | 0.2389% | -8.321% | 6.866% | -1.694% | -8.321% | 8.523% | -1.142% | -0.3134% | -6.112%
| Z-Score (Absolute) |  | 0.4 | 0.67 | 0.81 | 0.34 | 1.65 | NA | 1.35 | 0.46 | 1.43 | 0.16 | 1.16 | 3.45 | 0.03 | 0.05 | 0.78 | 0.7 | 0.13 | 0.78 | 0.86 | 0.08 | 0 | 0.57
Results from laboratory 8 were removed due to their 10 ppm reading resolution.
| Analytical results for W in OREAS 282 (Certified Value 4.46 ppm) | [26 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 4.98 | 7.5 | 14 | < 10 | 3.4 | 4 | 4.5 | 3.9 | 5.45 | 4.63 | 5.1 | 4.03 | 3.44 | 5.3 | 4.4 | 6.2 | 4.28 | 4.4 | 4.3 | 4.47 | < 20 | 5.51 | 4.14 | 4.97 |  | 1
| 1 | 2 | 4.82 | 7.4 | 11 | < 10 | 3.3 | 4 | 4 | 4 | 5.4 | 4.5 | 4.8 | 4.02 | 3.35 | 5 | 4.5 | 6.2 | 4.16 | 4.2 | 4.5 | 4.42 | < 20 | 5.3 | 4.19 | 4.72 |  | 31
| 1 | 3 | 4.9 | 7.6 | 11 | < 10 | 3.5 | 4 | 5 | 3.9 | 5.34 | 4.46 | 4.9 | 4.01 | 3.47 | 5.1 | 4.4 | 6.1 | 4.34 | 4.2 | 4.5 | 4.42 | < 20 | 5.29 | 4.27 | 4.77 |  | 16
| 1 | 4 | 4.92 | 7.5 | 12 | < 10 | 3.3 | 4 | 5 | 4 | 5.27 | 4.51 | 4.7 | 3.98 | 3.8 | 5.1 | 4.3 | 6 | 4.09 | 4.1 | 4.3 | 4.41 | < 20 | 5.21 | 4.27 | 4.69 |  | 4.4622
| 1 | 5 | 4.89 | 7.3 | 11 | < 10 | 3.5 | 4 | 5 | 3.9 | 5.32 | 4.41 | 5.1 | 3.95 | 3.76 | 5 | 4.4 | 6.4 | 4.27 | 4.1 | 4.3 | 4.19 | < 20 | 5.12 | 3.55 | 4.73 |  | 126
| 1 | 6 | 4.8 | 7.4 | 13 | < 10 | 3.4 | 4 | 4.5 | 3.9 | 5.09 | 4.36 | 4.8 | 4.05 | 3.56 | 5.3 | 4.5 | 6.2 | 4.17 | 4.4 | 4.5 | 4.49 | < 20 | 5.44 | 4.09 | 5.08
| Mean |  | 4.885 | 7.45 | 12 |  | 3.4 | 4 | 4.6667 | 3.9333 | 5.3117 | 4.4783 | 4.9 | 4.0067 | 3.5633 | 5.1333 | 4.4167 | 6.1833 | 4.2183 | 4.2333 | 4.4 | 4.4 |  | 5.3117 | 4.085 | 4.8267
| Median |  | 4.895 | 7.45 | 11.5 |  | 3.4 | 4 | 4.75 | 3.9 | 5.33 | 4.48 | 4.85 | 4.015 | 3.515 | 5.1 | 4.4 | 6.2 | 4.22 | 4.2 | 4.4 | 4.42 |  | 5.295 | 4.165 | 4.75
| Std Dev. |  | 0.066257 | 0.10488 | 1.2649 |  | 0.089443 | 0 | 0.40825 | 0.05164 | 0.12545 | 0.093256 | 0.16733 | 0.036148 | 0.18118 | 0.13663 | 0.075277 | 0.13292 | 0.093256 | 0.13663 | 0.10954 | 0.1077 |  | 0.14386 | 0.27157 | 0.15958
| Rel.Std.Dev. |  | 1.356% | 1.408% | 10.54% |  | 2.631% | 0% | 8.748% | 1.313% | 2.362% | 2.082% | 3.415% | 0.9022% | 5.085% | 2.662% | 1.704% | 2.15% | 2.211% | 3.227% | 2.49% | 2.448% |  | 2.708% | 6.648% | 3.306%
| PDM3 |  | 9.476% | 66.96% | 168.9% |  | -23.8% | -10.36% | 4.583% | -11.85% | 19.04% | 0.3623% | 9.812% | -10.21% | -20.14% | 15.04% | -1.02% | 38.57% | -5.464% | -5.128% | -1.393% | -1.393% |  | 19.04% | -8.453% | 8.169%
| Z-Score (Absolute) |  | 0.32 | 4.02 | 10.6 | 0.48 | 1.83 | NA | 0 | 1.06 | 0.93 | 0.27 | 0.34 | 0.95 | 1.59 | 0.67 | 0.36 | 2.19 | 0.65 | 0.63 | 0.39 | 0.39 | 7.71 | 0.93 | 0.84 | 0.23
Results from laboratory 7 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 282 (Certified Value 15 ppm) | [26 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 15.35 | 15.4 | 12 | 14.1 | 13.9 | 13.4 | 13.5 | 15 | 15.67 | 14.92 | 13.54 | 15.43 | 14.54 | 16.3 | 15.17 | 14.7 | 15.15 | 14.2 | 16 | 16.45 | 15.35 | 15.75 |  |  |  | 1
| 1 | 2 | 15.4 | 14.8 | 12 | 15.3 | 13.5 | 15.3 | 13.9 | 15 | 15.76 | 14.67 | 13.5 | 15.81 | 14.84 | 16.1 | 14.98 | 15.2 | 15.25 | 14.1 | 16 | 15.05 | 15.92 | 14.65 |  |  |  | 32
| 1 | 3 | 15.25 | 15.1 | 12 | 14 | 13.9 | 13.4 | 14.4 | 14.4 | 15.19 | 14.8 | 13.6 | 15.65 | 14.47 | 16.7 | 15.29 | 15.1 | 15.1 | 14.3 | 16 | 15.15 | 16.05 | 15.1 |  |  |  | 16
| 1 | 4 | 15.5 | 15.3 | 12 | 14.6 | 13.8 | 14.1 | 15.4 | 15.1 | 15.42 | 15.17 | 13.3 | 15.72 | 14.82 | 15.5 | 15.17 | 14.3 | 15.15 | 14.4 | 16 | 15.4 | 15.55 | 14.1 |  |  |  | 14.9715
| 1 | 5 | 15.2 | 15.5 | 12 | 14.6 | 14 | 14 | 14.8 | 15 | 15.65 | 15.11 | 13.35 | 15.33 | 14.38 | 15.9 | 14.92 | 15.1 | 15.1 | 14.4 | 16 | 14.7 | 15.2 | 15.25 |  |  |  | 127
| 1 | 6 | 14.65 | 15.1 | 13 | 14.4 | 13.8 | 14 | 13.5 | 14.9 | 15.4 | 15.09 | 13.21 | 15.63 | 14.6 | 16.9 | 14.94 | 14.6 | 15.6 | 14.2 | 16 | 15.65 | 14.8 | 15.9
| Mean |  | 15.225 | 15.2 | 12.1667 | 14.5 | 13.8167 | 14.0333 | 14.25 | 14.9 | 15.515 | 14.96 | 13.4167 | 15.595 | 14.6083 | 16.2333 | 15.0783 | 14.8333 | 15.225 | 14.2667 | 16 | 15.4 | 15.4783 | 15.125
| Median |  | 15.3 | 15.2 | 12 | 14.5 | 13.85 | 14 | 14.15 | 15 | 15.535 | 15.005 | 13.425 | 15.64 | 14.57 | 16.2 | 15.075 | 14.9 | 15.15 | 14.25 | 16 | 15.275 | 15.45 | 15.175
| Std Dev. |  | 0.30125 | 0.25298 | 0.40825 | 0.46476 | 0.17224 | 0.69474 | 0.76092 | 0.25298 | 0.21455 | 0.19738 | 0.15267 | 0.18086 | 0.18681 | 0.5164 | 0.15198 | 0.3559 | 0.1917 | 0.12111 | 0 | 0.60663 | 0.46499 | 0.67583
| Rel.Std.Dev. |  | 1.979% | 1.664% | 3.355% | 3.205% | 1.247% | 4.951% | 5.34% | 1.698% | 1.383% | 1.319% | 1.138% | 1.16% | 1.279% | 3.181% | 1.008% | 2.399% | 1.259% | 0.8489% | 0% | 3.939% | 3.004% | 4.468%
| PDM3 |  | 1.693% | 1.526% | -18.73% | -3.149% | -7.713% | -6.266% | -4.819% | -0.4775% | 3.63% | -0.07675% | -10.39% | 4.165% | -2.426% | 8.428% | 0.7136% | -0.9228% | 1.693% | -4.708% | 6.87% | 2.862% | 3.385% | 1.025%
| Z-Score (Absolute) |  | 0.32 | 0.28 | NA | 0.8 | 1.86 | 1.53 | 1.19 | 0.18 | 0.77 | 0.09 | 2.48 | 0.89 | 0.64 | 1.88 | 0.09 | 0.29 | 0.32 | 1.17 | NA | 0.59 | 0.71 | 0.16
Results from laboratories 3 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 282 (Certified Value 1.11 ppm) | [26 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.07 | 2.10 | 2.12 | 2.14 | 2.16 | 2.18 | 2.19 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [15 empty] | 2
|  |  | 04g | 0.25g | 0.25g | 25g | 0.5g | 25g | 0.25g | 0.5g | 0.25g | 0.4g | [15 empty] | 3
| 1 | 1 | 1.16 | 1 | 1.4 | 1.15 | 1.133 | 1.08 | 1.11 | 1.2 | 1.1 | 1 | [15 empty] | 1
| 1 | 2 | 1.14 | 1 | 1.5 | 1.25 | 1.115 | 1.06 | 1.13 | 1.2 | 1.1 | 1 | [15 empty] | 33
| 1 | 3 | 1.13 | 1 | 1.6 | 1.1 | 1.146 | 1.05 | 1.12 | 1.2 | 1.1 | 1 | [15 empty] | 16
| 1 | 4 | 1.14 | 1 | 1.4 | 1.15 | 1.111 | 1.06 | 1.05 | 1.1 | 1.1 | 1 | [15 empty] | 1.1081
| 1 | 5 | 1.15 | 1 | 1.3 | 1.15 | 1.14 | 1.07 | 1.1 | 1.2 | 1.1 | 1.1 | [15 empty] | 128
| 1 | 6 | 1.15 | 1 | 1.4 | 1.15 | 1.133 | 1.02 | 1.07 | 1.3 | 1.1 | 1
| Mean |  | 1.145 | 1 | 1.4333 | 1.1583 | 1.1297 | 1.0567 | 1.0967 | 1.2 | 1.1 | 1.0167
| Median |  | 1.145 | 1 | 1.4 | 1.15 | 1.133 | 1.06 | 1.105 | 1.2 | 1.1 | 1
| Std Dev. |  | 0.010488 | 0 | 0.10328 | 0.04916 | 0.013852 | 0.020656 | 0.030768 | 0.063246 | 2.4324e-16 | 0.040825
| Rel.Std.Dev. |  | 0.916% | 0% | 7.206% | 4.244% | 1.226% | 1.955% | 2.806% | 5.27% | 2.211e-14% | 4.016%
| PDM3 |  | 3.332% | -9.754% | 29.35% | 4.535% | 1.948% | -4.64% | -1.03% | 8.295% | -0.7295% | -8.25%
| Z-Score (Absolute) |  | 0.31 | NA | 6.21 | 0.59 | 0 | 1.49 | 0.67 | 1.44 | 0.61 | 2.31
Results from laboratory 4 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 282 (Certified Value 112 ppm) | [26 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.27 | 2.28 | 2.29 | 2.30 | 2.31 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.5g | 04g | 0.5g | 0.25g | 0.25g | 25g | 40g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 30g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g |  | 0
| 1 | 1 | 118 | 120 | 101 | 125 | 111 | 110 | 103 | 119 | 115.3 | 113 | 104 | 107 | 118 | 116 | 112 | 117 | 100 | 111 | 110 | 107 | 117 | 153 | 107 | 120 |  | 1
| 1 | 2 | 116 | 117 | 103 | 130.5 | 110 | 115 | 100 | 119 | 115.1 | 109 | 110 | 110 | 117 | 113 | 113 | 117 | 102 | 110.6 | 109 | 105 | 116 | 147 | 106 | 121 |  | 34
| 1 | 3 | 115 | 121 | 98 | 125.6 | 113 | 110 | 104 | 116 | 112.9 | 114 | 110 | 107 | 119 | 114 | 115 | 117 | 102 | 115.9 | 107 | 105 | 118 | 148 | 106 | 120 |  | 16
| 1 | 4 | 116 | 120 | 101 | 124.6 | 115 | 110 | 101 | 120 | 114.9 | 108 | 107 | 111 | 115 | 112 | 106 | 117 | 105 | 109.3 | 109 | 106 | 117 | 149 | 105 | 124 |  | 112.0736
| 1 | 5 | 118 | 122 | 101 | 125.4 | 110 | 110 | 104 | 120 | 116.3 | 109 | 105 | 108 | 115 | 113 | 122 | 119 | 104 | 107.9 | 110 | 105 | 116 | 143 | 106 | 121 |  | 129
| 1 | 6 | 117 | 118 | 102 | 127.7 | 111 | 110 | 103 | 117 | 114.4 | 106 | 107 | 110 | 117 | 115 | 108 | 116 | 100 | 104.6 | 108 | 108 | 118 | 147 | 105 | 121
| Mean |  | 116.6667 | 119.6667 | 101 | 126.4667 | 111.6667 | 110.8333 | 102.5 | 118.5 | 114.8167 | 109.8333 | 107.1667 | 108.8333 | 116.8333 | 113.8333 | 112.6667 | 117.1667 | 102.1667 | 109.8833 | 108.8333 | 106 | 117 | 147.8333 | 105.8333 | 121.1667
| Median |  | 116.5 | 120 | 101 | 125.5 | 111 | 110 | 103 | 119 | 115 | 109 | 107 | 109 | 117 | 113.5 | 112.5 | 117 | 102 | 109.95 | 109 | 105.5 | 117 | 147.5 | 106 | 121
| Std Dev. |  | 1.2111 | 1.8619 | 1.6733 | 2.25 | 1.9664 | 2.0412 | 1.6432 | 1.6432 | 1.1286 | 3.0605 | 2.4833 | 1.7224 | 1.6021 | 1.472 | 5.6451 | 0.98319 | 2.0412 | 3.7456 | 1.169 | 1.2649 | 0.89443 | 3.2506 | 0.75277 | 1.472
| Rel.Std.Dev. |  | 1.038% | 1.556% | 1.657% | 1.779% | 1.761% | 1.842% | 1.603% | 1.387% | 0.9829% | 2.786% | 2.317% | 1.583% | 1.371% | 1.293% | 5.01% | 0.8391% | 1.998% | 3.409% | 1.074% | 1.193% | 0.7645% | 2.199% | 0.7113% | 1.215%
| PDM3 |  | 4.098% | 6.775% | -9.881% | 12.84% | -0.3631% | -1.107% | -8.542% | 5.734% | 2.448% | -1.999% | -4.378% | -2.891% | 4.247% | 1.57% | 0.5292% | 4.544% | -8.84% | -1.954% | -2.891% | -5.419% | 4.396% | 31.91% | -5.568% | 8.113%
| Z-Score (Absolute) |  | 0.62 | 1.03 | 1.53 | 1.96 | 0.07 | 0.18 | 1.33 | 0.87 | 0.36 | 0.32 | 0.69 | 0.46 | 0.64 | 0.23 | 0.07 | 0.69 | 1.37 | 0.31 | 0.46 | 0.85 | 0.66 | 4.89 | 0.87 | 1.23
| Analytical results for Zr in OREAS 282 (Certified Value 12 ppm) | [26 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.28 | 2.29 | 2.30 | 2.31 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 04g | 0.5g | 0.25g | 25g | 40g | 0.25g | 0.5g | 25g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 12.7 | 13 | 12 | 11.4 | 10 | 10 | 10.7 | 12.85 | 16.1 | 10.8 | 8.44 | 14.1 | 15.5 | 12.2 | 10.9 | 13.5 | 11.6 | 14 | 13.5 | 12.2 | 11.7 |  |  |  |  | 1
| 1 | 2 | 12.9 | 13 | 12 | 12.9 | 10 | 10 | 11 | 12.9 | 16.1 | 10.3 | 8.8 | 14.4 | 14.9 | 11.6 | 11.1 | 13.6 | 11.7 | 13 | 12.7 | 11.3 | 10.9 |  |  |  |  | 35
| 1 | 3 | 12.6 | 13 | 11 | 11.6 | 9 | 10 | 10.8 | 12.56 | 16 | 10.7 | 8.99 | 13.9 | 16 | 11.8 | 10.7 | 13.1 | 12.1 | 13 | 13 | 12.6 | 11.4 |  |  |  |  | 16
| 1 | 4 | 12.8 | 13 | 11 | 11.3 | 10 | 10 | 11 | 12.64 | 16.4 | 9.9 | 8.63 | 14.1 | 14.6 | 12 | 10.4 | 13.7 | 11.6 | 13 | 12.8 | 12.6 | 10.6 |  |  |  |  | 12.0236
| 1 | 5 | 12.8 | 13 | 11 | 11.7 | 10 | 10 | 11.1 | 12.87 | 16.5 | 10.2 | 8.36 | 13.5 | 15 | 11.8 | 10.8 | 13.4 | 11.9 | 13 | 12.2 | 11.5 | 11.3 |  |  |  |  | 130
| 1 | 6 | 12.1 | 13 | 12 | 12.6 | 10 | 9 | 10.6 | 12.28 | 16.3 | 10 | 8.74 | 13.9 | 15.8 | 11.6 | 10.3 | 13.5 | 11 | 13 | 12.9 | 12 | 12.1
| Mean |  | 12.65 | 13 | 11.5 | 11.9167 | 9.8333 | 9.8333 | 10.8667 | 12.6833 | 16.2333 | 10.3167 | 8.66 | 13.9833 | 15.3 | 11.8333 | 10.7 | 13.4667 | 11.65 | 13.1667 | 12.85 | 12.0333 | 11.3333
| Median |  | 12.75 | 13 | 11.5 | 11.65 | 10 | 10 | 10.9 | 12.745 | 16.2 | 10.25 | 8.685 | 14 | 15.25 | 11.8 | 10.75 | 13.5 | 11.65 | 13 | 12.85 | 12.1 | 11.35
| Std Dev. |  | 0.2881 | 0 | 0.54772 | 0.66758 | 0.40825 | 0.40825 | 0.19664 | 0.24055 | 0.19664 | 0.3656 | 0.23418 | 0.29944 | 0.55136 | 0.23381 | 0.30332 | 0.20656 | 0.37283 | 0.40825 | 0.42308 | 0.5465 | 0.53914
| Rel.Std.Dev. |  | 2.277% | 0% | 4.763% | 5.602% | 4.152% | 4.152% | 1.81% | 1.897% | 1.211% | 3.544% | 2.704% | 2.141% | 3.604% | 1.976% | 2.835% | 1.534% | 3.2% | 3.101% | 3.292% | 4.542% | 4.757%
| PDM3 |  | 5.21% | 8.121% | -4.354% | -0.889% | -18.22% | -18.22% | -9.622% | 5.487% | 35.01% | -14.2% | -27.97% | 16.3% | 27.25% | -1.582% | -11.01% | 12% | -3.107% | 9.507% | 6.874% | 0.08132% | -5.741%
| Z-Score (Absolute) |  | 0.46 | NA | NA | 0.04 | NA | NA | 0.75 | 0.48 | 2.9 | 1.13 | 2.25 | 1.37 | 2.26 | 0.1 | 0.87 | 1.01 | 0.22 | NA | 0.59 | 0.04 | 0.44
Results from laboratories 2, 3, 7, 8 and 28 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C312: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C331: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C368: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C386: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C423: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C460: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C478: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C496: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C515: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C534: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C553: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C571: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C589: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C608: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C626: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C645: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C663: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C681: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C718: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C736: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C754: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C772: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C790: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C808: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C826: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C864: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C883: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C901: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C919: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C938: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C956: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C974: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C993: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1012: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1030: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1048: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1066: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1085: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1122: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1141: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1160: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 282 (Indicative Value 5.58 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.58 | 1
| 1 | 2 | 5.58 | 10
| Mean |  | 5.58 | 16
| Median |  | 5.58 | 5.58
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 282 (Indicative Value 25.65 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 25.63 | 1
| 1 | 2 | 25.66 | 11
| Mean |  | 25.645 | 16
| Median |  | 25.645 | 25.645
| Std Dev. |  | 0.021213 | 17
| Rel.Std.Dev. |  | 0.08272%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 282 (Indicative Value 1.54 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.53 | 1
| 1 | 2 | 1.54 | 12
| Mean |  | 1.535 | 16
| Median |  | 1.535 | 1.535
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.4607%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 282 (Indicative Value 2.18 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.17 | 1
| 1 | 2 | 2.18 | 13
| Mean |  | 2.175 | 16
| Median |  | 2.175 | 2.175
| Std Dev. |  | 0.0070711 | 19
| Rel.Std.Dev. |  | 0.3251%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 282 (Indicative Value 5.83 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.83 | 1
| 1 | 2 | 5.82 | 10
| Mean |  | 5.825 | 16
| Median |  | 5.825 | 5.825
| Std Dev. |  | 0.0070711 | 16
| Rel.Std.Dev. |  | 0.1214%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 282 (Indicative Value 0.03 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.03 | 1
| 1 | 2 | 0.03 | 11
| Mean |  | 0.03 | 16
| Median |  | 0.03 | 0.03
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 282 (Indicative Value 0.07 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.07 | 1
| 1 | 2 | 0.07 | 12
| Mean |  | 0.07 | 16
| Median |  | 0.07 | 0.07
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 282 (Indicative Value 0.074 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.073 | 1
| 1 | 2 | 0.074 | 13
| Mean |  | 0.0735 | 16
| Median |  | 0.0735 | 0.0735
| Std Dev. |  | 0.00070711 | 19
| Rel.Std.Dev. |  | 0.9621%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 282 (Indicative Value 32.96 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 32.94 | 1
| 1 | 2 | 32.97 | 10
| Mean |  | 32.955 | 16
| Median |  | 32.955 | 32.955
| Std Dev. |  | 0.021213 | 16
| Rel.Std.Dev. |  | 0.06437%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 282 (Indicative Value 1.68 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.68 | 1
| 1 | 2 | 1.678 | 11
| Mean |  | 1.679 | 16
| Median |  | 1.679 | 1.679
| Std Dev. |  | 0.0014142 | 17
| Rel.Std.Dev. |  | 0.08423%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 282 (Indicative Value 0.27 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.27 | 1
| 1 | 2 | 0.27 | 12
| Mean |  | 0.27 | 16
| Median |  | 0.27 | 0.27
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 282 (Indicative Value 24.59 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 24.61 | 1
| 1 | 2 | 24.56 | 15
| Mean |  | 24.585 | 16
| Median |  | 24.585 | 24.585
| Std Dev. |  | 0.035355 | 21
| Rel.Std.Dev. |  | 0.1438%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 282 (Indicative Value 6.76 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 2
2
| 1 | 1 | 6.72 | 1
| 1 | 2 | 6.79 | 17
| Mean |  | 6.755 | 16
| Median |  | 6.755 | 6.755
| Std Dev. |  | 0.049497 | 23
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 282 (Indicative Value 0.32 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.35 | 1
| 1 | 2 | 0.29 | 17
| Mean |  | 0.32 | 16
| Median |  | 0.32 | 0.32
| Std Dev. |  | 0.042426 | 23
| Rel.Std.Dev. |  | 13.26%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 282 (Indicative Value 0.95 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.9 | 1
| 1 | 2 | 1 | 19
| Mean |  | 0.95 | 16
| Median |  | 0.95 | 0.95
| Std Dev. |  | 0.070711 | 25
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 282 (Indicative Value 2760 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2750 | 1
| 1 | 2 | 2770 | 20
| Mean |  | 2760 | 16
| Median |  | 2760 | 2760
| Std Dev. |  | 14.1421 | 26
| Rel.Std.Dev. |  | 0.5124%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 282 (Indicative Value 105 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 104 | 1
| 1 | 2 | 106 | 21
| Mean |  | 105 | 16
| Median |  | 105 | 105
| Std Dev. |  | 1.4142 | 27
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 282 (Indicative Value 1 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1 | 22
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 282 (Indicative Value 1.2 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.2 | 23
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 282 (Indicative Value 1.45 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.5 | 24
| Mean |  | 1.45 | 16
| Median |  | 1.45 | 1.45
| Std Dev. |  | 0.070711 | 30
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 282 (Indicative Value 37.3 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.9 | 1
| 1 | 2 | 37.7 | 25
| Mean |  | 37.3 | 16
| Median |  | 37.3 | 37.3
| Std Dev. |  | 0.56569 | 31
| Rel.Std.Dev. |  | 1.517%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 282 (Indicative Value 4.9 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 5 | 26
| Mean |  | 4.9 | 16
| Median |  | 4.9 | 4.9
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 282 (Indicative Value 46 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46 | 1
| 1 | 2 | 46 | 27
| Mean |  | 46 | 16
| Median |  | 46 | 46
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 282 (Indicative Value 8.98 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.99 | 1
| 1 | 2 | 8.97 | 28
| Mean |  | 8.98 | 16
| Median |  | 8.98 | 8.98
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 0.1575%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 282 (Indicative Value 30 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26 | 1
| 1 | 2 | 34 | 29
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 5.6569 | 35
| Rel.Std.Dev. |  | 18.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 282 (Indicative Value 2.69 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.58 | 1
| 1 | 2 | 2.8 | 30
| Mean |  | 2.69 | 16
| Median |  | 2.69 | 2.69
| Std Dev. |  | 0.15556 | 36
| Rel.Std.Dev. |  | 5.783%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 282 (Indicative Value 1.54 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.54 | 1
| 1 | 2 | 1.53 | 31
| Mean |  | 1.535 | 16
| Median |  | 1.535 | 1.535
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 0.4607%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 282 (Indicative Value 0.7 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.68 | 1
| 1 | 2 | 0.72 | 32
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0.028284 | 38
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 282 (Indicative Value 8.75 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.7 | 1
| 1 | 2 | 8.8 | 33
| Mean |  | 8.75 | 16
| Median |  | 8.75 | 8.75
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 282 (Indicative Value 3.02 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.93 | 1
| 1 | 2 | 3.11 | 34
| Mean |  | 3.02 | 16
| Median |  | 3.02 | 3.02
| Std Dev. |  | 0.12728 | 40
| Rel.Std.Dev. |  | 4.215%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 282 (Indicative Value 0.55 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.55 | 1
| 1 | 2 | 0.55 | 35
| Mean |  | 0.55 | 16
| Median |  | 0.55 | 0.55
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 282 (Indicative Value 4.22 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.24 | 1
| 1 | 2 | 4.2 | 19
| Mean |  | 4.22 | 16
| Median |  | 4.22 | 4.22
| Std Dev. |  | 0.028284 | 25
| Rel.Std.Dev. |  | 0.6702%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 282 (Indicative Value 0.57 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.56 | 1
| 1 | 2 | 0.57 | 20
| Mean |  | 0.565 | 16
| Median |  | 0.565 | 0.565
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 1.252%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 282 (Indicative Value 0.063 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.05 | 1
| 1 | 2 | 0.1 | 21
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.0625
| Std Dev. |  |  | 27
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 282 (Indicative Value 21 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.9 | 1
| 1 | 2 | 21 | 22
| Mean |  | 20.95 | 16
| Median |  | 20.95 | 20.95
| Std Dev. |  | 0.070711 | 28
| Rel.Std.Dev. |  | 0.3375%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 282 (Indicative Value 0.21 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.21 | 23
| Mean |  | 0.205 | 16
| Median |  | 0.205 | 0.205
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 3.449%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 282 (Indicative Value 0.027 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0273 | 1
| 1 | 2 | 0.0274 | 24
| Mean |  | 0.02735 | 16
| Median |  | 0.02735 | 0.02735
| Std Dev. |  | 7.0711e-05 | 30
| Rel.Std.Dev. |  | 0.2585%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 282 (Indicative Value 13.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.4 | 1
| 1 | 2 | 13.8 | 25
| Mean |  | 13.6 | 16
| Median |  | 13.6 | 13.6
| Std Dev. |  | 0.28284 | 31
| Rel.Std.Dev. |  | 2.08%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 282 (Indicative Value 5.9 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.88 | 1
| 1 | 2 | 5.92 | 26
| Mean |  | 5.9 | 16
| Median |  | 5.9 | 5.9
| Std Dev. |  | 0.028284 | 32
| Rel.Std.Dev. |  | 0.4794%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 282 (Indicative Value 18.8 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.6 | 1
| 1 | 2 | 19 | 27
| Mean |  | 18.8 | 16
| Median |  | 18.8 | 18.8
| Std Dev. |  | 0.28284 | 33
| Rel.Std.Dev. |  | 1.504%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 282 (Indicative Value 34 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 28 | 1
| 1 | 2 | 40 | 28
| Mean |  | 34 | 16
| Median |  | 34 | 34
| Std Dev. |  | 8.4853 | 34
| Rel.Std.Dev. |  | 24.96%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 282 (Indicative Value 24 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 24 | 1
| 1 | 2 | 24 | 29
| Mean |  | 24 | 16
| Median |  | 24 | 24
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 282 (Indicative Value 4.89 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.83 | 1
| 1 | 2 | 4.94 | 30
| Mean |  | 4.885 | 16
| Median |  | 4.885 | 4.885
| Std Dev. |  | 0.077782 | 36
| Rel.Std.Dev. |  | 1.592%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 282 (Indicative Value 136 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 135 | 1
| 1 | 2 | 136 | 31
| Mean |  | 135.5 | 16
| Median |  | 135.5 | 135.5
| Std Dev. |  | 0.70711 | 37
| Rel.Std.Dev. |  | 0.5219%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 282 (Indicative Value 0.035 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.03 | 1
| 1 | 2 | 0.04 | 32
| Mean |  | 0.035 | 16
| Median |  | 0.035 | 0.035
| Std Dev. |  | 0.0070711 | 38
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 282 (Indicative Value 17.7 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.5 | 1
| 1 | 2 | 17.9 | 33
| Mean |  | 17.7 | 16
| Median |  | 17.7 | 17.7
| Std Dev. |  | 0.28284 | 39
| Rel.Std.Dev. |  | 1.598%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 282 (Indicative Value 4.15 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.1 | 1
| 1 | 2 | 4.2 | 34
| Mean |  | 4.15 | 16
| Median |  | 4.15 | 4.15
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 1.704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 282 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 35
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 282 (Indicative Value 3.55 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.63 | 1
| 1 | 2 | 3.46 | 19
| Mean |  | 3.545 | 16
| Median |  | 3.545 | 3.545
| Std Dev. |  | 0.12021 | 25
| Rel.Std.Dev. |  | 3.391%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 282 (Indicative Value 1.8 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.8 | 20
| Mean |  | 1.8 | 16
| Median |  | 1.8 | 1.8
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 282 (Indicative Value 283 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 281 | 1
| 1 | 2 | 285 | 21
| Mean |  | 283 | 16
| Median |  | 283 | 283
| Std Dev. |  | 2.8284 | 27
| Rel.Std.Dev. |  | 0.9994%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 282 (Indicative Value 0.44 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.42 | 1
| 1 | 2 | 0.45 | 22
| Mean |  | 0.435 | 16
| Median |  | 0.435 | 0.435
| Std Dev. |  | 0.021213 | 28
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 282 (Indicative Value 0.46 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.45 | 1
| 1 | 2 | 0.47 | 23
| Mean |  | 0.46 | 16
| Median |  | 0.46 | 0.46
| Std Dev. |  | 0.014142 | 29
| Rel.Std.Dev. |  | 3.074%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 282 (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 24
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 282 (Indicative Value 5.63 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.59 | 1
| 1 | 2 | 5.66 | 25
| Mean |  | 5.625 | 16
| Median |  | 5.625 | 5.625
| Std Dev. |  | 0.049497 | 31
| Rel.Std.Dev. |  | 0.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 282 (Indicative Value 0.163 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.163 | 1
| 1 | 2 | 0.163 | 26
| Mean |  | 0.163 | 16
| Median |  | 0.163 | 0.163
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 282 (Indicative Value 25.1 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 24 | 1
| 1 | 2 | 26.2 | 27
| Mean |  | 25.1 | 16
| Median |  | 25.1 | 25.1
| Std Dev. |  | 1.5556 | 33
| Rel.Std.Dev. |  | 6.198%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 282 (Indicative Value 0.21 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.21 | 1
| 1 | 2 | 0.21 | 28
| Mean |  | 0.21 | 16
| Median |  | 0.21 | 0.21
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 282 (Indicative Value 7.88 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.72 | 1
| 1 | 2 | 8.03 | 29
| Mean |  | 7.875 | 16
| Median |  | 7.875 | 7.875
| Std Dev. |  | 0.2192 | 35
| Rel.Std.Dev. |  | 2.784%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 282 (Indicative Value 116 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 116 | 1
| 1 | 2 | 116 | 30
| Mean |  | 116 | 16
| Median |  | 116 | 116
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 282 (Indicative Value 27 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 27 | 1
| 1 | 2 | 27 | 31
| Mean |  | 27 | 16
| Median |  | 27 | 27
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 282 (Indicative Value 17.6 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.5 | 1
| 1 | 2 | 17.6 | 32
| Mean |  | 17.55 | 16
| Median |  | 17.55 | 17.55
| Std Dev. |  | 0.070711 | 38
| Rel.Std.Dev. |  | 0.4029%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 282 (Indicative Value 1.51 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.52 | 1
| 1 | 2 | 1.5 | 33
| Mean |  | 1.51 | 16
| Median |  | 1.51 | 1.51
| Std Dev. |  | 0.014142 | 39
| Rel.Std.Dev. |  | 0.9366%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 282 (Indicative Value 140 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 140 | 1
| 1 | 2 | 140 | 34
| Mean |  | 140 | 16
| Median |  | 140 | 140
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 282 (Indicative Value 154 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 153 | 1
| 1 | 2 | 154 | 35
| Mean |  | 153.5 | 16
| Median |  | 153.5 | 153.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 0.4607%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.